LIGO Laboratory / LIGO Scientific Collaboration

LIGO-T0900322v1	Advanced LIGO UK          1 July 2009

Satellite Box test Plan

R.M. Cutler, University of Birmingham

Distribution of this document:

Inform aligo_sus

This is an internal working note

of the Advanced LIGO Project, prepared by members of the UK team.

| Institute for Gravitational Research University of Glasgow Phone +44 (0) 141 330 5884 Fax +44 (0) 141 330 6833 E-mail k.strain@physics.gla.ac.uk | School of Physics and Astronomy University of Birmingham Phone +44 (0) 121 414 6447 Fax +44 (0) 121 414 3722 E-mail av@star.sr.bham.ac.uk |
| --- | --- |
| Engineering Department CCLRC Rutherford Appleton Laboratory Phone +44 (0) 1235 445 297 Fax +44 (0) 1235 445 843 E-mail J.Greenhalgh@rl.ac.uk | Department of Physics University of Strathclyde Phone +44 (0) 1411 548 3360 Fax +44 (0) 141 552 2891 E-mail N.Lockerbie@phys.strath.ac.uk |

http://www.ligo.caltech.edu/

http://www.physics.gla.ac.uk/igr/sus/

http://www.sr.bham.ac.uk/research/gravity/rh,d,2.html

http://www.eng-external.rl.ac.uk/advligo/papers_public/ALUK_Homepage.htm

Contents

1. Introduction

2. Manual Satellite Box test Equipment

Table 1 DC measurements

Table 2 AC Measurements

Table 3 Test Results Sheet

3. Proposed Automatic Test Equipment

3.1 Test Unit

3.2 Auxiliary Equipment

4. Conclusion

1. Introduction

This report describes the tests conducted on each pre-production Satellite Box, using a specially designed manual test box.

For production, the plan is to use automatic test equipment derived from the manual system, where the switches are replaced by analogue switches under computer control. The automatic test equipment will be controlled by a PC, programmed in Visual Basic. The plan is to generate the test reports automatically.

The Automatic Test Equipment is being developed at the time of writing.

2. Manual Satellite Box test Equipment

The Pre Production satellite Boxes were tested by means of a specially made test box. On the front of the test box are two switches.

The unit under test is connected to the test box, and the switches are set at each position in turn. In each switch position, a parameter is measured by means of a precision DVM, and recorded on the test sheet.

Switch A is set to each position in turn. The last position of switch A brings Switch B into play. Switch B is then set to each position in turn.

The following measurements were made on each satellite box.

The tables on the next page give details of the tests performed.

Table 1 	DC Measurements

These measurements use the internal -10v bias reference with a series 1M resistor

| Switch A Position | Switch B Position | Parameter | Measurement |
| --- | --- | --- | --- |
| 1 |  | 5v supply | Measures the 5V rail |
| 2 |  | -ve supply | Measures the -14v rail |
| 3 |  | +ve supply | Measures the +14v rail |
| 4 |  | Ref A | Measures 2.5v reference A |
| 5 |  | Ref B | Measures 2.5v reference B |
| 6 |  | Ref C | Measures 2.5v reference C |
| 7 |  | Ref D | Measures 2.5v reference D |
| 8 |  | LED I A | Measures LED current source A |
| 9 |  | LED I B | Measures LED current source B |
| 10 |  | LED I C | Measures LED current source C |
| 11 |  | LED I D | Measures LED current source D |
| 12 | 1 | Amp gain A | Differential PD Amp A output with 10μA input |
| 12 | 2 | Mon gain A | Differential PD Mon A output with 10μA input |
| 12 | 3 | Bias V A | Measures the channel A -10v LIGO1 OSEM bias |
| 12 | 4 | Amp gain B | Differential PD Amp B output with 10μA input |
| 12 | 5 | Mon gain B | Differential PD Mon B output with 10μA input |
| 12 | 6 | Bias V B | Measures the channel B -10v LIGO1 OSEM bias |
| 12 | 7 | Amp gain C | Differential PD Amp C output with 10μA input |
| 12 | 8 | Mon gain C | Differential PD Mon C output with 10μA input |
| 12 | 9 | Bias V C | Measures the channel C -10v LIGO1 OSEM bias |
| 12 | 10 | Amp gain D | Differential PD Amp D output with 10μA input |
| 12 | 11 | Mon gain D | Differential PD Mon D output with 10μA input |
| 12 | 12 | Bias V D | Measures the channel D -10v LIGO1 OSEM bias |

Table 2 	AC Measurements

These measurements use a signal generator with a series 1M resistor.

| Switch A | Switch B |  |
| --- | --- | --- |
| 12 | 1 | Amp A frequency response |
| 12 | 2 | Mon A frequency response |
| 12 | 4 | Amp B frequency response |
| 12 | 5 | Mon B frequency response |
| 12 | 7 | Amp C frequency response |
| 12 | 8 | Mon C frequency response |
| 12 | 10 | Amp D frequency response |
| 12 | 11 | Mon D frequency response |

The r.m.s output is measured at 5Hz, 10Hz, 100Hz, 1KHz and 5 KHz with 1v r.m.s input.

The test report on the next page is designed to summarise the tests and record the results of the tests for each unit.

Table 3 	Test Result Sheet

E&OE								RWP

3. Proposed Automatic Test Unit

The idea of the auto test box is to make the testing of the satellite boxes easier, quicker and more consistent.

The unit will be connected to the satellite boxes via the four connectors J1, J2, J3 and J4 and will provide the necessary power.

Test unit will measure and record:

Time, date and serial number.

The +/-17V supplies current and voltages.

The internal regulated supply voltages (scaled +/-14V and 5V).

The internal reference voltages (1.00V)

The LED drive currents (35mA). The operation of the out of range indicators will      be done visually.

The gain of the amplifiers. Differential output, dc gain and ac gain at selected  frequencies

Photo diode bias voltages (10V)

# 3.1 Automatic Test Unit

It will carry out its own internal checks: amplifier dc and ac signal levels.

The signal selection will be done with analogue switches and relays controlled by an 8bit word from a computer interface.

There will be voltage regulators for the internal circuit and for output to the satellite box, also a voltage reference for the gain measurements together with signal conditioning amplifiers.

# 3.2 Auxiliary Equipment

Bench supply: +/- 24V

KEITHLEY DVM with IEEE interface

HP signal generator with IEEE interface

USB to IEEE pod

Computer

Connection leads

# 4. Conclusion

The manual test equipment was used to test the Pre-Production units quickly and efficiently. As large numbers of production units are planned, the planned Automatic test equipment will make the testing of the Production units easier and quicker, with consistent results.