LIGO Laboratory / LIGO Scientific Collaboration

LIGO- T1000133_v1	Advanced LIGO UK      March 2010

Triple Acquisition Driver Unit Test Report

R. M. Cutler, University of Birmingham

Distribution of this document:

Inform aligo_sus

This is an internal working note

of the Advanced LIGO Project, prepared by members of the UK team.

| Institute for Gravitational Research University of Glasgow Phone +44 (0) 141 330 5884 Fax +44 (0) 141 330 6833 E-mail k.strain@physics.gla.ac.uk | School of Physics and Astronomy University of Birmingham Phone +44 (0) 121 414 6447 Fax +44 (0) 121 414 3722 E-mail av@star.sr.bham.ac.uk |
| --- | --- |
| Engineering Department CCLRC Rutherford Appleton Laboratory Phone +44 (0) 1235 445 297 Fax +44 (0) 1235 445 843 E-mail J.Greenhalgh@rl.ac.uk | Department of Physics University of Strathclyde Phone +44 (0) 1411 548 3360 Fax +44 (0) 141 552 2891 E-mail N.Lockerbie@phys.strath.ac.uk |

http://www.ligo.caltech.edu/

http://www.physics.gla.ac.uk/igr/sus/

http://www.sr.bham.ac.uk/research/gravity/rh,d,2.html

http://www.eng-external.rl.ac.uk/advligo/papers_public/ALUK_Homepage.htm

TRIPLE ACQUISITION DRIVER UNIT BOARD TEST REPORT

Unit……………………………………Serial No ………………………………

Test Engineer ………

Date …………………

Drive Card ID……………

Monitor Card ID ………..

Contents

1	 Description

2	 Test Equipment

3	 Inspection

4	 Continuity Checks

5	 Test Set Up

6	 Power

7	 Relay operation

8 	Monitor Outputs

8.1 	Amplifier Monitors

8.2 	Coil Monitors

8.3	R.M.S Monitors

8.4	Noise Monitors

9. 	Distortion

10 	Load Tests

10.1 	Noisy Mode

10.2 	Low noise Mode

10.3 	Acquisition Mode

11 	Noise Measurements

12 	Final Assembly Tests

1. Description

Block diagram

Description

The Acquisition unit consists of four identical channels and the power regulators which provide regulated power to the four channels. Each channel consists of a coil drive channel, and monitor circuitry.

The driver has 3 main modes of operation, selectable by two external relay commands: Noisy Mode, Quiet Mode and Acquisition Mode. There is also a mode which switches the channel off.

Unit……………………………………Serial No ………………………………

Test Engineer ………

Date …………………

2. Test equipment

Power supplies (At least +/- 20v variable, 1A)

Signal generator (capable of delivering 10v peak, 0.1Hz to 10 KHz))

Analogue oscilloscope

Agilent Dynamic Signal Analyser (or similar)

Low noise Balanced Driver circuit

Relay test box

Record the Models and serial numbers of the test equipment used below.

| Unit (e.g. DVM) | Manufacturer | Model | Serial Number |
| --- | --- | --- | --- |

Unit……………………………………Serial No ………………………………

Test Engineer ………

Date …………………

3. Inspection

Workmanship

Inspect the general workmanship standard and comment:

Links:

Check that the links W2 and W4 are present on each channel.

Unit……………………………………Serial No ………………………………

Test Engineer ………

Date …………………

4. Continuity Checks

Continuity to the V, I and R.M.S Monitor (J1)

PD out to AA

| PIN | SIGNAL | DESCRIPTION | To J1 PIN | OK? |
| --- | --- | --- | --- | --- |
| 1 | PD1P | Photodiode A+ | 1 |  |
| 2 | PD2P | Photodiode B+ | 2 |  |
| 3 | PD3P | Photodiode C+ | 3 |  |
| 4 | PD4P | Photodiode D+ | 4 |  |
| 5 | 0V |  |
| 6 | PD1N | Photodiode A- | 14 |  |
| 7 | PD2N | Photodiode B- | 15 |  |
| 8 | PD3N | Photodiode C- | 16 |  |
| 9 | PD4N | Photodiode D- | 17 |  |

LED Mon

| PIN | SIGNAL |  | To J1 PIN | OK? |
| --- | --- | --- | --- | --- |
| 1 | Imon1P |  | 5 |  |
| 2 | Imon2P |  | 6 |  |
| 3 | Imon3P |  | 7 |  |
| 4 | Imon4P |  | 8 |  |
| 5 | 0V |  |
| 6 | Imon1N |  | 18 |  |
| 7 | Imon2N |  | 19 |  |
| 8 | Imon3N |  | 20 |  |
| 9 | Imon4N |  | 21 |  |

Pd from Sat

| PIN | SIGNAL | DESCRIPTION | OK? |
| --- | --- | --- | --- |
| 9 | V+ (TP1) | +17v Supply |  |
| 10 | V+ (TP1) | +17v Supply |  |
| 11 | V- (TP2) | -17v Supply |  |
| 12 | V- (TP2) | -17v Supply |  |
| 13 | 0V (TP3) |  |  |
| 22 | 0V (TP3) |  |  |
| 23 | 0V (TP3) |  |  |
| 24 | 0V (TP3) |  |  |
| 25 | 0V (TP3) |  |  |

5. TEST SET UP

Note:

(1)  Input signal to differential amplifier is generally stated in the tests below.

There is therefore an inherent gain of 2 in the system.

(2) Some signal generators will indicate 1vpk/pk when the output is in fact 1v Peak into the high impedance Differential driver used. The test procedure refers to the actual voltage out of the signal generator.

Connections:

Differential signal inputs to the board under test:

Drive Input J3 pins 1, 2, 3, 4 = positive input

Drive Input J3 pins 6, 7, 8, 9 = negative input

Drive Input J3 pin 5 = ground

Power

DC IN J1 pin 9, 10 = +16.5v

DC IN J1 pin 11,12 = -16.5

DC IN J1 pins 22, 23, 24, 25 = 0v

Outputs

Coil Out to Sat (J4)

Ch1+ = J4 pin 1		Ch1- = J4 pin 9

Ch2+ = J4 pin 3		Ch2- = J4 pin 11

Ch3+ = J4 pin 5		Ch3- = J4 pin 13

Ch4+ = J4 pin 7 		Ch4- = J4 pin 15

Unit……………………………………Serial No ………………………………

Test Engineer ………

Date …………………

6. Power

Check that the 3 pin power connector is wired correctly: A1 positive, A2 return, A3 Negative.

Set the power supply outputs to zero.

Connect power to the unit

Increase the voltages on the supplies to +/-3V.

Determine that the supply polarities are correct on TP1 and TP2.

If they are, increase input voltages to +/- 16.5v.

Record the output voltages, measured on a DVM with 4 or more digits, from each regulator

Observe the output on an analogue oscilloscope, set to AC. Measure and record the peak to peak noise on each regulator output.

Record regulator outputs:

| Regulator | Output voltage | Nominal +/- 0.5v? | Output noise |
| --- | --- | --- | --- |
| +12v TP5 |  |  |  |
| +15v TP4 |  |  |  |
| -15v TP6 |  |  |  |

| All Outputs smooth DC, no oscillation? |  |
| --- | --- |

Record Power Supply Currents

| Supply | Current |
| --- | --- |
| +16.5v |  |
| -16.5v |  |

If the supplies are correct, proceed to the next test.

Unit……………………………………Serial No ………………………………

Test Engineer ………

Date …………………

7. Relay Operation

Operate each relay in turn.

Observe its operation. LEDs should illuminate when the relays are operated.

Filter

| Channel | Indicator | OK? |
| --- | --- | --- |
|  | ON | OFF |  |
| Ch1 |  |  |  |
| Ch2 |  |  |  |
| Ch3 |  |  |  |
| Ch4 |  |  |  |

TEST RELAYS

| Channel | Indicator | OK? |
| --- | --- | --- |
|  | ON | OFF |  |
| Ch1 |  |  |  |
| Ch2 |  |  |  |
| Ch3 |  |  |  |
| Ch4 |  |  |  |

ACQUISITION RELAYS

| Channel | Indicator | OK? |
| --- | --- | --- |
|  | ON | OFF |  |
| Ch1 |  |  |  |
| Ch2 |  |  |  |
| Ch3 |  |  |  |
| Ch4 |  |  |  |

Unit……………………………………Serial No ………………………………

Test Engineer ………

Date …………………

8. Monitor Outputs

Switch out the filters and set the unit to Acquisition Mode.

With a 20 ohm dummy load on each channel, apply an input from the signal generator at 1 KHz, and adjust the amplitude until the output is 1vr r.m.s as measured between TP4 and TP5.

Measure the Voltage Monitor outputs with respect to 0v for each channel.

8.1 Voltage Monitors

| Ch. | Output: | V, I and R.M.S Monitor | Expected value | Pass/Fail: Equal? (+/- 0.1v) |
| --- | --- | --- | --- | --- |
| 1 |  | 10 | 0.33v |  |
| 2 |  | 7 | 0.33v |  |
| 3 |  | 4 | 0.33v |  |
| 4 |  | 1 | 0.33v |  |

Adjust the input voltage until the voltage across the load resistor = 1v rms.  Record the current monitor output values.

8.2 Current Monitors

| Ch. | Output | V, I and R.M.S Monitor | Expected Value | Pass/Fail: Equal? (+/- 0.1v) |
| --- | --- | --- | --- | --- |
| 1 |  | 11 | 1.86v r.m.s |  |
| 2 |  | 8 | 1.86v r.m.s |  |
| 3 |  | 5 | 1.86v r.m.s |  |
| 4 |  | 2 | 1.86v r.m.s |  |

8.3 R.M.S Monitors

| Ch. | Output | V, I and R.M.S Monitor | Expected Value | Pass/Fail: Equal? (+/- 0.1v) |
| --- | --- | --- | --- | --- |
| 1 |  | 12 | 1.86v dc |  |
| 2 |  | 9 | 1.86v dc |  |
| 3 |  | 6 | 1.86v dc |  |
| 4 |  | 3 | 1.86v dc |  |

8.4 Noise Monitors

Using the Pre-Amplifier with a gain of 10 and Dynamic Signal Analyser, measure the noise monitor outputs in uV/√Hz on the noise monitor outputs. Correct for the pre-amplifier gain. 10pA/rt Hz should give 0.825μV/μHz out.

| Ch. | Output | /(Pre-amplifier gain) | Expected Value | Comparison |
| --- | --- | --- | --- | --- |
| 1 |  |  | 0.825μV/μHz |  |
| 2 |  |  | 0.825μV/μHz |  |
| 3 |  |  | 0.825μV/μHz |  |
| 4 |  |  | 0.825μV/μHz |  |

Unit……………………………………Serial No ………………………………

Test Engineer ………

Date …………………

9. Distortion

Switch the filters out. Increase input voltage to 5v peak, f = 1KHz. Use 20 Ohm loads. Observe the voltage across each load with an oscilloscope in both Acquisition and Non-Acquisition modes.

|  | Acquisition Mode: Distortion Free? | Non-Acquisition Mode: Distortion Free? |
| --- | --- | --- |
| Ch1 |  |  |
| Ch2 |  |  |
| Ch3 |  |  |
| Ch4 |  |  |

Unit……………………………………Serial No ………………………………

Test Engineer ………

Date …………………

10 Load tests and Frequency response check

Plug in the 20 Ohm 5W loads. Ensure the links W4 are in place.

10.1 Noisy Mode

With the acquisition mode switched out, and filters switched out, apply 5v peak at the input to the drive unit. Measure the r.m.s differential voltage across each load resistor in turn using a true r.m.s meter, at the frequencies below. Calculate the output current in each case (Vout/20).

1Hz

|  | Vo r.m.s | Peak Io (Vo/20) x 1.414 | Specification | Pass/Fail |
| --- | --- | --- | --- | --- |
| Ch1 |  |  | 2.5 mA peak |  |
| Ch2 |  |  | 2.5 mA peak |  |
| Ch3 |  |  | 2.5 mA peak |  |
| Ch4 |  |  | 2.5 mA peak |  |

10Hz

|  | Vo r.m.s | Peak Io (Vo/20) x 1.414 | Specification | Pass/Fail |
| --- | --- | --- | --- | --- |
| Ch1 |  |  | 2.5 mA peak |  |
| Ch2 |  |  | 2.5 mA peak |  |
| Ch3 |  |  | 2.5 mA peak |  |
| Ch4 |  |  | 2.5 mA peak |  |

100Hz

|  | Vo r.m.s | Peak Io (Vo/20) x 1.414 | Specification | Pass/Fail |
| --- | --- | --- | --- | --- |
| Ch1 |  |  | 2.5 mA peak |  |
| Ch2 |  |  | 2.5 mA peak |  |
| Ch3 |  |  | 2.5 mA peak |  |
| Ch4 |  |  | 2.5 mA peak |  |

200Hz

|  | Vo r.m.s | Peak Io (Vo/20) x 1.414 | Specification | Pass/Fail |
| --- | --- | --- | --- | --- |
| Ch1 |  |  | 2.5 mA peak |  |
| Ch2 |  |  | 2.5 mA peak |  |
| Ch3 |  |  | 2.5 mA peak |  |
| Ch4 |  |  | 2.5 mA peak |  |

1 KHz

|  | Vo r.m.s | Peak Io (Vo/20) x 1.414 | Specification | Pass/Fail |
| --- | --- | --- | --- | --- |
| Ch1 |  |  | 2.5 mA peak |  |
| Ch2 |  |  | 2.5 mA peak |  |
| Ch3 |  |  | 2.5 mA peak |  |
| Ch4 |  |  | 2.5 mA peak |  |

Unit……………………………………Serial No ………………………………

Test Engineer ………

Date …………………

10.2 Low noise Mode

With the acquisition mode switched out and filters switched in, apply 5v peak at the input to the drive unit. Measure the r.m.s differential voltage across each load resistor in turn using a true r.m.s, at the frequencies below.

Calculate the output current in each case (Vout/20).

1Hz

|  | Vo r.m.s | Peak Io (Vo/20) x 1.414 | Specification | Pass/Fail |
| --- | --- | --- | --- | --- |
| Ch1 |  |  | 2.5 mA peak |  |
| Ch2 |  |  | 2.5 mA peak |  |
| Ch3 |  |  | 2.5 mA peak |  |
| Ch4 |  |  | 2.5 mA peak |  |

10Hz

|  | Vo r.m.s | Peak Io (Vo/20) x 1.414 | Specification | Pass/Fail |
| --- | --- | --- | --- | --- |
| Ch1 |  |  | 2.5 mA peak |  |
| Ch2 |  |  | 2.5 mA peak |  |
| Ch3 |  |  | 2.5 mA peak |  |
| Ch4 |  |  | 2.5 mA peak |  |

100Hz

|  | Vo r.m.s | Peak Io (Vo/20) x 1.414 | Specification | Pass/Fail |
| --- | --- | --- | --- | --- |
| Ch1 |  |  | 2.5 mA peak |  |
| Ch2 |  |  | 2.5 mA peak |  |
| Ch3 |  |  | 2.5 mA peak |  |
| Ch4 |  |  | 2.5 mA peak |  |

200Hz

|  | Vo r.m.s | Peak Io (Vo/20) x 1.414 | Specification | Pass/Fail |
| --- | --- | --- | --- | --- |
| Ch1 |  |  | 2.5 mA peak |  |
| Ch2 |  |  | 2.5 mA peak |  |
| Ch3 |  |  | 2.5 mA peak |  |
| Ch4 |  |  | 2.5 mA peak |  |

1 KHz

|  | Vo r.m.s | Peak Io (Vo/20) x 1.414 | Specification | Pass/Fail |
| --- | --- | --- | --- | --- |
| Ch1 |  |  | 2.5 mA peak |  |
| Ch2 |  |  | 2.5 mA peak |  |
| Ch3 |  |  | 2.5 mA peak |  |
| Ch4 |  |  | 2.5 mA peak |  |

Unit……………………………………Serial No ………………………………

Test Engineer ………

Date …………………

10.3 Acquisition Mode

With the acquisition mode switched in, and filters switched out, apply 5v peak at the input to the drive unit. Measure the r.m.s differential voltage across each load resistor in turn using a true r.m.s meter at the frequencies below. Calculate the peak voltages, then the peak output current in each case (Vout/20).

100Hz

|  | Vo r.m.s | Vo pk. | Peak Io (Vo/20) x 1.414 | Specification | Pass/Fail |
| --- | --- | --- | --- | --- | --- |
| Ch1 |  |  |  | 125mA peak |  |
| Ch2 |  |  |  | 125mA peak |  |
| Ch3 |  |  |  | 125mA peak |  |
| Ch4 |  |  |  | 125mA peak |  |

200Hz

|  | Vo r.m.s | Vo pk. | Peak Io (Vo/20) x 1.414 | Specification | Pass/Fail |
| --- | --- | --- | --- | --- | --- |
| Ch1 |  |  |  | 125mA peak |  |
| Ch2 |  |  |  | 125mA peak |  |
| Ch3 |  |  |  | 125mA peak |  |
| Ch4 |  |  |  | 125mA peak |  |

1 KHz

|  | Vo r.m.s | Vo pk. | Peak Io (Vo/20) x 1.414 | Specification | Pass/Fail |
| --- | --- | --- | --- | --- | --- |
| Ch1 |  |  |  | 125mA peak |  |
| Ch2 |  |  |  | 125mA peak |  |
| Ch3 |  |  |  | 125mA peak |  |
| Ch4 |  |  |  | 125mA peak |  |

5 KHz

|  | Vo r.m.s | Vo pk. | Peak Io (Vo/20) x 1.414 | Specification | Pass/Fail |
| --- | --- | --- | --- | --- | --- |
| Ch1 |  |  |  | 125mA peak |  |
| Ch2 |  |  |  | 125mA peak |  |
| Ch3 |  |  |  | 125mA peak |  |
| Ch4 |  |  |  | 125mA peak |  |

Unit……………………………………Serial No ………………………………

Test Engineer ………

Date …………………

11.  Noise Measurements

As the previous test involves non – representative temperature rises, allow the unit to cool before performing this test.

Replace the filter links W4, on each channel.

Connect the filter test box, and switch in all filters.

Switch it out of Test Mode and out of the Acquisition mode

Use the HP 35670A Dynamic Signal Analyser.

Connect a shorting plug to the demand input to short all positive and negative demands together and to 0v. Connect 20 Ohm loads to the outputs.

Use Stuart Aston’s noise measurement set up, loaded from disc.

Measure the noise output from each channel in turn at the amplifier outputs (TP4 and TP5). The Low Pass filter on the SR650 may be used to reduce mains interference, to prevent the Signal Analyser from overloading. Ideally the filter corner frequency should be set to 3 KHz. Set the amplifier gain to 1000, and check that the overload light is not on before each measurement.

Measure the noise output at 10 Hz.

|  | Spec in dB V/√Hz | Measured @ 10Hz | -60dB = | Measured in nV/√Hz | OK? |
| --- | --- | --- | --- | --- | --- |
| Ch1 | -143.5 |  |  |  |  |
| Ch2 | -143.5 |  |  |  |  |
| Ch3 | -143.5 |  |  |  |  |
| Ch4 | -143.5 |  |  |  |  |

Notes:

Specified noise output current at 10 Hz = 10pA/root Hz (worst case)

Total resistance at 10Hz, in Low noise mode = 6.7K

Amplifier noise voltage should therefore be = 67 nV/√Hz

67 nV/√Hz = -143.5 dB/√Hz

Unit……………………………………Serial No ………………………………

Test Engineer ………

Date …………………

12. Final Assembly Tests

1. Remove the lid of the box.

2. Unplug all external connections.

3. Check that the 9mm pillars are in place in the corners of the Monitor Board towards the centre of the box.

4. Check that all internal connectors are firmly mated.

5. Tighten the screw-locks holding all the external connectors.

6. Check that the nuts holding the tabs of the power drivers are secure – tighten as necessary. Test with a DVM that none of the tabs are shorted to chassis.

7. Check that all the LEDs are nicely centred.

8. Check that all links W4 and W2 are in place.

9. Check that the boards are labelled with their Drawing Number, Issue Number, and serial number. Record below:

| UoB box ID |  |
| --- | --- |
| Driver board ID |  |
| Driver board Drawing No/Issue No |  |
| Driver board Serial Number |  |
| Monitor board ID |  |
| Monitor board Drawing No/Issue No |  |
| Monitor board Serial Number |  |

10. Check the security of any modification wires.

11. Visually inspect.

12. Put the lid on and fasten all screws,

Check all external screws for tightness.

# Text boxes

Signal

Generator

Differential

Driver

Unit Under Test

Oscilloscope

Power Supply

Mode

Control

Filter

In/out

Acquisition Mode

Triple Acquisition drive board

Monitor

Board

Unit under Test

Power Supply

HP 35670A Signal Analyser

Demand

Inputs

0V

SR560