## Slide 1: Advanced-LIGO optical lever progress
Richard Abbott, Eric Black, Tara Chalermsongsak, Craig Conley, Vladimir Dergachev,
Doug Cook, Riccardo DeSalvo, Andrea Lottarini, Gerardo Moreno, Mohana Mageswaran,
Rick Savage, Robert Schofield, Michael Smith, Cheryl Vorvick, Hiro Yamamoto
LIGO-G1000309

## Slide 2: Initial LIGO Optical Lever Setup
Steering mirror
Collimator
Diode Laser
Quadrand Photodiode
Fiber optic
mirror

## Slide 3
1) Straight -> Angle cut fiber
3) use direct pointing,better support structure
2) Laser -> SLED
AdvLIGO Optical Lever Setup
Superluminescent LED
Quadrand Photodiode
With x-y translational
stage
Fiber optic
3/15/10
3
LSC meeting in Pasadena
mirror

## Slide 4: How to reconstruct beam position from QPD’s signal
V1
V4
V2
V3
Vsum (S)= V1+V2+V3+V4
Vsum corresponds to the total power of the beam
Right (R) = V1+ V4
Left (L) = V2 + v3
Top (T) = V1+V2
Bottom (B) = V3 + V4
Vx_normalized = (R-L) / S
Vy_normalized = (T-B) / S

## Slide 5
Scan the beam across the QPD, the signal (R-L)/S will be
the Error function.
slope at center ~ const/w
Near center
(R-L)/S = const * d
Then the voltage readout
Can be reconstructed to
distance from the center.
angle = d/armlength.
(R – L)/S =
(R-L)/S
d
Notes: Once we get the slope, spot size
Can be calculated.

## Slide 6: 1) Changing from straight cut to angle cut fiber connector
Why?
  Vibration on PC (Physical Contact)fiber causes the fluctuation on apparent beam position
  Back reflelction -> mode jump in the laser
  Angle cut connector,APC, (Angled Physical Contact) reduces the effect of the reflection-> immune to vibration
Notes:   Reflection of light back to the laser and cause it to mode jump

## Slide 7: Test of APC and PC fiber
A loud speaker, driven by a function generator, shakes the test fiber.
Fiber ends are strain relieved

## Slide 8: Result 1) the beam at center
3/15/10
draft
8
APC -- S
PC -- , S
S[Volt]
APC Position [micro rad]
PC – Position [micro rad]
quiet
shake
position micro rad
Volt
Time [s]
Time [s]
Notes: PC sum sees the fluctuation, but its signal does not because it is at the center

## Slide 9: Result 2) the beam off center
Plot of reconstructed x-position vs time[s] between PC (red) and APC (blue), 100 micron away from the gap. The APC fiber can maintain power stability much better than the PC fiber during the shake ( 30s < t < 70s.)
Quiet shake quiet
APC – position micro rad
PC – position micro rad
quiet
quiet
shake
position micro rad
Time [s]

## Slide 10: Another Instabilities
The oscilloscope is in x-y mode
X axis -> R – L (hirozontal)
Y axis -> T – B (vertical)
At certain areas, the apparent
beam position fluctuates.
The fluctuation subsides as
The beam goes to the center
Or far away in one quad
Notes: It seems that there is mode hopping in the laser, So we decide to find a new light source.

## Slide 11: 2) SLED as a Light Source
Current/temperature fluctuation -> mode hopping -> position fluctuation
With SLED -> no mode hopping -> no current-fluctuation induced instability
Q-photonics Super-luminescent Diodes, pigtail single mode fiber 14-Pin DIL, typically driven by a ITC502-IEEE remote control - Benchtop Laser Diode and Temperature Controller ±200 mA / 16 W

## Slide 12: Current Modulation
CCD will measure x-y positions and
spot size.
Triangular function 32+/- 2mA period of 1000 seconds
collimator

## Slide 13: Sum Voltage readout vs Time During Current Modulation
SLED
Diode Laser
Volt
Time [s]
Volt
Time [s]

## Slide 14: Influence of input current on Normalized Beam Position Data from QPD
3/15/10
14
draft
SLED
Laser

## Slide 15: Trace on the Oscilloscope
R-L and T-B (un-normalized Voltage) are displayed on x and y axes of
the scope. The spot is off center. The current is modulated by
the triangular function.
SLED
Laser
Notes: A perfect gaussian beam should grow linearly with powre, the deviation can beseen when the spot is close to the center, away from the center, the spot grow linearly for both laser and SLED

## Slide 16: Trace on the Oscilloscope (laser)

## Slide 17: SLED results from LHO
Red is SLED, others are lasers.
Kawabe K, Hanford elog, Feb 23 19:08:05 2010 UTC

## Slide 18: SLED results from LHO (2)
Atkinson D. Hanford Elog, Feb22 22:57:46 2010 UTC
free swinging noise read out by the optical levers

## Slide 19: SLED results from LHO (3)
Power decreases
By ~ 7 % over 16 days
Rodruck M. Hanford elog, Mar 11 19:51:47 2010 URC

## Slide 20: 3) Mounting Mechanism
Fifty mm diameter test mass telescope launcher on its yaw-tilt micrometric pointing mechanism; note the locking mechanisms.

## Slide 21: Pylon
The pylon is made with sheet metal laser cut, bent and stitch welded.
The pylon will be grouted and bolted on the ground

## Slide 22: Observing Thermal Distortion
3/15/10
22
LSC meeting in Pasadena
Main beam
Two auxiliary beams
Mirror

## Slide 23: summary
Use angle cut fiber connector to reduce vibration induced instabilites
Use SLED instead of a diode laser -> no mode hopping effect due to current fluctuation.
More rigid mounting and locking mechanism to ensure better stability.
Add auxiliary beams to monitor ROC of the mirrors. Fitting algorithm and CCDs will be used for obtaining the data