## Sheet1
FOR USE WITH A TRIPLE PENDULUM |  |  | HSTS BLADES | T1000352-v2
updated 10 Oct 2008 NAR | [5 empty] | KEY: - INPUT BLUE
|  |  |  | v2 - updated max stress |  | OUTPUT RED | [5 empty] | HAVE L and h NEED TO GET SAME
1.CALCULATION OF VARIOUS BLADE PARAMETERS, |  |  |  | 2010-08-01
updated 13 March 2009 NAR |  | lower blades, high stress |  | v3 updated 6 Aug2010 | comment on m3 | [6 empty] | VALUE OF h for a VALUE OF R
further updated 30 March 2009 - lower blades |  |  |  | NAR and CIT
Further updated 8th October 2009 revised masses from M Meyer SW design |  |  |  |  | NAR | [6 empty] | Length,L | 250 | mm
Further updated June 2010 for production blades. |  |  |  |  | NAR | [6 empty] | deflection,l | 222.6462 | mm
l
1.UPPER BLADES | [10 empty] | R | mm | deflection, h | mm
[11 empty] | 20 |  |  | 0.044034
thickness | 1.3 | mm | 0.0013 | m | masstotal / blade | 4.483 | kg |  |  |  | 30 |  |  | 43.8361
length,l | 25 | cm | 0.25 | m | mass supported | 1.564 | kg |  |  |  | 35 |  |  | 12.156
width,a | 3.9878 | cm | 0.039878 | m | load, P | 43.9782 |  |  |  |  | 45 |  |  | 11.3961
[5 empty] | cross sectional | 7.301e-12 |  |  |  |  | 55 |  |  | 64.1388
mass1 | 3.128 | kg | [8 empty] | 65 |  |  | 114.5232
clamps etc… | 0 | kg | [8 empty] | 75 |  |  | 148.6256
MASS 1 | 3.128 | kg |  |  | Deflection | 22.2646 | cm | 8.7656 | inch |  | 85 |  |  | 168.2986
mass2 | 2.969 | kg | [8 empty] | 95 |  |  | 177.9101
mass3 | 2.869 | kg | this is for one degree wedge which is lighter |  | frequency | 1.7886 | Hz |  |  |  | 105 |  |  | 181.0615
no. of blades | 2 |  | version of HSTS. 0.5 deg wedge mass is 20gm heavier. | [7 empty] | 115 |  |  | 180.2294
|  |  | Can adjust m2 to keep overall mass same |  | Stress | 978.8337 | Mpa |  |  |  | 125 |  |  | 177.0184
Young's modulus | 186000000000 |  |  |  |  | Good |  |  |  |  | 135 |  |  | 172.4449
shape factor | 1.32 |  | used for IMC blade prototypes |  | MAX STRESS | 1000 |  |  |  |  | 145 |  |  | 167.1475
[11 empty] | 155 |  |  | 161.5246
shape / radius / deflection |  |  |  |  | INTERNAL MODE (SCALLED FROM ORIGINAL BLADE) | [5 empty] | 165 |  |  | 155.8233
1.3 194 140 |  |  |  |  | length |  | 0.37 |  |  |  | 175 |  |  | 150.1945
19th April: Norna and Calum using inferred |  |  |  |  | thickness |  | 0.002 | 78.3068 | Hz |  | 185 |  |  | 144.7277
result of the GEO-SR blades - Stefan |  |  |  |  | internal mode |  | 55 |  |  |  | 195 |  |  | 139.4739
MEASURED | [7 empty] | radius from R=EI/M |  | 0.12351
length | 25 | 0.25 |  | inferred shape factor |  |  |  |  | where M =Pl
thickness | 1.5 | 0.0015 |  | 1.2981
large width | 4 | 0.04 |  | small width | 0.016
deflection | 16.2 | 0.162 |  | ratio of width, beta | 0.4
cross-sectional |  | 1.125e-11
Load | 5111 | 50.1389 |  | shape factor | 1.2029
1b. LOWER BLADES | [11 empty] | HAVE L and h NEED TO GET SAME
thickness | 0.76 | mm | 0.00076 | m | masstotal / blade | 1.4595 | kg |  |  |  |  | VALUE OF h for a VALUE OF R
length,l | 12 | cm | 0.12 | m | mass supported | 0.74225 | kg
width,a | 1.8 | cm | 0.018 | m | load, P | 14.3177 | [5 empty] | Length,L | 120 | mm
[5 empty] | cross sectional | 6.5846e-13 | [5 empty] | deflection,l |  | mm
mass2 | 2.969 | kg | [11 empty] | l
mass3 | 2.869 | kg |  |  | Deflection | 10.3698 | cm | 4.0826 | inch |  | R | mm | deflection, h | mm
no. of blades | 4 | [9 empty] | 20 |  |  | 0.79659
Young's modulus | 186000000000 |  |  |  | frequency | 2.1707 | Hz |  |  |  | 30 |  |  | 49.6093
shape factor | 1.54 |  | used for IMC blade prototypes | [7 empty] | 35 |  |  | 68.5686
[5 empty] | Stress | 991.5301 | Mpa |  |  |  | 45 |  |  | 85.0197
[6 empty] | Good |  |  |  |  | 55 |  |  | 86.5538
[11 empty] | 65 |  |  | 82.6729
[11 empty] | 75 |  |  | 77.19
|  |  | INTERNAL MODE (SCALLED FROM ORIGINAL BLADE) | [7 empty] | 85 |  |  | 71.5392
|  |  | length |  | 0.37 | [5 empty] | 95 |  |  | 66.2332
|  |  | thickness |  | 0.002 | 198.6951 | Hz |  |  |  | 105 |  |  | 61.4253
|  |  | internal mode |  | 55 | [5 empty] | 115 |  |  | 57.13
[11 empty] | 125 |  |  | 53.31
[11 empty] | 152 |  |  | 44.9587
[5 empty] | radius from R=EI/M | 0.071284 |  |  |  |  | 162 |  |  | 42.449
[11 empty] | 172 |  |  | 40.1898
[11 empty] | 182 |  |  | 38.1479
[11 empty] | 192 |  |  | 36.2951
[11 empty] | 202 |  |  | 34.6076
2.92
2.92
5.84