LIGO Laboratory / LIGO Scientific Collaboration

LIGO- E1000327 	        LIGO                          April 11th, 2011

aLIGO HAM-ISI, Pre integration Test report, Phase 1,

LLO Unit #3

E1000327 – V2

Joe Hanson, Adrien Le Roux, Michael Vargas, Celine Ramet, Vincent Lhuillier

Distribution of this document:

Advanced LIGO Project

This is an internal working note

of the LIGO Laboratory

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 1970 Mail Stop S9-02 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

.

Table of contents:

Introduction	3

I.	Pre-Assembly Testing	4

	Step 1: Position Sensors	4

	Step 2: GS13	5

o	Step 2.1 – Horizontal GS-13s	6

o	Step 2.2 – Vertical GS-13s	7

	Step 3: Actuators	9

II.	Tests to be performed during assembly	10

	Step 1: Check torques on all bolts	10

	Step 2: Check gaps under Support Posts	10

	Step 3: Pitchfork/Boxwork flatness before Optical Table install	10

	Step 4: Blade spring profile	10

	Step 5: Gap checks on actuators	11

	Step 6: Check level of Stage 1 optical table	11

	Step 7: Mass budget	12

	Step 8: Lockers adjustment	14

III.	Tests to be performed after assembly	15

	Step 1 - Actuators Inventory	15

	Step 2 - Sensors Inventory	15

	Step 3 - Electronics Inventory	16

	Step 4 - Set up sensors gap	16

	Step 5 - Measure the Sensor gap	17

	Step 6 - Check Sensor gaps after the platform release	17

	Step 7 - Check range of motion (hand pushing)	18

o	Step 7.1 – Test Nº1	18

o	Step 7.2 – Test Nº2	18

	Step 8 - Capacitive position sensor Power Spectrum	19

	Step 9 - GS13 power spectrums	20

o	Step 9.1 –Table locked and unlocked (free)	20

o	Step 9.2 – Table tilted	21

	Step 10 - Coil Driver, cabling and resistance check	23

	Step 11 - Actuators Sign and range of motion (Local drive)	24

o	Step 11.1 - Actuators sign	24

o	Step 11.2 - Range of motion - Local drive	24

	Step 12 - Vertical Sensor Calibration	25

	Step 13 - Vertical Spring Constant	26

	Step 14 - Static Testing (Tests in the local basis)	27

	Step 15 - Linearity test	28

	Step 16 – Static tests in the general coordinate basis	30

	Step 17 - Frequency response	30

o	Step 17.1 - Local to local measurements	30

o	Step 17.2 - Cartesian to Cartesian measurements	33

	Step 18 - Transfer function comparison	36

o	Step 18.1 - Local to local - Comparison with HAM6	36

o	Step 18.2 - Cartesian to Cartesian - Comparison with HAM6	39

	Step 19 - Lower Zero Moment Plane	43

	Step 20 - Damping loops	44

o	Step 20.1 - Transfer functions - Simulation	44

o	Step 20.2 - Powerspectrum – Experimental	44

Introduction

This document presents the tests performed to characterize and validate the “HAM-ISI LLO Unit #3”. This unit is the 3rd unit assembled for aLIGO at LLO.

The procedure document used to perform this test is:

-   E1000309 -V3 - aLIGO HAM-ISI, Pre-Integration Testing Procedure, Phase I (post assembly, before storage)

Other useful information can be found in:

-  E1000300 - HAM-ISI LLO test stand: software and electronic check

# Pre-Assembly Testing

## Step 1: Position Sensors

Note: The back panel reads 0.508V/0.001"

| S/N sensor | S/N board | ADE Gap Standoff(mm) | Location on the Jig | Gap Standoff on Jig(mm/in) | Voltage before zeroing | Voltage after zeroing. Prebake | Voltage after zeroing. Post bake |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 12057 | NR | NR | NR | NR | NR | NR | NR |
| 12026 | NR | NR | NR | NR | NR | NR | NR |
| 12013 | NR | NR | NR | NR | NR | NR | NR |
| 12059 | NR | NR | NR | NR | NR | NR | NR |
| 12056 | NR | NR | NR | NR | NR | NR | NR |
| 12083 | NR | NR | NR | NR | NR | NR | NR |

NR: not recorded

Will be measured for the next units.

Sensors noise spectra measured before baking:

| Figure - H1 and V1 sensor noise | Figure - H2 and V2 sensor noise |
| --- | --- |

Figure  - H3 and V3 sensor noise

Acceptance Criteria:

- Power spectrum magnitudes must be lower than:

  - 9.e-10 m/√Hz at 0.1Hz

  - 6.e-10 m/√Hz at 1Hz

Issues/difficulties/comments regarding this test: Values of sensor gaps and zeroing were not recorded. Waived for this unit.

12013 (used for H2) is noisier than other sensors, still passes.

Test result:	Passed:   X   	Failed:         .

## Step 2: GS13

All the data related to GS-13 post podding testing can be found in the SVN at : SeismicSVN\seismic\Common\Data\aLIGO_GS13_TestData\PostMod_TestResults_PDFs.

E1000058 spreadsheet provides the status of each individual GS-13 at LLO site during aLIGO HAM assembly

Data files in SVN at:

/opt/svncommon/seisvn/seismic/Common/Data/aLIGO_GS13_TestData/PostMod_TestResults_RawASCII

Scripts files for processing and plotting in SVN at:

/opt/svncommon/seisvn/seismic/Common/MatlabTools

- gs13qatest.m

Figures in SVN at:

/opt/svncommon/seisvn/seismic/Common/Data/aLIGO_GS13_TestData/PostMod_TestResults_PDFs

## Step 2.1 – Horizontal GS-13s

Figure  - Huddle testing of Horiz GS-13 804,795, and 792 after aLIGO modifications

## Step 2.2 – Vertical GS-13s

Huddle testing

Figure  - Huddle testing of Vert GS-13 679,683 and 699 after aLIGO modifications

Driven testing

Figure  - Driven Transfer Function of Vert GS-13 679,683 and 699 after aLIGO modifications

Acceptance Criteria:

- GS13 should have been already modified and tested. GS-13 Inspection/Pod Assembly (D047810). Checklist is defined in F090070-v6

- Resonant frequency at 1 Hz (huddle testing)

- No spring resonance on vertical driven tests

Test result:	Passed:   X  	Failed:         .

## Step 3: Actuators

Actuator data can be found at: T0900564. Actuator inventory is made at Section II – Step 1.

| Actuator Serial #: L087 Operator Name: Gordon, Matt Date: 11/22/2009 Time: 11:48 AM Actuator Coil Resistance: 6.28 Ohms, PASS Ambient Temperature: 71.1 F Hi Pot Test Results: 1000 MOhms, PASS X Travel Limit (inches): 0.526 Y Travel Limit (inches): 0.205 Z Travel Limit (inches): 0.508 | Actuator Serial #: L088 Operator Name: Gordon, Matt Date: 11/21/2009 Time: 4:48 PM Actuator Coil Resistance: 6.31 Ohms, PASS Ambient Temperature: 70.0 F Hi Pot Test Results: 1000 MOhms, PASS X Travel Limit (inches): 0.530 Y Travel Limit (inches): 0.206 Z Travel Limit (inches): 0.506 |
| --- | --- |
| Actuator Serial #: L096 Operator Name: Gordon, Matt Date: 11/23/2009 Time: 3:50 PM Actuator Coil Resistance: 6.26 Ohms, PASS Ambient Temperature: 71.1 F Hi Pot Test Results: 1000 MOhms, PASS X Travel Limit (inches): 0.530 Y Travel Limit (inches): 0.206 Z Travel Limit (inches): 0.509 | Actuator Serial #: L134 Operator Name: Gordon, Matt Date: 4/12/2010 Time: 3:25 PM Actuator Coil Resistance: 6.34 Ohms, PASS Ambient Temperature: 73.3 F Hi Pot Test Results: 1000 MOhms, PASS X Travel Limit (inches): 0.536 Y Travel Limit (inches): 0.205 Z Travel Limit (inches): 0.506 |
| Actuator Serial #: L137 Operator Name: Gordon, Matt Date: 4/12/2010 Time: 4:19 PM Actuator Coil Resistance: 6.42 Ohms, PASS Ambient Temperature: 73.3 F Hi Pot Test Results: 1000 MOhms, PASS X Travel Limit (inches): 0.530 Y Travel Limit (inches): 0.206 Z Travel Limit (inches): 0.501 | Actuator Serial #: L145 Operator Name: Gordon, Matt Date: 4/13/2010 Time: 11:44 AM Actuator Coil Resistance: 6.45 Ohms, PASS Ambient Temperature: 73.1 F Hi Pot Test Results: 1000 MOhms, PASS X Travel Limit (inches): 0.526 Y Travel Limit (inches): 0.205 Z Travel Limit (inches): 0.506 |

Table 1 Extract of T0900564-testing at Qinetiq

Acceptance Criteria:

- Actuators were previously tested and results are reported in T900564.

Test result:	Passed:    X  	Failed:         .

# Tests to be performed during assembly

## Step 1: Check torques on all bolts

Acceptance Criteria:

- All bolts should trip the wrench, and start moving immediately after. If any bolts in a pattern move before torque is reached, recheck after all bolts are brought to spec.

Test result:	Passed:   X    	Failed:         .

## Step 2: Check gaps under Support Posts

Acceptance Criteria:

- The test is passed if: a 0.001 inch shim cannot be passed freely through any connection to Stage 0 or between post and gussets. If shim can pass through, loosen all constraining bolts, and then retighten iteratively from the center of the part to the edges. Retest.

Test result:	Passed:   X   	Failed:         .

## Step 3: Pitchfork/Boxwork flatness before Optical Table install

Acceptance Criteria:

- Shim inserted won’t pass between parts.

Test result:	Passed:   X   	Failed:         .

## Step 4: Blade spring profile

| Blade # | Base (") | Tip(") | Flatness (mils) |
| --- | --- | --- | --- |
| 1 | .495 | .491 | +4 |
| 2 | .501 | .489 | +12 |
| 3 | .498 | .490 | +8 |

Table 2 - Blade profile

Acceptance Criteria

- Blades must be flat within 0.020" inches.

Test result:	Passed:   X   	Failed:         .

Issues/difficulties/comments regarding this test:

Passed on first try. No adjustment necessary.

## Step 5: Gap checks on actuators

Acceptance Criteria

- Gaps must be within 0.010” of design (i.e. 0.090” and .070” pass, but 0.095” and 0.065” doesn’t).

Test result:	Passed:   X   	Failed:         .

## Step 6: Check level of Stage 1 optical table

Max angle=(.5/64)/85.59= 91 urad

Acceptance Criteria 100 urad.

Test result:	Passed:   X   	Failed:         .

## Step 7: Mass budget

Figure  - Optical table masses distribution

|  | Mass (kg) |
| --- | --- |
| t1 | 45.00 |
| t2 | 45.00 |
| t3 | 45.00 |
| t4 | 45.00 |
| t5 | 45.00 |
| t6 | 45.00 |
| t7 | 35.00 |
| total | 305.00 |

Figure – Keel Masses and Wall masses location

|  | 00 | 01 | 02 | 03 | 04 | 05 | 06 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 0.6 | 1.1 | 2.2 | 4.5 | 7.9 | 15.6 | 27.2 | lbs | kgs |
| w9 |  |  |  |  |  | 1 | 1 | 42.8 | 19.41 |
| w1 | 1 |  |  |  |  | 1 | 1 | 43.4 | 19.69 |
| w2 |  |  |  |  |  | 1 | 1 | 42.8 | 19.41 |
| w3 |  |  |  |  |  | 1 | 1 | 42.8 | 19.41 |
| w4 |  | 1 |  | 1 |  | 1 | 1 | 48.4 | 21.95 |
| w5 |  | 1 |  |  |  | 1 | 1 | 43.9 | 19.91 |
| w6 | 2 |  |  |  |  | 1 | 1 | 44 | 19.96 |
| w7 |  |  |  |  |  | 1 | 1 | 42.8 | 19.41 |
| w8 |  |  |  |  |  | 1 | 1 | 42.8 | 19.41 |
| Side Masses Total | 3 | 2 | 0 | 1 | 0 | 9 | 9 | 393.7 | 178.58 |

Table – Wall masses distribution

|  | 00 | 01 | 02 | 03 | 04 | 05 | 06 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 0.6 | 1.1 | 2.2 | 4.5 | 7.9 | 15.6 | 27.2 | lbs | kgs |
| k1 |  |  |  |  | 1 |  | 1 | 35.1 | 15.92 |
| k2 | 1 | 1 | 1 |  |  |  | 1 | 31.1 | 14.11 |
| k3 |  |  |  |  | 1 |  | 1 | 35.1 | 15.92 |
| k4 | 1 | 1 | 1 |  |  |  | 1 | 31.1 | 14.11 |
| k5 |  |  |  |  | 1 |  | 1 | 35.1 | 15.92 |
| k6 | 1 | 1 | 1 |  |  |  | 1 | 31.1 | 14.11 |
|  | 3 | 3 | 3 | 0 | 3 | 0 | 6 | 198.6 | 90.08 |

Table – Keel masses distribution

| Lockers | Shim thickness (mil) |
| --- | --- |
| A | 125 |
| B | 125 |
| C | 125 |
| D | 125 |

Table – Shims Thickness

|  | Side | Keel | Top | Total |
| --- | --- | --- | --- | --- |
| Weigh (kg) | 178.58 | 90.08 | 305.00 | 573.66 |

Table - Masses distribution

Acceptance Criteria

The Mass budget must be

- 596.7Kg +/-25Kg (4%)

Test result:	Passed:   X   	Failed:         .

## Step 8: Lockers adjustment

| D.I at Lockers | Dial indicators V | Dial indicators H |
| --- | --- | --- |
| A | TBU | TBU |
| B | TBU | TBU |
| C | TBU | TBU |
| D | TBU | TBU |

Table – Dial indicators read-out

Issues/difficulties encountered during this test :

Acceptance Criteria

- Vertical and horizontal displacement near the lockers must be lower than 2 mils

Test result:	Passed:   X    	Failed:         .

# Tests to be performed after assembly

## Step 1 - Actuators Inventory

| Actuator | Actuator S/N |
| --- | --- |
| H1 | L 145 |
| H2 | L 088 |
| H3 | L 134 |
| V1 | L 096 |
| V2 | L 137 |
| V3 | L 087 |

Table - Actuators' inventory

Issues/difficulties/comments regarding this test:

Test result:	Passed:   X   	Failed:        .

## Step 2 - Sensors Inventory

| Sensor | CPS S/N | ADE board serial # |
| --- | --- | --- |
| H1 | 12057 |  |
| H2 | 12013 |  |
| H3 | 12056 |  |
| V1 | 12026 |  |
| V2 | 12059 |  |
| V3 | 12083 |  |

Table - Capacitive position sensors' inventory

| Geophones GS13 | Serial Number | POD |
| --- | --- | --- |
| H1 | 792/809 | 66/46 |
| H2 | 795 | 92 |
| H3 | 804 | 26 |
| V1 | 699 | 56 |
| V2 | 679 | 2 |
| V3 | 683 | 60 |

Table - Geophones' inventory (cf E1000058)

Issues/difficulties/comments regarding this test:

Vertical GS-13 are staying on this unit, whereas horizontals are gotten off to test Unit #2

Test result:	Passed:   X    	Failed:         .

## Step 3 - Electronics Inventory

Write down serial number of coil driver, Anti aliasing chassis, Anti image chassis and interface chassis used for this test are listed below:

| Hardware | Ligo reference | S/N |
| --- | --- | --- |
| Coil driver | D0902744 | S1000317 |
|  |  | S1000316 |
| Anti Image filter | D070081 | S1000251 |
| Anti aliasing filter | D1000269 | S1000244 |
|  |  | S1000245 |
| Interface chassis | D1000067 | S1000311 |
|  |  | S1000312 |
|  |  | S1000314 |

Table - Inventory electronics

Issues/difficulties/comments regarding this test:

We had to change one Anti aliasing filter prior to start testing on this unit.

Test result:	Passed:   X   	Failed:         .

## Step 4 - Set up sensors gap

|  | 10 Kg masses at each corners | No mass |
| --- | --- | --- |
| Table locked | ADE boxes on | ADE boxes on |
| Sensors | Offset (Mean) | Std deviation | Offset (Mean) | Std deviation |
| H1 | -36.967 | 1.2 | -109.44 | 1.5 |
| H2 | 254.8 | 1.1 | 243.71 | 0.8 |
| H3 | -23.343 | 0.7 | -91.761 | 1.1 |
| V1 | -264.62 | 0.6 | -52.461 | 1.6 |
| V2 | -148.73 | 1.8 | 24.719 | 1.5 |
| V3 | 196.35 | 1.4 | 296.66 | 1.2 |

Table - Set-up sensors gap

Issues/difficulties/comments regarding this test: HAM-ISI – LLO unit#3 uses synchronized satellite boxes

Acceptance criteria:

- All mean values must be lower than 400 cts (a bit less than .0005”).

- All standard deviations below 5 counts.

Test result:	Passed:   X   	Failed:         .

## Step 5 - Measure the Sensor gap

| Sensors | Gap measured on the Jig | Gap measured on the table | % of change | Offset sensors (counts) |
| --- | --- | --- | --- | --- |
| H1 | NR | 0.080" |  | 100 |
| H2 | NR | 0.085" |  | -107 |
| H3 | NR | 0.085" |  | -36 |
| V1 | NR | 0.085" |  | 28 |
| V2 | NR | 0.085" |  | -27 |
| V3 | NR | 0.085" |  | 449 |

Table  - Sensors gap

Issues/difficulties/comments regarding this test: No information of gaps measured on the Jig. Difficult to measure without scratching the target.

Acceptance criteria:

- Change of gaps lower than 2% (reference is the gap measured on the Jig)

Test result:	Passed:     X   	Failed:         .

## Step 6 - Check Sensor gaps after the platform release

|  | Table locked | Table unlocked |
| --- | --- | --- |
| Sensors | Offset (Mean) | Std deviation | Offset (Mean) | Difference |
| H1 | -274.44 | 0.83234 | 664.2 | 938.64 |
| H2 | -43.197 | 0.75333 | -327.25 | -284.053 |
| H3 | 159.76 | 0.74358 | -561.14 | -720.9 |
| V1 | -296.64 | 1.0669 | 583.93 | 880.57 |
| V2 | 245.48 | 1.4453 | 495.1 | 249.62 |
| `V3 | -362.15 | 1.627 | -1144.2 | -782.05 |

Table – Sensor gaps after platform release

Issues/difficulties/comments regarding this test:

Acceptance criteria:

- Absolute values of the difference between the unlocked and the locked table must be below:

  - 1600 cts for horizontal sensors (~0.002”)

  - 1600 cts for vertical sensors (~0.002”)

- Considering the acceptance criteria of step 4, all mean values must be lower than

  - 2000 cts for horizontal sensors (~0.0025”)

  - 2000 cts for vertical sensors (~0.0025”)

Test result:	Passed:   X   	Failed:         .

## Step 7 - Check range of motion (hand pushing)

## Step 7.1 – Test Nº1

|  | CPS read out | Calculated after calibration |
| --- | --- | --- |
| Sensors | UP (Counts) | Down (Counts) | UP (mil) | Down (mil) |
| V1 | 20269 | -20331 | 24.0 | -24.1 |
| V2 | 20234 | -20363 | 23.9 | -24.1 |
| V3 | 19885 | -19746 | 23.5 | -23.4 |
|  | CPS read out | Calculated after calibration |
| Sensors | CW(-RZ) | CCW (+RZ) | CW (mil) | CCW (mil) |
| H1 | 18281 | -22331 | 21.6 | -26.4 |
| H2 | 24413 | -19937 | 28.9 | -23.6 |
| H3 | 18099 | -22126 | 21.4 | -26.2 |

Table - Optic table range of motion

Issues/difficulties/comments regarding this test: the results indicate that all displacement sensor signals have the same sign all the way through the read chain.

## Step 7.2 – Test Nº2

Displacement sensors and actuator gap check:

|  | Push in positive direction | Push in negative direction | Railing | Actuator Gap Check |
| --- | --- | --- | --- | --- |
| H1 | 20691 | -26431 |  | X |
| H2 | 24701 | -24100 |  | X |
| H3 | 25028 | -22929 |  | X |
| V1 | 19785 | -20422 |  | X |
| V2 | 31629 | -32519 |  | X |
| V3 | 19762 | -21787 |  | X |

Table - Sensor and Actuator gap check

Acceptance criteria:

- The vertical sensor readout be positive when the optic table is pushed in the +Z direction

- The horizontal sensor readout be negative when the optic table is pushed in the +RZ direction

- Step 7.1

  - Absolutes value of all estimated motions must be higher than 16000counts (~0.020”)

- Step 7.2

  - No contact point on sensors

  - Absolute value of sensor read out must be higher than 16000counts (~0.020”)

  - No contact point on actuators

Test result:	Passed:   X   	Failed:         .

## Step 8 - Capacitive position sensor Power Spectrum

Data files in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Powerspectra/Undamped

- LLO_HAM_ISI_Unit_3_Calibrated_PSD_CPS_GS13_Unlocked_Locked_2011-03-25.mat

Scripts files for processing and plotting in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Powerspectra/Undamped

- Powerspectra_Measurement_HAM_ISI_Locked_Unlocked.m

Figures in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Figures/Powerspectra/Undamped

LLO_HAM_Unit_3_Calibrated_PSD_CPS_Unlocked_Locked_2011-03-25.pdf

CPS calibration:

The CPS power spectrums are calibrated by using a sensitivity of 30.2 nm/count.

Figure - Calibrated CPS power spectrum

(+ signs are Horizontal requirements, x signs are Vertical requirements)

Acceptance criteria:

- No cross talk (peaks at low frequencies + harmonics on measurements)

- Magnitude of power spectrums must be lower than

|  | Locked | Unlocked |
| --- | --- | --- |
|  | at 0.1Hz | at 1Hz | at 0.1Hz | at 1Hz |
| Horizontal | 1.00E-07 | 2.00E-08 | 5.00E-07 | 8.00E-08 |
| Vertical | 4.00E-07 | 5.00E-08 | 8.00E-07 | 4.00E-07 |

Test result:	Passed:   X   	Failed:          .

## Step 9 - GS13 power spectrums

## Step 9.1 –Table locked and unlocked (free)

Figure – Power spectrum Calibrated GS13

## Step 9.2 – Table tilted

The figure below presents the GS13 power spectrum when the table is unlocked and loaded with a 20Kg mass at one of its corner.

Figure – Power spectrum Calibrated GS13 with mass at corner

Data files in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/ Powerspectra/Undamped

- LLO_HAM_ISI_Unit_3_Calibrated_PSD_CPS_GS13_Unlocked_Locked_2011-03-25.mat

- LLO_HAM_ISI_Unit_3_Calibrated_PSD_GS13_Table_Tilted_2011-03-29.mat

Scripts files for taking and processing the data, and plotting it in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Scripts/Data_collection/Old_Scripts

- Powerspectra_Measurement_HAM_ISI_Locked_Unlocked.m

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Scripts/Data_collection/

- Powerspectra_Measurements_Tilted_HAM_ISI.m

Figures in SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Figures/ Powerspectra/Undamped

- LLO_HAM_Unit_3_Calibrated_PSD_GS13_Unlocked_Locked_2011-03-25.pdf

- LLO_HAM_ISI_Unit_3_Calibrated_PSD_GS13_Table_Tilted_2011-03-29.pdf

Acceptance criteria:

- With table locked, magnitudes of power spectra must be lower than:

|  | Table locked |
| --- | --- |
|  | at 0.1Hz | at 1Hz | at 10Hz |
| H & V Geophones | 8.E-05 | 3.E-08 | 2.E-10 |

- With table unlocked, magnitudes of power spectra must be lower than:

|  | Table unlocked |
| --- | --- |
|  | at 0.1Hz | at 1Hz | at 10Hz |
| Horizontal Geophones | 3.E-04 | 1.E-07 | 1.E-11 |
| Vertical Geophones | 5.E-05 | 1.E-07 | 1.E-11 |

- With table unlocked and tilted, magnitudes of power spectra must be lower than:

|  | Table unlocked (tilted with masses) |
| --- | --- |
|  | at 0.1Hz | at 1Hz | at 10Hz |
| H & V Geophones | 8.E-05 | 3.E-08 | 2.E-10 |

Test result:	Passed:   X   	Failed:       .

## Step 10 - Coil Driver, cabling and resistance check

| Actuator | V1 | H1 | V2 |
| --- | --- | --- | --- |
| Coil driver | S1000317 - Coarse 1 | S1000317 - Fine 1 | S1000317 - Fine 2 |
| Anti image pin # |  |  |  |
| Cable # | 28 | 25 | 30 |
| Resistance (Ohm) | P1 - P2 | P2 - P3 | P1 - P2 | P2 - P3 | P1 - P2 | P2 - P3 |
|  | 6.3 | O.L (infinity) | 6.5 | O.L (infinity) | 6.4 | O.L (infinity) |
| MEDM offset (1000 counts) | Measurement P2 (-) ; P1&P3 (+) | Measurement P2 (-) ; P1&P3 (+) | Measurement P2 (-) ; P1&P3 (+) |
|  | 0.3 | 0.303 | 0.3 |
| Actuator | H2 | V3 | H3 |
| Coil driver | S1000317 - Coarse 2 | S1000316 - Coarse 1 | S1000316 - Fine 1 |
| Anti image pin # |  |  |  |
| Cable # | 29 | 26 | 27 |
| Resistance (Ohm) | P1 - P2 | P2 - P3 | P1 - P2 | P2 - P3 | P1 - P2 | P2 - P3 |
|  | 6.4 | O.L (infinity) | 6.4 | O.L (infinity) | 6.3 | O.L (infinity) |
| MEDM offset (1000 counts) | Measurement P2 (-) ; P1&P3 (+) | Measurement P2 (-) ; P1&P3 (+) | Measurement P2 (-) ; P1&P3 (+) |
|  | 0.303 | 0.296 | 0.298 |

Table - Actuators resistance check

Issues/difficulties/comments regarding this test:

Acceptance criteria:

- The measured resistance between the middle pin and one side pin must be 6.5 +/-1 ohms

- Actuator neutral pins must be connected on pin #1 (left side pin of the plug)

- Actuator drive pins must be connected on pin #2 (middle pin of the plug)

- Actuator ground shield pins must be connected on pin #3 (right pin of the plug)

- All LEDs on the coil driver front panel must be green

Test result:	Passed:   X   	Failed:         .

## Step 11 - Actuators Sign and range of motion (Local drive)

## Step 11.1 - Actuators sign

Acceptance criteria: A positive offset drive on one actuator must give positive sensor readout on the collocated sensor. Signs will also be tested when measuring local to local transfer functions.

Test result:	Passed:   X    	Failed:         .

## Step 11.2 - Range of motion - Local drive

|  | Positive and negative offset Drive(+/-30000 counts) |
| --- | --- |
|  | V1 | V2 | V3 | H1 | H2 | H3 |
| V1 readout (counts) | -24840 |  |  | X | X | X |
|  | 23704 |  |  |  |  |  |
| V2 readout (counts) |  | -23504 |  | X | X | X |
|  |  | 24473 |  |  |  |  |
| V3 readout (counts) |  |  | -25079 | X | X | X |
|  |  |  | 24232 |  |  |  |
| H1 readout (counts) | X | X | X | -19988 |  |  |
|  |  |  |  | 19535 |  |  |
| H2 readout (counts) | X | X | X |  | -25296 |  |
|  |  |  |  |  | 27191 |  |
| H3 readout (counts) | X | X | X |  |  | -22424 |
|  |  |  |  |  |  | 21599 |

Table - Range of motion - Local drive

Data files in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Static_Tests

- LLO_HAM-ISI_Unit_3_Range_Of_Motion_20110329.mat

Scripts files for taking and processing the data, and plotting it in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Scripts/Data_collection

- Range_Motion_HAM_ISI.m

Issues/difficulties/comments regarding this test:

Acceptance criteria:

- Main couplings read out must be at least +/-16000 counts (~0.020”)

- Signs of actuators drive and sensors read out have to be the same

Test result:	Passed:   X   	Failed:         .

## Step 12 - Vertical Sensor Calibration

| Lockers | D.I readout for a negative drive | D.I readout without any drive | D.I readout for positive drive |  |
| --- | --- | --- | --- | --- |
| A | 19 | 0 | -19 |  |
| B | 19.25 | 0 | -19 |  |
| C | 18.5 | 0 | -19 |  |
| D | 19 | 0 | -19 |  |
| Average | 18.9375 | 0 | -19 | -37.9 |
| Sensors | Counts | Counts | Counts | Difference (Counts) |
| V1 | -15407 | 703 | 16730 | -32137 |
| V2 | -15595 | 551 | 16437 | -32032 |
| V3 | -17675 | -1772 | 14357 | -32032 |
|  |  |  | Average | -32067 |

Table - Calibration of capacitive position sensors

Vertical sensitivity:	32067/37.91 = 845.3 count/mil

or	845.3 count/mil * 1/1638 V/count = 0.516V/mil

or	25400nm/mil * 1/845.3 mil/count = 30.05 nm/count

Difference with Nominal sensitivity = (845.3-840)/840=-0.63%

Issues/difficulties/comments regarding this test:

Acceptance criteria: Deviation from nominal value < 2%. Nominal value is 840 count/mil.

Test result:	Passed:   X   	Failed:          .

## Step 13 - Vertical Spring Constant

| Sensors | Mean diff counts | Mean diff m | K (N/m) |
| --- | --- | --- | --- |
| V1 | 7370 | 2.23E-04 | 8.81E+04 |
| V2 | 7538 | 2.28E-04 | 8.62E+04 |
| V3 | 8553 | 2.58E-04 | 7.60E+04 |
|  |  | Average (N/m) | 2.50E+05 |

Table - Vertical spring constant

The measured vertical stiffness is 2.503e5N/m. The nominal vertical spring constant is 2.428e5N/m. The measured error on the vertical stiffness is 3.10%.

Issues/difficulties/comments regarding this test:

Acceptance criteria:

- Spring constant is within +/- 10% of 2.428e5 N/m (HPD FEA Results).

Test result:	Passed:   X   	Failed:         .

## Step 14 - Static Testing (Tests in the local basis)

|  |  | H1 | H2 | H3 | V1 | V2 | V3 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Actuators (1000 counts) | H1 | 1967.669 | 1210.834 | 1224.792 | -3.392 | 12.744 | -33.92 |
|  | H2 | 1207.9290 | 2017.293 | 1256.419 | 11.939 | 21.785 | -21.05 |
|  | H3 | 1224.299 | 1259.544 | 2015.90 | 37.030 | 16.072 | -30.849 |
|  | V1 | 201.7530 | 172.320 | -313.735 | 1415.569 | -43.636 | -546.33 |
|  | V2 | -276.287 | 261.940 | 230.952 | -554.369 | 1437.609 | -44.9600 |
|  | V3 | 159.632 | -385.887 | 142.389 | 44.354 | -612.036 | 1403.714 |

Table - Main and cross coupling

Data files in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Static_Tests

- LLO_HAM_ISI_Unit_3_Sensor_Readout_Local_20110331.mat

Scripts files for taking data in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Scripts/Data_Collection

- Sensor_Readout_Local_Drive_MEDM_HAM_ISI.m

Issues/difficulties/comments regarding this test:

Acceptance criteria:

- Vertical

For a +1000 count offset drive on vertical actuators

  - Collocated sensors must be 1400 counts +/- 10%

- Horizontal

For a +1000 count offset drive on horizontal actuators

  - Collocated sensors must be 2000 counts +/- 10%

  - Non-collocated horizontal sensors must be 1250 counts +/-10%

Test result:	Passed:   X    	Failed:         .

## Step 15 - Linearity test

|  | Slope | Offset | Average slope | Variation from average(%) |
| --- | --- | --- | --- | --- |
| H1 | 2.077 | 127.6 | 2.0845 | -0.37 |
| H2 | 2.102 | -763 |  | 0.83 |
| H3 | 2.075 | -711 |  | -0.46 |
| V1 | 1.498 | -282 | 1.4872 | 0.74 |
| V2 | 1.479 | 368.3 |  | -0.53 |
| V3 | 1.484 | -1024 |  | -0.20 |

Table - Slopes and offset of the triplet Actuators - HAM-ISI - Sensors

Scripts files for taking data in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Scripts/Data_Collection

- Linearity_Test_Awgstream_HAM_ISI.m

Figure - Horizontal and vertical actuators x HAM-ISI x sensors

Data files in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Linearity_test

- LLO_HAM_ISI_Unit_3_Linearity_test_20110407.mat

Figures in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Figures/Linearity_test

- LLO_HAM_ISI_Unit_3_Linearity_test_20110407.fig

LLO_HAM_ISI_Unit_3_Linearity_test_20110407.pdf

Issues/difficulties/comments regarding this test:

Acceptance criteria:

- Horizontal and vertical slopes of the triplet actuators x HAM-ISI x sensors:  Average slope +/- 3%.

Test result:	Passed:   X   	Failed:         .

## Step 16 – Static tests in the general coordinate basis

|  | 1000 counts Drive | X Drive | Y Drive | Z Drive | Rx Drive | Ry Drive | Rz Drive |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Sensors readout (count) | H1 | 257 | -434 | 11 | -433 | -236 | -1970 |
|  | H2 | 260 | 469 | 15 | 497 | -246 | -1971 |
|  | H3 | -513 | 13 | 17 | 22 | 543 | -1980 |
|  | V1 | -41 | -53 | 191 | -565 | -1710 | -98 |
|  | V2 | -39 | -25 | 230 | 1634 | 334 | -71 |
|  | V3 | -49 | -53 | 222 | -1315 | 1178 | -89 |
|  | Direction read out | 514 | 512 | 236 | 2631 | 2518 | 2547 |

Table - Tests in the general coordinate basis

Issues/difficulties/comments regarding this test:

Acceptance criteria:

For a positive drive in the Cartesian basis:

- Local sensor readout must have the same sign that the reference table (CONT2ACT check)

- Cartesian sensors read out must be positive (DISP2CEN check) in the drive direction

Test result:	Passed:   X   	Failed:         .

## Step 17 - Frequency response

Compensation filters of the new GS13 interface chassis are located in the geophone pre-filters bank. Powerspectra were measured with masses on the optic table not bolted.

## Step 17.1 - Local to local measurements

Local to local transfer functions have been measured with 90 repetitions.

Data files in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_Functions/Measurements/

Undamped/

- LLO_HAM_ISI_Unit_3_Data_TF_L2L_50mHz_500mHz_20110406-181629.mat

- LLO_HAM_ISI_Unit_3_Data_TF_L2L_500mHz_5Hz_20110406-151108.mat

- LLO_HAM_ISI_Unit_3_Data_TF_L2L_200Hz_800Hz_20110406-120426.mat

- LLO_HAM_ISI_Unit_3_Data_TF_L2L_5Hz_200Hz_20110406-133747.mat

Data collection script files:

opt/svncommon/seisvn/seismic/HAM-ISI/X2/Scripts/Data_Collection

- Run_TF_L2L_50mHz_500mHz.m

- Run_TF_L2L_500mHz_5Hz.m

- Run_TF_L2L_5mHz_200Hz.m

- Run_TF_L2L_200Hz_800Hz.m

Scripts files for processing and plotting in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_Functions/Measurements/

Undamped/

- Plot_LLO_HAM_ISI_Unit_3_TF_L2L_2011_04_06.m

Figures in SVN at:

opt/svncommon/seisvn/seismic/HAMISI/X2/Data/Unit_3/Figures/Transfer_Functions/Measurements/

Undamped/

- LLO_HAM_ISI_Unit_3_TF_L2L_H_CPS_50mHz_800Hz_2011_04_06.fig

- LLO_HAM_ISI_Unit_3_TF_L2L_V_CPS_50mHz_800Hz_2011_04_06.fig

- LLO_HAM_ISI_Unit_3_TF_L2L_H_GS13_50mHz_800Hz_2011_04_06.fig

- LLO_HAM_ISI_Unit_3_TF_L2L_V_GS13_50mHz_800Hz_2011_04_06.fig

Storage of measured transfer functions in the SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_functions/ Measurements/

Undamped/

- LLO_HAM_ISI_Unit_3_Data_TF_L2L_2011_04_06.mat

The local to local transfer functions are presented below.

Figure - Local to Local Measurements – Horizontal capacitive sensors

Figure - Local to Local Measurements – Horizontal inertial sensors

Figure - Local to Local Measurements – Vertical capacitive sensors

Figure - Local to Local Measurements – Vertical inertial sensors

Issues/difficulties/comments regarding this test:

Cables contacting Stage 1 and Stage 0 created issues in the .5 to 5 Hz.

At low frequencies, the horizontal GS-13s present some differences.

## Step 17.2 - Cartesian to Cartesian measurements

Cartesian to Cartesian transfer functions have been measured with 90 repetitions.

Data files in SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_Functions/Measurements/

Undamped/

- LLO_HAM_ISI_Unit_3_Data_TF_C2C_50mHz_500mHz_20110405-215335.mat

- LLO_HAM_ISI_Unit_3_Data_TF_C2C_500mHz_5Hz_20110405-184814.mat

- LLO_HAM_ISI_Unit_3_Data_TF_C2C_5Hz_200Hz_20110405-171453.mat

- LLO_HAM_ISI_Unit_3_Data_TF_C2C_200Hz_800Hz_20110405-154132.mat

Scripts files for processing and plotting in SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_Functions/Measurements/

Undamped/

- Plot_LLO_HAM_ISI_Unit_3_TF_C2C_2011_04_06.m

Figures in SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Figures/Transfer_Functions/

Measurements/Undamped/

- LLO_HAM_ISI_Unit_3_TF_C2C_X_Y_RZ_CPS_50mHz_800Hz_2011_04_05.fig

- LLO_HAM_ISI_Unit_3_TF_C2C_X_Y_RZ_GS13_50mHz_800Hz_2011_04_05.fig

- LLO_HAM_ISI_Unit_3_TF_C2C_Z_RX_RY_CPS_50mHz_800Hz_2011_04_05.fig

- LLO_HAM_ISI_Unit_3_TF_C2C_Z_RX_RY_GS13_50mHz_800Hz_2011_04_05.fig

Storage of measured transfer functions in the SVN at:

/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_functions/Cartesian_to_Cartesian

- LLO_HAM_ISI_Unit_3_Data_TF_C2C_2011_04_05

Figure - Cartesian to Cartesian CPS measurements – X, Y, RZ directions

Figure - Cartesian to Cartesian GS-13 measurements – X, Y, RZ directions

Figure - Cartesian to Cartesian measurements – Z, RX, RY directions

Issues/difficulties/comments regarding this test:

Acceptance criteria:

- Local to local measurement

  - On CPS, the phase must be 0º at DC

  - On Geophones, the phase must be -90º at DC

- Cartesian to cartesian measurement

  - On CPS, the phase must be 0º at DC

  - On Geophones, the phase must be -90º at DC

Test result:	Passed:   X   	Failed:         .

## Step 18 - Transfer function comparison

## Step 18.1 - Local to local - Comparison with HAM6

Scripts files for processing and plotting in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_Functions/Measurements/

Undamped/

- Plot_LLO_HAM_ISI_Unit_3_TF_L2L_2011_04_06.m

Local to local figures in SVN at:

/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Figures/Transfer_Functions/Measurements/

- LLO_HAM_ISI_Unit_3_TF_L2L_H_CPS_50mHz_800Hz_wHAM6_2011_04_06.fig

- LLO_HAM_ISI_Unit_3_TF_L2L_V_CPS_50mHz_800Hz_wHAM6_2011_04_06.fig

- LLO_HAM_ISI_Unit_3_TF_L2L_H_GS13_50mHz_800Hz_wHAM6_2011_04_06.fig

- LLO_HAM_ISI_Unit_3_TF_L2L_V_GS13_50mHz_800Hz_wHAM6_2011_04_06.fig

GS13, Local to local measurement

Figure - Local to local measurements comparison with HAM6 – Horizontal GS13

Figure - Local to local measurements comparison with HAM6 – Vertical GS13

CPS, Local to local measurement, Undamped

Figure - Local to local measurements comparison – Horizontal Position sensors

Figure - Local to local measurements comparison – Horizontal Position sensors

Difference with HAM6:

- Sign difference on CPS and Geophones.

## Step 18.2 - Cartesian to Cartesian - Comparison with HAM6

Scripts files for processing and plotting in SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_Functions/Measurements/

Undamped/

- Plot_LLO_HAM_ISI_Unit_3_TF_C2C_2011_04_06.m

Cartesian to Cartesian figures in SVN at :

/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_functions/Cartesian_to_Cartesian

- LLO_HAM_ISI_Unit_3_TF_C2C_X_Y_RZ_CPS_50mHz_800Hz_wHAM6_2011_04_05.fig

- LLO_HAM_ISI_Unit_3_TF_C2C_Z_RX_RY_CPS_50mHz_800Hz_wHAM6_2011_04_05.fig

- LLO_HAM_ISI_Unit_3_TF_C2C_X_Y_RZ_GS13_50mHz_800Hz_wHAM6_2011_04_05.fig

- LLO_HAM_ISI_Unit_3_TF_C2C_Z_RX_RY_GS13_50mHz_800Hz_wHAM6_2011_04_05.fig

GS13, Cartesian to Cartesian measurement, Undamped

Figure - Cartesian to Cartesian measurements comparison with HAM6 – Horiz GS13

Figure - Cartesian to Cartesian measurements comparison with HAM6 – Vertical GS13

CPS, Cartesian to Cartesian measurement, Undamped

Figure - Cartesian to Cartesian measurements comparison – Horizontal Position sensors

Figure - Cartesian to Cartesian measurements comparison – Vertical Position sensors

Difference with HAM6:

- Resonance at 1 Hz is slightly higher than HAM6

- Sign difference on Geophones

Acceptance criteria:

- No difference with the reference transfer functions (SVN)

  - Phase – less than 10º - In Phase – Out of Phase

  - Damping (fit by eye with the reference transfer functions)

  - DC gain

  - Eigen frequencies shift less than 5%

Test result:	Passed:    X   	Failed:         .

## Step 19 - Lower Zero Moment Plane

Data collection script files:

opt/svncommon/seisvn/seismic/HAM-ISI/X2/Scripts/Data_Collection

- Run_Cart2Cart_10mHz_100mHz.m

Data files in SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_Functions/Measurements/

Undamped/

- LZMP_LLO_HAM-ISI-Unit_3_2011_04_06.mat

Scripts files for processing and plotting in SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_Functions/Measurements/

Undamped/

- LZMP_2011_04_06.m

Figures in SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Figures/Transfer_Functions/

Measurements/Undamped/

- LZMP.fig

X & Y offsets:

| X offset (mm) | 0.399 |
| --- | --- |
| Y offset (mm) | 0.738 |

Table – Offset of the Lower Zero Moment Plane

The results from two measurements are presented on the figure below:

Figure - Lower Zero moment plane – Main and cross couplings at low frequency

Acceptance criteria:

- X offset must be less than 1 mm

- Y offset must be less than 1mm

Test result:	Passed:    X   	Failed:         .

## Step 20 - Damping loops

Filters used by Damping loops in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X1/FilterDesign/Unit_3/txt_file

- G1ISIHAM_Back_up_2010_09_01.txt (digitalized filters copied and rename to G1ISIHAM.txt in /opt/rtcds/geo/g1/chans)

## Step 20.1 - Transfer functions - Simulation

Scripts files for processing and plotting in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_Functions/Simulation/Damping

- HAM_ISI_LLO_Unit_3_Damping_TF_2011_04_07.m

Figures in SVN at:

/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/ Transfer_Functions/Simulations/

Damping/

- Damping_LOOP_H1_H2_H3.fig

- Damping_LOOP_V1_V2_V3.fig

Results are saved in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_Functions/Simulations/

Damping/

The following figures present the plant, controller, open loop, closed loop and sensitivity of vertical and horizontal damping loops. H1/V1 are plotted in solid line, H2/V2 are plotted in dash line, H3/V3 are plotted in dash-dot line.

## Step 20.2 - Powerspectrum – Experimental

Data files in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Powerspectra/Damping/

Scripts files for taking data and plotting in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Scripts/Data_Collection/

- Powerspectra_Measurements_Undamped_Damped_HAM_ISI.m

Figures in SVN at:

/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Figures/Powerspectra/Damping/

- LLO_HAM_ISI_Unit_3_Calibrated_PSD_CPS_Undamped_Damped_2011_04_08.fig

Powerspectrum:

The figure below shows power spectrum of Geophones when there when there is no damping loop and when all damping loops are engaged.

Figure – Horizontal (left) and vertical (right) damping loops - Experimental

Acceptance criteria:

- All Ham 6 damping loops must implemented and stable

Test result:	Passed:   X   	Failed:         .