LIGO Laboratory / LIGO Scientific Collaboration

LIGO- E1000327 	                 LIGO                   July 20th, 2011

aLIGO HAM-ISI,

LLO Assembly Unit# 3 (HAM3 Chamber) Testing Report,

Phase I: Assembly validation

E1000327 – V7

Joe Hanson, Celine Ramet, Michael Vargas, Adrien Le Roux

Distribution of this document:

Advanced LIGO Project

This is an internal working note

of the LIGO Laboratory

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 1970 Mail Stop S9-02 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

Table of contents:

Introduction	4

I.	Pre-Assembly Testing	5

	Step 1: Position Sensors	5

	Step 2: GS13	7

	Step 2.1 – Horizontal GS-13s	7

	Step 2.2 – Vertical GS-13s	9

	Step 3: Actuators	11

II.	Tests to be performed during assembly	12

	Step 1: Inventory (E1000052)	12

	Step 2: Check torques on all bolts	13

	Step 3: Check gaps under Support Posts	13

	Step 4: Pitchfork/Boxwork flatness before Optical Table install	13

	Step 5: Blade spring profile	13

	Step 6: Gap checks on actuators-after installation on Stage 1	14

	Step 7: Check level of Stage 0	14

	Step 8: Check level of Stage 1 Optical Table	15

	Step 9: Mass budget	16

	Step 10: Shim thickness	18

	Step 11: Lockers adjustment	18

III.	Tests to be performed after assembly	19

	Step 1 - Electronics Inventory	19

	Step 2 - Set up sensors gap	19

	Step 3 - Measure the Sensor gap	20

	Step 4 - Check Sensor gaps after the platform release	20

	Step 5 – Performance of the limiter	21

	Step 5.1 - Test Nº1 - Push “in the general coordinates”	21

	Step 5.2 - Test Nº2 – Push “locally”	21

	Step 6 - Position Sensors unlocked/locked Power Spectrum	22

	Step 7 - GS13 power spectrum -tabled tilted	24

	Step 8- GS13 pressure readout	26

	Step 9 - Coil Driver, cabling and resistance check	27

	Step 10 - Actuators Sign and range of motion (Local drive)	28

	Step 11 - Vertical Sensor Calibration	29

	Step 12 - Vertical Spring Constant	30

	Step 13 - Static Testing (Tests in the local basis)	30

	Step 14 - Linearity test	32

	Step 15 - Cartesian Basis Static Testing	34

	Step 16- Frequency response	35

	Step 16.1 - Local to local measurements	35

	Step 16.2 - Cartesian to Cartesian measurements	38

	Step 17 - Transfer function comparison with Reference	41

o	Step 17.1 - Local to local - Comparison with Reference	41

	Step 17.2 - Cartesian to Cartesian - Comparison with Reference	43

	Step 18 - Lower Zero Moment Plane	47

	Step 19 - Damping loops	49

	Step 19.1 - Transfer functions - Simulation	49

	Step 19.2 - Powerspectra – Experimental	51

Conclusion	1

Introduction

This document presents the tests performed to characterize and validate the “HAM-ISI LLO Unit #3”. This unit was the 3rd unit assembled for aLIGO at LLO. This unit was partially assembled in Fall 2010, but following the discovery of unauthorized repairs in the parts, the assembly was interrupted. All parts in questions were disassembled. This unit was the 1st to be re-assembled when the assembly started back in Spring 2011.

Following that testing, the testing procedure was modified (going to v7), that latest version of the test report (v5)was created afterwards, some data being re-processed. In other words, v4 was still using the older requirements.

The procedure document used to perform this test is:

-   E1000309 –V7 - aLIGO HAM-ISI, Pre-Integration Testing Procedure, Phase I (post assembly, before storage)

Other useful information can be found in:

-  E1000300 - HAM-ISI LLO test stand: software and electronic check

# Pre-Assembly Testing

## Step 1: Position Sensors

Note: The back panel reads 0.508V/0.001"

| S/N sensor | S/N board | ADE Gap Standoff(mm) | Location on the Jig | Gap Standoff on Jig(mm/in) | Voltage before zeroing | Voltage after zeroing. Prebake | Voltage after zeroing. Post bake |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 12057 | NR | NR | NR | NR | NR | NR | NR |
| 12026 | NR | NR | NR | NR | NR | NR | NR |
| 12013 | NR | NR | NR | NR | NR | NR | NR |
| 12059 | NR | NR | NR | NR | NR | NR | NR |
| 12056 | NR | NR | NR | NR | NR | NR | NR |
| 12083 | NR | NR | NR | NR | NR | NR | NR |

NR: not recorded

Will be measured for the next units.

Sensors noise spectra measured before baking:

| Figure - H1 and V1 sensor noise | Figure - H2 and V2 sensor noise |
| --- | --- |

Figure  - H3 and V3 sensor noise

Acceptance Criteria:

- Power spectrum magnitudes must be lower than:

  - 9.e-10 m/√Hz at 0.1Hz

  - 6.e-10 m/√Hz at 1Hz

Issues/difficulties/comments regarding this test: Values of sensor gaps and zeroing were not recorded. Waived for this unit.

12013 (used for H2) is noisier than other sensors, still passes.

Test result:	Passed:   X   	Failed:         .

## Step 2: GS13

All the data related to GS-13 post podding testing can be found in the SVN at : SeismicSVN\seismic\Common\Data\aLIGO_GS13_TestData\PostMod_TestResults_PDFs.

E1000058 spreadsheet provides the status of each individual GS-13 at LLO site during aLIGO HAM assembly

Data files in SVN at:

/opt/svncommon/seisvn/seismic/Common/Data/aLIGO_GS13_TestData/PostMod_TestResults_RawASCII

Scripts files for processing and plotting in SVN at:

/opt/svncommon/seisvn/seismic/Common/MatlabTools

- gs13qatest.m

Figures in SVN at:

/opt/svncommon/seisvn/seismic/Common/Data/aLIGO_GS13_TestData/PostMod_TestResults_PDFs

## Step 2.1 – Horizontal GS-13s

Huddle testing

Figure  - Huddle testing of Horiz GS-13 804,795, and 792 after aLIGO modifications

## Step 2.2 – Vertical GS-13s

Huddle testing

Figure  - Huddle testing of Vert GS-13 679,683 and 699 after aLIGO modifications

Driven testing

Figure  - Driven Transfer Function of Vert GS-13 679,683 and 699 after aLIGO modifications

Acceptance Criteria:

- GS13 should have been already modified and tested. GS-13 Inspection/Pod Assembly (D047810). Checklist is defined in F090070-v6

- Resonant frequency at 1 Hz (huddle testing)

- No spring resonance on vertical driven tests

Test result:	Passed:   X  	Failed:         .

## Step 3: Actuators

Actuator data can be found at: T0900564. Actuator inventory is made at Section II – Step 1.

| Actuator Serial #: L087 Operator Name: Gordon, Matt Date: 11/22/2009 Time: 11:48 AM Actuator Coil Resistance: 6.28 Ohms, PASS Ambient Temperature: 71.1 F Hi Pot Test Results: 1000 MOhms, PASS X Travel Limit (inches): 0.526 Y Travel Limit (inches): 0.205 Z Travel Limit (inches): 0.508 | Actuator Serial #: L088 Operator Name: Gordon, Matt Date: 11/21/2009 Time: 4:48 PM Actuator Coil Resistance: 6.31 Ohms, PASS Ambient Temperature: 70.0 F Hi Pot Test Results: 1000 MOhms, PASS X Travel Limit (inches): 0.530 Y Travel Limit (inches): 0.206 Z Travel Limit (inches): 0.506 |
| --- | --- |
| Actuator Serial #: L096 Operator Name: Gordon, Matt Date: 11/23/2009 Time: 3:50 PM Actuator Coil Resistance: 6.26 Ohms, PASS Ambient Temperature: 71.1 F Hi Pot Test Results: 1000 MOhms, PASS X Travel Limit (inches): 0.530 Y Travel Limit (inches): 0.206 Z Travel Limit (inches): 0.509 | Actuator Serial #: L134 Operator Name: Gordon, Matt Date: 4/12/2010 Time: 3:25 PM Actuator Coil Resistance: 6.34 Ohms, PASS Ambient Temperature: 73.3 F Hi Pot Test Results: 1000 MOhms, PASS X Travel Limit (inches): 0.536 Y Travel Limit (inches): 0.205 Z Travel Limit (inches): 0.506 |
| Actuator Serial #: L137 Operator Name: Gordon, Matt Date: 4/12/2010 Time: 4:19 PM Actuator Coil Resistance: 6.42 Ohms, PASS Ambient Temperature: 73.3 F Hi Pot Test Results: 1000 MOhms, PASS X Travel Limit (inches): 0.530 Y Travel Limit (inches): 0.206 Z Travel Limit (inches): 0.501 | Actuator Serial #: L145 Operator Name: Gordon, Matt Date: 4/13/2010 Time: 11:44 AM Actuator Coil Resistance: 6.45 Ohms, PASS Ambient Temperature: 73.1 F Hi Pot Test Results: 1000 MOhms, PASS X Travel Limit (inches): 0.526 Y Travel Limit (inches): 0.205 Z Travel Limit (inches): 0.506 |

Acceptance Criteria:

- Actuators were previously tested and results are reported in T900564.

The tests report must contain:

Test results (Passed:       	Failed:        ).””

# Tests to be performed during assembly

## Step 1: Inventory (E1000052)

| aLIGO HAM-ISI Serial Number Registration |
| --- |
| Assembly Site | LLO | Assembly S/N | 003 | Destination | L1-H3 |
| DCC/Vendor number | Part name | Configuration | S/N | S/N | S/N |
| D071001 | Stage 0 base |  | 6 |  |  |
| D071051 | Stage 1 base |  | 5 |  |  |
| D071050 | Optical table |  | 6 |  |  |
| D071002 | Spring Post |  | 42 | 28 | 34 |
| D071100 | Spring |  | 42 | 15 | 14 |
| D071102 | Flexure |  | 14 | 26 | 15 |
| ADE | Position sensor | Horizontal | 12057 | 12013 | 12056 |
|  |  | Vertical | 12026 | 12059 | 12083 |
| D047812 | GS-13 pod | Horizontal | 66/46 | 26 | 92 |
|  |  | Vertical | 56 | 2 | 60 |
| D047823 | L4C pod | Horizontal | N/A | N/A | N/A |
|  |  | Vertical | N/A | N/A | N/A |
| D0902749 | Actuator | Horizontal | 145 | 088 | 134 |
|  |  | Vertical | 096 | 137 | 087 |

## Step 2: Check torques on all bolts

Acceptance Criteria:

- All bolts should trip the wrench, and start moving immediately after. If any bolts in a pattern move before torque is reached, recheck after all bolts are brought to spec.

Test result:	Passed:   X    	Failed:         .

## Step 3: Check gaps under Support Posts

Acceptance Criteria:

- A 0.001 inch shim cannot be passed freely through any connection to Stage 0 or between post and gussets. If shim can pass through, loosen all constraining bolts, and then retighten iteratively from the center of the part to the edges. Retest.

Test result:	Passed:   X   	Failed:         .

## Step 4: Pitchfork/Boxwork flatness before Optical Table install

Acceptance Criteria:

- Shim inserted won’t pass between parts.

Test result:	Passed:   X   	Failed:         .

## Step 5: Blade spring profile

| Blade # | Base (") | Tip(") | Flatness (mils) |
| --- | --- | --- | --- |
| 1 | .495 | .491 | +4 |
| 2 | .501 | .489 | +12 |
| 3 | .498 | .490 | +8 |

Table 2 - Blade profile

Acceptance Criteria:

- Blades must be flat within 0.015" inches.

Test result:	Passed:   X   	Failed:         .

## Step 6: Gap checks on actuators-after installation on Stage 1

| Actuator | Front Gap (1/1000") | Back Gap (1/1000") |
| --- | --- | --- |
| V1 | 80 | 90 |
| V2 | 90 | 75 |
| V3 | 95/85 | 65/80 |
| H1 | 85 |  |
| H2 | 80 |  |
| H3 | 80 |  |

Acceptance Criteria

- Gaps must be within 0.010” of design (i.e. 0.090” and .070” pass, but 0.095” and 0.065” doesn’t).

Test result:	Passed:      	Failed:     X    .

## Step 7: Check level of Stage 0

Not recorded.

Test result:	Passed:      	Failed:     X    .

## Step 8: Check level of Stage 1 Optical Table

Optical Level measurement of Stage 1 at large (12 - 15) number of points.

Max angle=(.5/64)/85.59= 91 urad

Acceptance Criteria

- The maximum angle of the table with the horizontal mustn’t exceed  ~100µrad

Test result:	Passed:   X   	Failed:         .

## Step 9: Mass budget

Figure – Keel Masses and Wall masses location

Figure  - Optical table masses distribution

|  | 00 | 01 | 02 | 03 | 04 | 05 | 06 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 0.6 | 1.1 | 2.2 | 4.5 | 7.9 | 15.6 | 27.2 | lbs | kgs |
| w9 |  |  |  |  |  | 1 | 1 | 42.8 | 19.41 |
| w1 | 1 |  |  |  |  | 1 | 1 | 43.4 | 19.69 |
| w2 |  |  |  |  |  | 1 | 1 | 42.8 | 19.41 |
| w3 |  |  |  |  |  | 1 | 1 | 42.8 | 19.41 |
| w4 |  | 1 |  | 1 |  | 1 | 1 | 48.4 | 21.95 |
| w5 |  | 1 |  |  |  | 1 | 1 | 43.9 | 19.91 |
| w6 | 2 |  |  |  |  | 1 | 1 | 44 | 19.96 |
| w7 |  |  |  |  |  | 1 | 1 | 42.8 | 19.41 |
| w8 |  |  |  |  |  | 1 | 1 | 42.8 | 19.41 |
| Side Masses Total | 3 | 2 | 0 | 1 | 0 | 9 | 9 | 393.7 | 178.58 |

Table – Wall masses distribution

|  | 00 | 01 | 02 | 03 | 04 | 05 | 06 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 0.6 | 1.1 | 2.2 | 4.5 | 7.9 | 15.6 | 27.2 | lbs | kgs |
| k1 |  |  |  |  | 1 |  | 1 | 35.1 | 15.92 |
| k2 | 1 | 1 | 1 |  |  |  | 1 | 31.1 | 14.11 |
| k3 |  |  |  |  | 1 |  | 1 | 35.1 | 15.92 |
| k4 | 1 | 1 | 1 |  |  |  | 1 | 31.1 | 14.11 |
| k5 |  |  |  |  | 1 |  | 1 | 35.1 | 15.92 |
| k6 | 1 | 1 | 1 |  |  |  | 1 | 31.1 | 14.11 |
|  | 3 | 3 | 3 | 0 | 3 | 0 | 6 | 198.6 | 90.08 |

|  | Side | Keel | Top | Total |
| --- | --- | --- | --- | --- |
| Weigh (kg) | 178.58 | 90.08 | 305.00 | 573.66 |
| Torque x at O (N.m) | -13.85 | 0.00 | -8.72 | -22.57 |
| Torque y at O (N.m) | -24.30 | 5.26 | 11.21 | -7.83 |

Table - Masses distribution (computed using T1100261)

Acceptance Criteria

The Mass budget must be

- 579.1 Kg (cf E1100427)+/-25Kg (5%)

Test result:	Passed:   X   	Failed:         .

## Step 10: Shim thickness

| Lockers | Shim thickness (mil) |
| --- | --- |
| A | 125 |
| B | 125 |
| C | 125 |
| D | 125 |

Table – Shims Thickness

Acceptance Criteria

- Inventory is complete

Test result:	Passed:   X   	Failed:         .

## Step 11: Lockers adjustment

| D.I at Lockers | Dial indicators V | Dial indicators H |
| --- | --- | --- |
| A | 0.5 | -1 |
| B | 2 | 0 |
| C | .25 | 0 |
| D | -1.5 | -1.5 |

Table – Dial indicators read-out (in thousands of an inch)

Issues/difficulties encountered during this test : N/A

Acceptance Criteria

- Vertical and horizontal displacement near the lockers must be lower than 2 mils (0.002”)

Test result:	Passed:   X    	Failed:         .

# Tests to be performed after assembly

## Step 1 - Electronics Inventory

| Hardware | LIGO reference | S/N |
| --- | --- | --- |
| Coil driver | D0902744 | S1000317 |
|  |  | S1000316 |
| Anti Image filter | D070081 | S1000251 |
| Anti aliasing filter | D1000269 | S1000244 |
|  |  | S1000245 |
| Interface chassis | D1000067 | S1000311 |
|  |  | S1000312 |
|  |  | S1000314 |

Table - Inventory electronics

Acceptance Criteria

- Inventory is complete

Test result:	Passed:   X   	Failed:         .

## Step 2 - Set up sensors gap

|  | 10 Kg masses at each corners | No mass |
| --- | --- | --- |
| Table locked | ADE boxes on | ADE boxes on |
| Sensors | Offset (Mean) | Std deviation | Offset (Mean) | Std deviation |
| H1 | -36.967 | 1.2 | -109.44 | 1.5 |
| H2 | 254.8 | 1.1 | 243.71 | 0.8 |
| H3 | -23.343 | 0.7 | -91.761 | 1.1 |
| V1 | -264.62 | 0.6 | -52.461 | 1.6 |
| V2 | -148.73 | 1.8 | 24.719 | 1.5 |
| V3 | 196.35 | 1.4 | 296.66 | 1.2 |

Table – Capacitive position sensor readout after gap set-up

Issues/difficulties/comments regarding this test: HAM-ISI – LLO unit#3 uses synchronized satellite boxes

Acceptance criteria:

- All mean values must be lower than 400 cts (a bit less than .0005”).

- All standard deviations below 5 counts.

- No cross talk

Test result:	Passed: X 	Failed:       .

## Step 3 - Measure the Sensor gap

| Sensors | Gap measured on the Jig | Gap measured on the table |
| --- | --- | --- |
| H1 | NR | 0.085" |
| H2 | NR | 0.085" |
| H3 | NR | 0.085" |
| V1 | NR | 0.082" |
| V2 | NR | 0.085" |
| V3 | NR | 0.082" |

Acceptance criteria:

Sensors gap measured on the jig and on the optic table must be:

- 0.080” +/-0.002”

Test result:	Passed:   X  	Failed:         .

## Step 4 - Check Sensor gaps after the platform release

|  | Table locked | Table unlocked |
| --- | --- | --- |
| Sensors | Offset (Mean) | Std deviation | Offset (Mean) | Difference |
| H1 | -274.44 | 0.83234 | 664.2 | 938.64 |
| H2 | -43.197 | 0.75333 | -327.25 | -284.053 |
| H3 | 159.76 | 0.74358 | -561.14 | -720.9 |
| V1 | -296.64 | 1.0669 | 583.93 | 880.57 |
| V2 | 245.48 | 1.4453 | 495.1 | 249.62 |
| `V3 | -362.15 | 1.627 | -1144.2 | -782.05 |

Table – Sensor gaps after platform release

Acceptance criteria:

- Absolute values of the difference between the unlocked and the locked table must be below:

  - 1600 cts for horizontal sensors (~0.002”)

  - 1600 cts for vertical sensors (~0.002”)

- Considering the acceptance criteria of step 4, all mean values must be lower than

  - 2000 cts for horizontal sensors (~0.0025”)

  - 2000 cts for vertical sensors (~0.0025”)

Test result:	Passed:  X    	Failed:         .

## Step 5 – Performance of the limiter

## Step 5.1 - Test Nº1 - Push “in the general coordinates”

|  | CPS read out | Calculated after calibration |
| --- | --- | --- |
| Sensors | UP (Counts) | Down (Counts) | UP (mil) | Down (mil) |
| V1 | 20269 | -20331 | 24.0 | -24.1 |
| V2 | 20234 | -20363 | 23.9 | -24.1 |
| V3 | 19885 | -19746 | 23.5 | -23.4 |
|  | CPS read out | Calculated after calibration |
| Sensors | CW(-RZ) | CCW (+RZ) | CW (mil) | CCW (mil) |
| H1 | 18281 | -22331 | 21.6 | -26.4 |
| H2 | 24413 | -19937 | 28.9 | -23.6 |
| H3 | 18099 | -22126 | 21.4 | -26.2 |

Table - Optic table range of motion

## Step 5.2 - Test Nº2 – Push “locally”

|  | Push in positive direction | Push in negative direction | Railing | Actuator Gap Check |
| --- | --- | --- | --- | --- |
| H1 | 20691 | -26431 |  | X |
| H2 | 24701 | -24100 |  | X |
| H3 | 25028 | -22929 |  | X |
| V1 | 19785 | -20422 |  | X |
| V2 | 31629 | -32519 |  | X |
| V3 | 19762 | -21787 |  | X |

Table - Optic table range of motion

Acceptance criteria:

- The vertical sensor readout be positive when the optic table is pushed in the +Z direction

- The horizontal sensor readout be negative when the optic table is pushed in the +RZ direction

- Step 7.1

  - Absolutes value of all estimated motions must be higher than 16000counts (~0.020”)

- Step 7.2

  - No contact point on sensors

  - Absolute value of sensor read out must be higher than 16000counts (~0.020”)

  - No contact point on actuators

Test result:	Passed:   X   	Failed:         .

## Step 6 - Position Sensors unlocked/locked Power Spectrum

Data files in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/ Powerspectra/Undamped

- LLO_HAM_ISI_Unit_3_Calibrated_PSD_CPS_GS13_Unlocked_Locked_2011-03-25.mat

- LLO_HAM_ISI_Unit_3_Calibrated_PSD_GS13_Table_Tilted_2011-03-29.mat

Scripts files for processing and plotting in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Powerspectra/Undamped

- Powerspectra_Measurement_HAM_ISI_Locked_Unlocked.m

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Scripts/Data_collection/

- - Powerspectra_Measurements_Tilted_HAM_ISI.m

Figures in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/ Powerspectra/Undamped

- LLO_HAM_ISI_Unit_3_Calibrated_PSD_CPS_GS13_Unlocked_Locked_2011-03-25.mat

- LLO_HAM_ISI_Unit_3_Calibrated_PSD_GS13_Table_Tilted_2011-03-29.mat

CPS calibration:

The CPS power spectrums are calibrated by using a sensitivity of 30.2 nm/count.

Figure - Calibrated CPS power spectrum

Figure – Power spectrum Calibrated GS13

Acceptance criteria:

- No cross talk (peaks at low frequencies + harmonics on measurements)

- Magnitudes of power spectra must be between requirement curves such as in the following figures (dashed lines)

| Sensors | ISI state | Frequency (Hz) | 2x10-2 | 1x10-1 | 1 | 10 | 20 | 100 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| GS-13 | Table locked | Max | 3x10-1 | 3x10-4 | 3x10-7 | 10-7 |  | 10-11 | 10-14 |
|  |  | Min | 3x10-4 | 3x10-7 | 3x10-10 | 10-12 |  | 10-14 | 10-17 |
|  | Table unlocked | Max | 1 | 3x10-3 | 10-5 | 10-9 |  | 10-11 | 10-14 |
|  |  | Min | 10-4 | 3x10-7 | 10-9 | 10-13 |  | 10-15 | 10-18 |
| CPS | Table locked | Max | 2*10-7 | 2x10-8 | 10-8 | 5x10-8 | 2x10-7 | 5x10-9 | 10-9 |
|  |  | Min | 5x10-9 | 2x10-9 | 8x10-10 | 5x10-10 |  | 10-10 | 5x10-11 |
|  | Table unlocked | Max | 2x10-6 | 8x10-7 | 8x10-7 | 5x10-8 | 2x10-7 | 2x10-8 | 10-9 |
|  |  | Min | 10-7 | 5x10-8 | 8x10-9 | 5x10-10 |  | 2x10-10 | 10-10 |

Table - Step 6 -Normal conditions-Sensors power spectra requirements

Test result:	Passed:   X   	Failed:          .

## Step 7 - GS13 power spectrum -tabled tilted

The figure below presents the GS13 power spectrum when the table is unlocked and loaded with a 20Kg mass at one of its corner.

Figure – Power spectrum Calibrated GS13 with mass at corner

Data files in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/ Powerspectra/Undamped

- LLO_HAM_ISI_Unit_3_Calibrated_PSD_CPS_GS13_Unlocked_Locked_2011-03-25.mat

- LLO_HAM_ISI_Unit_3_Calibrated_PSD_GS13_Table_Tilted_2011-03-29.mat

Scripts files for taking and processing the data, and plotting it in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Scripts/Data_collection/

- Powerspectra_Measurements_Tilted_HAM_ISI.m

Figures in SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Figures/ Powerspectra/Undamped

- LLO_HAM_Unit_3_Calibrated_PSD_GS13_Unlocked_Locked_2011-03-25.pdf

- LLO_HAM_ISI_Unit_3_Calibrated_PSD_GS13_Table_Tilted_2011-03-29.pdf

Acceptance criteria:

- With table unlocked and tilted, magnitudes of power spectra must be fully included within:

| Sensor | ISI State | Frequency | 5x10-1 Hz | 1 | 10 | 100 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| GS-13 | Table Tilted | Max | 2x10--5 | 2x10--6 | 8x10-8 | 4x10-11 | 3x10-14 |
|  |  | Min | 10-8 | 10-9 | 2x10-11 | 10-14 | 10-17 |

Table - Table Tilted- Sensors power spectra requirements

Test result:	Passed:   X   	Failed:          .

## Step 8- GS13 pressure readout

Scripts files for taking and processing the data, and plotting it in SVN at: seismicSVN/Common/MatlabTools

- gs13Presstest.m

Figures in SVN at:

- seismicSVN/Common/Data/Pressure_Plots

Figure – Pressure Plots

Issues/difficulties/comments regarding this test

We know that the drop to 0 at the end of the plot is due to issues with getdata.

We can notice issues with V2 dropping a few hundreds of counts now and then.

We could not measure V3, due to an issue with the read-out electronics.

Acceptance criteria:

- The pressure on all channels must be 25000 counts +/- 3000 counts

- All channels must follow comparable trend

Test result:	Passed:      	Failed:   X       .

## Step 9 - Coil Driver, cabling and resistance check

| Actuator | V1 | H1 | V2 |
| --- | --- | --- | --- |
| Coil driver | S1000317 - Coarse 1 | S1000317 - Fine 1 | S1000317 - Fine 2 |
| Anti image pin # |  |  |  |
| Cable # | 28 | 25 | 30 |
| Resistance (Ohm) | P1 - P2 | P2 - P3 | P1 - P2 | P2 - P3 | P1 - P2 | P2 - P3 |
|  | 6.3 | O.L (infinity) | 6.5 | O.L (infinity) | 6.4 | O.L (infinity) |
| MEDM offset (1000 counts) | Measurement P2 (-) ; P1&P3 (+) | Measurement P2 (-) ; P1&P3 (+) | Measurement P2 (-) ; P1&P3 (+) |
|  | 0.3 | 0.303 | 0.3 |
| Actuator | H2 | V3 | H3 |
| Coil driver | S1000317 - Coarse 2 | S1000316 - Coarse 1 | S1000316 - Fine 1 |
| Anti image pin # |  |  |  |
| Cable # | 29 | 26 | 27 |
| Resistance (Ohm) | P1 - P2 | P2 - P3 | P1 - P2 | P2 - P3 | P1 - P2 | P2 - P3 |
|  | 6.4 | O.L (infinity) | 6.4 | O.L (infinity) | 6.3 | O.L (infinity) |
| MEDM offset (1000 counts) | Measurement P2 (-) ; P1&P3 (+) | Measurement P2 (-) ; P1&P3 (+) | Measurement P2 (-) ; P1&P3 (+) |
|  | 0.303 | 0.296 | 0.298 |

Table - Actuators resistance check

Acceptance criteria:

- The measured resistance between the middle pin and one side pin must be 6.5 +/-1 ohms

- Actuator neutral pins must be connected on pin #1 (left side pin of the plug)

- Actuator drive pins must be connected on pin #2 (middle pin of the plug)

- Actuator ground shield pins must be connected on pin #3 (right pin of the plug)

- All LEDs on the coil driver front panel must be green

The tests report must contain:

- The table “Actuators resistance check”

- Issues/difficulties/comments regarding this test

- Test result (Passed:       	Failed:          ).

## Step 10 - Actuators Sign and range of motion (Local drive)

|  | Negative drive | Positive drive |
| --- | --- | --- |
| H1 readout (count) | -24840 | 23704 |
| H2 readout (count) | -23504 | 24473 |
| H3 readout (count) | -25079 | 24232 |
| V1 readout (count) | -19988 | 19535 |
| V2 readout (count) | -25296 | 27191 |
| V3 readout (count) | -22424 | 21599 |

Table - Range of motion - Local drive

Data files in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Static_Tests

- LLO_HAM-ISI_Unit_3_Range_Of_Motion_20110329.mat

Scripts files for taking and processing the data, and plotting it in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Scripts/Data_collection

- Range_Motion_HAM_ISI.m

Acceptance criteria:

- Main couplings sensors readout must be at least 16000 counts (~0.02”)

- A positive offset drive on one actuator must give positive sensor readout on the collocated sensor. Signs will also be tested when measuring local to local transfer functions.

Test result:	Passed:   X   	Failed:         .

## Step 11 - Vertical Sensor Calibration

| Lockers | D.I readout for a negative drive | D.I readout without any drive | D.I readout for positive drive |  |
| --- | --- | --- | --- | --- |
| A | 19 | 0 | -19 |  |
| B | 19.25 | 0 | -19 |  |
| C | 18.5 | 0 | -19 |  |
| D | 19 | 0 | -19 |  |
| Average | 18.9375 | 0 | -19 | -37.9 |
| Sensors | Counts | Counts | Counts | Difference (Counts) |
| V1 | -15407 | 703 | 16730 | -32137 |
| V2 | -15595 | 551 | 16437 | -32032 |
| V3 | -17675 | -1772 | 14357 | -32032 |
|  |  |  | Average | -32067 |

Table - Calibration of capacitive position sensors

Vertical sensitivity:	32067/37.91 = 845.3 count/mil

or	845.3 count/mil * 1/1638 V/count = 0.516V/mil

or	25400nm/mil * 1/845.3 mil/count = 30.05 nm/count

Nominal Calibration

CPS Sensitivity: 	20V/0.039" = 20V/39mils = 0.513V/mil

Calibration in counts: 	215 / 20 * 20/39 = 840 count/mil

or	25400 nm/mil * 1/840 mil/count = 30.2nm/count

Difference with Nominal sensitivity = (845.3-840)/840=-0.63%

Acceptance criteria:

- Deviation from nominal value < 2%. Nominal value is 840 count/mil.

Test result:	Passed:   X   	Failed:          .

## Step 12 - Vertical Spring Constant

Results presented below are obtained after the initial sensors calibration.

| Sensors | Mean diff counts | Mean diff m | K (N/m) | Error with average |
| --- | --- | --- | --- | --- |
| V1 | 7370 | 2.23E-04 | 8.81E+04 | 5.59 % |
| V2 | 7538 | 2.28E-04 | 8.62E+04 | 3.32 % |
| V3 | 8553 | 2.58E-04 | 7.60E+04 | -8.91 % |
|  |  | Average (N/m) | 8.34E+04 |  |
|  |  | Total Stiffness (N/m) | 2.50E+05 |  |

Table - Vertical spring constant

Acceptance criteria:

- +/-2 % of 2.4704e5 N/m (i.e. between 2.421e5 and 2.520e5 N/m)

- +/- 5% of variation between each spring and the average

The measured error on the vertical stiffness is 1.32%.

Test result:	Passed:      	Failed:   X

## Step 13 - Static Testing (Tests in the local basis)

|  |  | H1 | H2 | H3 | V1 | V2 | V3 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Actuators (1000 counts) | H1 | 1967.669 | 1210.834 | 1224.792 | -3.392 | 12.744 | -33.92 |
|  | H2 | 1207.9290 | 2017.293 | 1256.419 | 11.939 | 21.785 | -21.05 |
|  | H3 | 1224.299 | 1259.544 | 2015.90 | 37.030 | 16.072 | -30.849 |
|  | V1 | 201.7530 | 172.320 | -313.735 | 1415.569 | -43.636 | -546.33 |
|  | V2 | -276.287 | 261.940 | 230.952 | -554.369 | 1437.609 | -44.9600 |
|  | V3 | 159.632 | -385.887 | 142.389 | 44.354 | -612.036 | 1403.714 |

Table - Main and cross coupling

Data files in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Static_Tests

- LLO_HAM_ISI_Unit_3_Sensor_Readout_Local_20110331.mat

Scripts files for taking data in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Scripts/Data_Collection

- Sensor_Readout_Local_Drive_MEDM_HAM_ISI.m

Issues/difficulties/comments regarding this test:

Acceptance criteria:

- Vertical

For a +1000 count offset drive on vertical actuators

  - Collocated sensors must be 1400 counts +/- 10%

- Horizontal

For a +1000 count offset drive on horizontal actuators

  - Collocated sensors must be 2000 counts +/- 10%

  - Non-collocated horizontal sensors must be 1250 counts +/-10%

Test result:	Passed:   X    	Failed:         .

## Step 14 - Linearity test

|  | Slope | Offset | Average slope | Variation from average(%) |
| --- | --- | --- | --- | --- |
| H1 | 2.077 | 127.6 | 2.0845 | -0.37 |
| H2 | 2.102 | -763 |  | 0.83 |
| H3 | 2.075 | -711 |  | -0.46 |
| V1 | 1.498 | -282 | 1.4872 | 0.74 |
| V2 | 1.479 | 368.3 |  | -0.53 |
| V3 | 1.484 | -1024 |  | -0.20 |

Table - Slopes and offset of the triplet Actuators - HAM-ISI - Sensors

Scripts files for taking data in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Scripts/Data_Collection

- Linearity_Test_Awgstream_HAM_ISI.m

Figure - Horizontal and vertical actuators x HAM-ISI x sensors

Data files in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Linearity_test

- LLO_HAM_ISI_Unit_3_Linearity_test_20110407.mat

Figures in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Figures/Linearity_test

- LLO_HAM_ISI_Unit_3_Linearity_test_20110407.fig

- LLO_HAM_ISI_Unit_3_Linearity_test_20110407.pdf

Acceptance criteria:

- Horizontal and vertical slopes of the triplet actuators x HAM-ISI x sensors:  Average slope +/- 1%

Test result:	Passed:   X   	Failed:         .

## Step 15 - Cartesian Basis Static Testing

|  | 1000 counts drive | X Drive | Y Drive | Z Drive | Rx Drive | Ry Drive | Rz Drive |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Sensors readout (count) | H1 | 263.528 | -390.4432 | 39.232 | -351.599 | -234.314 | -1870.593 |
|  | H2 | 232.73 | 510.05 | 51.46 | 511.84 | -214.09 | -1926.44 |
|  | H3 | -492.32 | 23.53 | 10.56 | 70.06 | 532.44 | -1901.82 |
|  | V1 | -5.871 | 6.292 | 248.899 | -510.236 | -1619.426 | 11.019 |
|  | V2 | -21.28 | -33.566 | 239.421 | 1633.514 | 398.43 | -57.855 |
|  | V3 | 2.8 | -18.2 | 270.36 | -1169.8 | 1208.911 | 29.8 |
|  | Direction read out | 492.38 | 524.71 | 256.965 | 2516.66 | 2506.73 | 2404.763 |

Table - Tests in the general coordinate basis

Issues/difficulties/comments regarding this test:

Acceptance criteria:

For a positive drive in the Cartesian basis:

- Local sensor readout must have the same sign that the reference table (CONT2ACT check)

- Cartesian sensors read out must be positive (DISP2CEN check) in the drive direction

Test result:	Passed:   X   	Failed:         .

## Step 16- Frequency response

Compensation filters of the new GS13 interface chassis are located in the geophone pre-filters bank. Powerspectra were measured with masses on the optic table not bolted.

## Step 16.1 - Local to local measurements

Local to local transfer functions have been measured with 90 repetitions.

Data files in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_Functions/Measurements/

Undamped/

- LLO_HAM_ISI_Unit_3_Data_TF_L2L_50mHz_500mHz_20110406-181629.mat

- LLO_HAM_ISI_Unit_3_Data_TF_L2L_500mHz_5Hz_20110406-151108.mat

- LLO_HAM_ISI_Unit_3_Data_TF_L2L_200Hz_800Hz_20110406-120426.mat

- LLO_HAM_ISI_Unit_3_Data_TF_L2L_5Hz_200Hz_20110406-133747.mat

Data collection script files:

opt/svncommon/seisvn/seismic/HAM-ISI/X2/Scripts/Data_Collection

- Run_TF_L2L_50mHz_500mHz.m

- Run_TF_L2L_500mHz_5Hz.m

- Run_TF_L2L_5mHz_200Hz.m

- Run_TF_L2L_200Hz_800Hz.m

Scripts files for processing and plotting in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_Functions/Measurements/

Undamped/

- Plot_LLO_HAM_ISI_Unit_3_TF_L2L_2011_04_06.m

Figures in SVN at:

opt/svncommon/seisvn/seismic/HAMISI/X2/Data/Unit_3/Figures/Transfer_Functions/Measurements/

Undamped/

- LLO_HAM_ISI_Unit_3_TF_L2L_H_CPS_50mHz_800Hz_2011_04_06.fig

- LLO_HAM_ISI_Unit_3_TF_L2L_V_CPS_50mHz_800Hz_2011_04_06.fig

- LLO_HAM_ISI_Unit_3_TF_L2L_H_GS13_50mHz_800Hz_2011_04_06.fig

- LLO_HAM_ISI_Unit_3_TF_L2L_V_GS13_50mHz_800Hz_2011_04_06.fig

Storage of measured transfer functions in the SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_functions/ Measurements/

Undamped/

- LLO_HAM_ISI_Unit_3_Data_TF_L2L_2011_04_06.mat

The local to local transfer functions are presented below.

Figure - Local to Local Measurements – Horizontal capacitive sensors

Figure - Local to Local Measurements – Horizontal inertial sensors

Figure - Local to Local Measurements – Vertical capacitive sensors

Figure - Local to Local Measurements – Vertical inertial sensors

Issues/difficulties/comments regarding this test:

Around 1 Hz, the inertial sensors seem to indicate different behaviors of each corner.

## Step 16.2 - Cartesian to Cartesian measurements

Cartesian to Cartesian transfer functions have been measured with 90 repetitions.

Data files in SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_Functions/Measurements/

Undamped/

- LLO_HAM_ISI_Unit_3_Data_TF_C2C_50mHz_500mHz_20110405-215335.mat

- LLO_HAM_ISI_Unit_3_Data_TF_C2C_500mHz_5Hz_20110405-184814.mat

- LLO_HAM_ISI_Unit_3_Data_TF_C2C_5Hz_200Hz_20110405-171453.mat

- LLO_HAM_ISI_Unit_3_Data_TF_C2C_200Hz_800Hz_20110405-154132.mat

Scripts files for processing and plotting in SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_Functions/Measurements/

Undamped/

- Plot_LLO_HAM_ISI_Unit_3_TF_C2C_2011_04_06.m

Figures in SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Figures/Transfer_Functions/

Measurements/Undamped/

- LLO_HAM_ISI_Unit_3_TF_C2C_X_Y_RZ_CPS_50mHz_800Hz_2011_04_05.fig

- LLO_HAM_ISI_Unit_3_TF_C2C_X_Y_RZ_GS13_50mHz_800Hz_2011_04_05.fig

- LLO_HAM_ISI_Unit_3_TF_C2C_Z_RX_RY_CPS_50mHz_800Hz_2011_04_05.fig

- LLO_HAM_ISI_Unit_3_TF_C2C_Z_RX_RY_GS13_50mHz_800Hz_2011_04_05.fig

Storage of measured transfer functions in the SVN at:

/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_functions/Cartesian_to_Cartesian

- LLO_HAM_ISI_Unit_3_Data_TF_C2C_2011_04_05

Figure - Cartesian to Cartesian CPS measurements – X, Y, RZ directions

Figure - Cartesian to Cartesian GS-13 measurements – X, Y, RZ directions

Figure - Cartesian to Cartesian measurements – Z, RX, RY directions

Issues/difficulties/comments regarding this test:

Acceptance criteria:

- Local to local measurements

  - On CPS, the phase must be 0º at DC

  - On Geophones, the phase must be -90º at DC

  - Identical shape in each corner

- Cartesian to Cartesian measurements

  - On CPS, the phase must be 0º at DC

  - On Geophones, the phase must be -90º at DC

  - Identical shape X/Y and RX/RY

Test result:	Passed:   X   	Failed:         .

## Step 17 - Transfer function comparison with Reference

## Step 17.1 - Local to local - Comparison with Reference

This is the 1st unit compared to LHO Unit #2 instead of LLO HAM 6 (v4 was comparing to LLO HAM 6 and both data can be found on the SVN).

Scripts files for processing and plotting in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_Functions/Measurements/

Undamped/

- Plot_LLO_HAM_ISI_Unit_3_TF_L2L_2011_04_06.m

Local to local figures in SVN at:

/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Figures/Transfer_Functions/Measurements/

- LLO_HAM_ISI_Unit_3_TF_L2L_H_CPS_50mHz_800Hz_wRef_2011_04_06.fig

- LLO_HAM_ISI_Unit_3_TF_L2L_V_CPS_50mHz_800Hz_wRef_2011_04_06.fig

- LLO_HAM_ISI_Unit_3_TF_L2L_H_GS13_50mHz_800Hz_wRef_2011_04_06.fig

- LLO_HAM_ISI_Unit_3_TF_L2L_V_GS13_50mHz_800Hz_wRef_2011_04_06.fig

GS13, Local to local measurement

Figure - Local to local measurements comparison with LHO UNIT 2 – Horizontal GS-13

Figure - Local to local measurements comparison with LHO UNIT 2 – Vertical GS13

CPS, Local to local measurement, Undamped

Figure - Local to local measurements comparison – Horizontal Position sensors

Figure - Local to local measurements comparison – Vertical Position sensors

## Step 17.2 - Cartesian to Cartesian - Comparison with Reference

Scripts files for processing and plotting in SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_Functions/Measurements/

Undamped/

- Plot_LLO_HAM_ISI_Unit_3_TF_C2C_2011_04_06.m

Cartesian to Cartesian figures in SVN at :

/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_functions/Cartesian_to_Cartesian

- LLO_HAM_ISI_Unit_3_TF_C2C_X_Y_RZ_CPS_50mHz_800Hz_wRef_2011_04_05.fig

- LLO_HAM_ISI_Unit_3_TF_C2C_Z_RX_RY_CPS_50mHz_800Hz_wRef_2011_04_05.fig

- LLO_HAM_ISI_Unit_3_TF_C2C_X_Y_RZ_GS13_50mHz_800Hz_wRef_2011_04_05.fig

- LLO_HAM_ISI_Unit_3_TF_C2C_Z_RX_RY_GS13_50mHz_800Hz_wRef_2011_04_05.fig

GS13, Cartesian to Cartesian measurement, Undamped

Figure - Cartesian to Cartesian measurements comparison with LHO Unit2 – Horiz GS13

Figure - Cartesian to Cartesian measurements comparison with LHO Unit2 – Vertical GS13

CPS, Cartesian to Cartesian measurement, Undamped

Figure - Cartesian to Cartesian measurements comparison with LHO Unit2 – Horizontal Position sensors

Figure - Cartesian to Cartesian measurements comparison with LHO Unit2 – Vertical Position sensors

Acceptance criteria:

- No difference with the reference transfer functions (SVN)

  - Phase – less than 10º - In Phase – Out of Phase

  - Damping (fit by eye with Reference transfer functions)

  - DC gain

  - Eigen frequencies shift less than 10%

Test result:	Passed:   X   	Failed:         .

## Step 18 - Lower Zero Moment Plane

Data collection script files:

opt/svncommon/seisvn/seismic/HAM-ISI/X2/Scripts/Data_Collection

- Run_Cart2Cart_10mHz_100mHz.m

Data files in SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_Functions/Measurements/

Undamped/

- LZMP_LLO_HAM-ISI-Unit_3_2011_04_06.mat

Scripts files for processing and plotting in SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_Functions/Measurements/

Undamped/

- LZMP_2011_04_06.m

Figures in SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Figures/Transfer_Functions/

Measurements/Undamped/

- LZMP.fig

X & Y offsets:

| X offset (mm) | 0.399 |
| --- | --- |
| Y offset (mm) | 0.738 |

Table – Offset of the Lower Zero Moment Plane

The results from two measurements are presented on the figure below:

Figure - Lower Zero moment plane – Main and cross couplings at low frequency

Issues/difficulties/comments regarding this test

It was discovered after the testing was complete and the unit stored away, that this measurement was done with an insufficient amount of averages. This measurement should be redone during Phase #2 of the testing.

Acceptance criteria:

- X offset must be less than 2 mm

- Y offset must be less than 2 mm

Test result:	Passed:   X    	Failed:         .

## Step 19 - Damping loops

In this step, HAM6 damping loops are implemented. First, damping performances are evaluated in simulation. Second, Damping loops are implemented and performance is experimentally measured.

## Step 19.1 - Transfer functions - Simulation

Scripts files for processing and plotting in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_Functions/Simulation/Damping

- HAM_ISI_LLO_Unit_3_Damping_TF_2011_04_18.m

Figures in SVN at:

/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_Functions/Simulations/

Damping/

- Damping_LOOP_H1_H2_H3.fig

- Damping_LOOP_V1_V2_V3.fig

Results are saved in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_Functions/Simulations/

Damping/

The following figures present the plant, controller, open loop, closed loop and sensitivity of vertical and horizontal damping loops. H1 (respectively V1) are plotted in solid line, H2 (respectively V2) are plotted in dash line, H3 (respectively V3) are plotted in dash-dot line.

Figure - Horizontal damping loops - Simulation

Figure - Vertical damping loops - Simulation

Acceptance criteria:

- HAM6 damping loops must implemented and stable with

  - Phase margin must be at least 45º

  - Gain margin must be at least 20dB

Test result:	Passed:   X   	Failed:         .

## Step 19.2 - Powerspectra – Experimental

Data files in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Powerspectra/Damping/

Scripts files for taking data and plotting in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Scripts/Data_Collection/

- Powerspectra_Measurements_Undamped_Damped_HAM_ISI.m

Figures in SVN at:

/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Figures/Powerspectra/Damping/

- LLO_HAM_ISI_Unit_3_Calibrated_PSD_CPS_Undamped_Damped_2011_04_08.fig

- Simulation_vs_experimental_Suppression.fig

Sensitivity:

The figure below compare the sensitivity ('Undamped/Damped') of LLO HAM (Aug 2008) and LHO Unit 2. Performances are very similar, which confirms that we can use the damping loop as they are (modulo electronics change compensation). The plot also shows that the measured performance matches with the prediction.

Figure – Horizontal (left) and vertical (right) damping loops - Experimental

Acceptance criteria:

- HAM6 damping loop must stable when all damping loops are engaged

- Similar damping effect than in simulated plots

Test result:	Passed:   X   	Failed:         .

Conclusion

A few issues were found during the testing of this unit. Some were only discovered after reprocessing of the data following the requirements update. This unit was approved based on the precedent version of this document (v4). The known issues are summed up here:

- sensor gaps not recorded on the jig

- actuator gaps do not meet requirements, it was decided during the approval meeting to

- LZMP measurements were done with only 1 average (high uncertainty), they should be redone on the side of the chamber

- Vertical spring constant: one spring appears to have a much lower constant than the others (-8.91% than the average of the 3). Because no significant difference can be seen between the 3 corners in the vertical local to local transfer functions, it's likely that it's due

- Could not check pressure on V3 GS-13 and V2 GS-13 gives some weird results.