## Outgassing Rates
A-OSEM Assembly
Species | AMU | Outgassing Rate (Torr-L/s per AOSEM body)
H | 2 | #N/A
He | 4 | 9.2683e-13
C+ | 12 | 2.0281e-09
N+, N2+ | 14 | 5.3845e-10
water (O+) | 16 | 5.8343e-09
water (HO+) | 17 | 2.4752e-08
water (H2O+) | 18 | 9.4023e-08
CO2++ | 22 | 2.3061e-10
N2 | 28 | 1.2178e-08
O2 | 32 | 1.3007e-10
Ar | 40 | 1.5442e-10
hydrocarbon | 41 | 7.4558e-13
hydrocarbon | 43 | 4.2233e-10
CO2 | 44 | 2.7386e-08
hydrocarbon | 53 | 2.7547e-13
hydrocarbon | 55 | 5.7113e-13
hydrocarbon | 57 | 3.7003e-13
Sum HC |  | 4.243e-10
Comparison of A-OSEM Related Outgassing Rates:
| AMU |  | Outgassing Rate (Torr-L/s/AOSEM) | Compare to: AOSEM Wound Heads (E1000374) 404 units | Compare to: Esd 480 material (E1000298) equivalent to ~25 A-OSEM heads | Compare to: AOSEM circuit bd + connector sub-assemblies (E1000376)
hydrogen (H2) | 2 |  | #N/A | 3.6702e-12 | 7.035e-12 | 3.8675e-13
water (H2O) at 1000 hr | max of 16, 17, 18 |  | 9.4023e-08 | 2.951e-14 | 2.7225e-11 | 1.4519e-12
air (N2) | 28 |  | 1.2178e-08 | 3.707e-13 | 5.6069e-11 | 1.799e-11
CO2 | 44 |  | 2.7386e-08 | 9.6659e-15 | 4.6079e-10 | 3.6772e-12
hydrocarbons | sum of 41, 43, 53, 55, 57 |  | 4.243e-10 | 4.6268e-15 | 1.4555e-11 | 4.1192e-13
N.B.: For the time being (at least) the outgassing rates from the RGA results reported in this document for the first 7 A-OSEM production units wil be used in the aLIGO Residual Gas Estimate (E0900398)

## VBOC+Bake+324+RGA+Oven
2010-07-14 | 2010 09:21:42 AM
Residual Gas Analyzer Software
Software Version | 3.205.003
A
| -1.27e-19 |  | T-L/s
4 | 1.99e-19 | 3.2727e-19 | 3.945e-12
4.1 | 2.18e-19 | 3.4627e-19 | 4.3217e-12
4.2 | 1.34e-20 | 1.4167e-19 | 2.6564e-13
4.3 | 5.4e-20 | 1.8227e-19 | 1.0705e-12
4.4 | 9.07e-20 | 2.1897e-19 | 1.798e-12
4.5 | 1.29e-19 | 2.5727e-19 | 2.5573e-12
4.6 | 1.91e-20 | 1.4737e-19 | 3.7864e-13
4.7 | 2.48e-20 | 1.5307e-19 | 4.9164e-13
4.8 | -5.85e-20 | 6.977e-20 | -1.1597e-12
4.9 | -4.63e-20 | 8.197e-20 | -9.1786e-13
5 | -4.86e-20 | 7.967e-20 | -9.6345e-13
5.1 | 8.03e-20 | 2.0857e-19 | 1.5919e-12
5.2 | 1e-19 | 2.2827e-19 | 1.9824e-12
5.3 | 1.79e-20 | 1.4617e-19 | 3.5485e-13
5.4 | 9.85e-20 | 2.2677e-19 | 1.9527e-12
5.5 | 3.1e-20 | 1.5927e-19 | 6.1455e-13
5.6 | 2.36e-20 | 1.5187e-19 | 4.6785e-13
5.7 | -1.76e-20 | 1.1067e-19 | -3.489e-13
5.8 | 1.12e-19 | 2.4027e-19 | 2.2203e-12
5.9 | -5.25e-20 | 7.577e-20 | -1.0408e-12
6 | -7.07e-20 | 5.757e-20 | -1.4016e-12
6.1 | 7.46e-21 | 1.3573e-19 | 1.4789e-13
6.2 | 4.24e-20 | 1.7067e-19 | 8.4054e-13
6.3 | 1.91e-20 | 1.4737e-19 | 3.7864e-13
6.4 | 9.55e-20 | 2.2377e-19 | 1.8932e-12
6.5 | -1.37e-20 | 1.1457e-19 | -2.7159e-13
6.6 | -1.07e-20 | 1.1757e-19 | -2.1212e-13
6.7 | -1.46e-20 | 1.1367e-19 | -2.8943e-13
6.8 | -5.67e-21 | 1.226e-19 | -1.124e-13
6.9 | 5.82e-20 | 1.8647e-19 | 1.1538e-12
7 | 1.85e-20 | 1.4677e-19 | 3.6675e-13
7.1 | 7.76e-21 | 1.3603e-19 | 1.5383e-13
7.2 | 4.98e-20 | 1.7807e-19 | 9.8724e-13
7.3 | -7.79e-20 | 5.037e-20 | -1.5443e-12
7.4 | 2.42e-20 | 1.5247e-19 | 4.7974e-13
7.5 | 4.24e-20 | 1.7067e-19 | 8.4054e-13
7.6 | 2e-20 | 1.4827e-19 | 3.9648e-13
7.7 | 4.3e-20 | 1.7127e-19 | 8.5244e-13
7.8 | 6.27e-20 | 1.9097e-19 | 1.243e-12
7.9 | 1.2e-19 | 2.4827e-19 | 2.3789e-12
8 | 5.97e-20 | 1.8797e-19 | 1.1835e-12
8.1 | 1.04e-19 | 2.3227e-19 | 2.0617e-12
8.2 | 9.25e-21 | 1.3752e-19 | 1.8337e-13
8.3 | 7.34e-20 | 2.0167e-19 | 1.4551e-12
8.4 | 4.54e-20 | 1.7367e-19 | 9.0001e-13
8.5 | -2.84e-20 | 9.987e-20 | -5.63e-13
8.6 | -1.61e-20 | 1.1217e-19 | -3.1917e-13
8.7 | -2.39e-20 | 1.0437e-19 | -4.738e-13
8.8 | -2.54e-20 | 1.0287e-19 | -5.0353e-13
8.9 | -1.64e-20 | 1.1187e-19 | -3.2512e-13
9 | 9.37e-20 | 2.2197e-19 | 1.8575e-12
9.1 | 5.76e-20 | 1.8587e-19 | 1.1419e-12
9.2 | -3.85e-20 | 8.977e-20 | -7.6323e-13
9.3 | 5.88e-20 | 1.8707e-19 | 1.1657e-12
9.4 | 4.78e-20 | 1.7607e-19 | 9.4759e-13
9.5 | 3.91e-20 | 1.6737e-19 | 7.7512e-13
9.6 | 1.45e-19 | 2.7327e-19 | 2.8745e-12
9.7 | 1.41e-19 | 2.6927e-19 | 2.7952e-12
9.8 | 2.89e-20 | 1.5717e-19 | 5.7292e-13
9.9 | 3.46e-20 | 1.6287e-19 | 6.8591e-13
10 | 1.23e-19 | 2.5127e-19 | 2.4384e-12
10.1 | 5.73e-20 | 1.8557e-19 | 1.1359e-12
10.2 | 1.23e-19 | 2.5127e-19 | 2.4384e-12
10.3 | 1.84e-19 | 3.1227e-19 | 3.6476e-12
10.4 | 3.44e-19 | 4.7227e-19 | 6.8195e-12
10.5 | 3.56e-19 | 4.8427e-19 | 7.0574e-12
10.6 | 4.18e-19 | 5.4627e-19 | 8.2865e-12
10.7 | 6.92e-19 | 8.2027e-19 | 1.3718e-11
10.8 | 1.02e-18 | 1.1483e-18 | 2.0221e-11
10.9 | 1.64e-18 | 1.7683e-18 | 3.2512e-11
11 | 2.9e-18 | 3.0283e-18 | 5.749e-11
11.1 | 6.18e-18 | 6.3083e-18 | 1.2251e-10
11.2 | 1.36e-17 | 1.3728e-17 | 2.6961e-10
11.3 | 3.78e-17 | 3.7928e-17 | 7.4935e-10
11.4 | 1.25e-16 | 1.2513e-16 | 2.478e-09
11.5 | 2.9e-16 | 2.9013e-16 | 5.749e-09
11.6 | 4.05e-16 | 4.0513e-16 | 8.0288e-09
11.7 | 5.04e-16 | 5.0413e-16 | 9.9913e-09
11.8 | 6.18e-16 | 6.1813e-16 | 1.2251e-08
11.9 | 6.85e-16 | 6.8513e-16 | 1.3579e-08
12 | 7.16e-16 | 7.1613e-16 | 1.4194e-08
12.1 | 6.21e-16 | 6.2113e-16 | 1.2311e-08
12.2 | 5.04e-16 | 5.0413e-16 | 9.9913e-09
12.3 | 3.19e-16 | 3.1913e-16 | 6.3239e-09
12.4 | 4.38e-17 | 4.3928e-17 | 8.683e-10
12.5 | 4.92e-18 | 5.0483e-18 | 9.7535e-11
12.6 | 4.13e-18 | 4.2583e-18 | 8.1873e-11
12.7 | 4.7e-18 | 4.8283e-18 | 9.3173e-11
12.8 | 5.7e-18 | 5.8283e-18 | 1.13e-10
12.9 | 7.19e-18 | 7.3183e-18 | 1.4254e-10
13 | 8.14e-18 | 8.2683e-18 | 1.6137e-10
13.1 | 8.61e-18 | 8.7383e-18 | 1.7069e-10
13.2 | 9.54e-18 | 9.6683e-18 | 1.8912e-10
13.3 | 1.31e-17 | 1.3228e-17 | 2.597e-10
13.4 | 3.21e-17 | 3.2228e-17 | 6.3635e-10
13.5 | 8.47e-17 | 8.4828e-17 | 1.6791e-09
13.6 | 1.21e-16 | 1.2113e-16 | 2.3987e-09
13.7 | 1.41e-16 | 1.4113e-16 | 2.7952e-09
13.8 | 1.67e-16 | 1.6713e-16 | 3.3106e-09
13.9 | 1.85e-16 | 1.8513e-16 | 3.6675e-09
14 | 1.9e-16 | 1.9013e-16 | 3.7666e-09
14.1 | 1.73e-16 | 1.7313e-16 | 3.4296e-09
14.2 | 1.42e-16 | 1.4213e-16 | 2.815e-09
14.3 | 1.01e-16 | 1.0113e-16 | 2.0022e-09
14.4 | 1.82e-17 | 1.8328e-17 | 3.608e-10
14.5 | 3.48e-18 | 3.6083e-18 | 6.8988e-11
14.6 | 3.69e-18 | 3.8183e-18 | 7.3151e-11
14.7 | 4.6e-18 | 4.7283e-18 | 9.1191e-11
14.8 | 6.31e-18 | 6.4383e-18 | 1.2509e-10
14.9 | 8.85e-18 | 8.9783e-18 | 1.7544e-10
15 | 1.35e-17 | 1.3628e-17 | 2.6763e-10
15.1 | 2.5e-17 | 2.5128e-17 | 4.956e-10
15.2 | 5.54e-17 | 5.5528e-17 | 1.0983e-09
15.3 | 1.38e-16 | 1.3813e-16 | 2.7357e-09
15.4 | 4.14e-16 | 4.1413e-16 | 8.2072e-09
15.5 | 1.06e-15 | 1.0601e-15 | 2.1014e-08
15.6 | 1.46e-15 | 1.4601e-15 | 2.8943e-08
15.7 | 1.68e-15 | 1.6801e-15 | 3.3304e-08
15.8 | 1.91e-15 | 1.9101e-15 | 3.7864e-08
15.9 | 2.01e-15 | 2.0101e-15 | 3.9846e-08
16 | 2.06e-15 | 2.0601e-15 | 4.0838e-08
16.1 | 1.89e-15 | 1.8901e-15 | 3.7468e-08
16.2 | 1.7e-15 | 1.7001e-15 | 3.3701e-08
16.3 | 1.56e-15 | 1.5601e-15 | 3.0926e-08
16.4 | 1.73e-15 | 1.7301e-15 | 3.4296e-08
16.5 | 4.26e-15 | 4.2601e-15 | 8.4451e-08
16.6 | 5.92e-15 | 5.9201e-15 | 1.1736e-07
16.7 | 7.04e-15 | 7.0401e-15 | 1.3956e-07
16.8 | 8.12e-15 | 8.1201e-15 | 1.6097e-07
16.9 | 8.54e-15 | 8.5401e-15 | 1.693e-07
17 | 8.74e-15 | 8.7401e-15 | 1.7326e-07
17.1 | 7.96e-15 | 7.9601e-15 | 1.578e-07
17.2 | 7.28e-15 | 7.2801e-15 | 1.4432e-07
17.3 | 6.57e-15 | 6.5701e-15 | 1.3024e-07
17.4 | 5.89e-15 | 5.8901e-15 | 1.1676e-07
17.5 | 1.5e-14 | 1.5e-14 | 2.9736e-07
17.6 | 2.26e-14 | 2.26e-14 | 4.4802e-07
17.7 | 2.7e-14 | 2.7e-14 | 5.3525e-07
17.8 | 3.11e-14 | 3.11e-14 | 6.1653e-07
17.9 | 3.24e-14 | 3.24e-14 | 6.423e-07
18 | 3.32e-14 | 3.32e-14 | 6.5816e-07
18.1 | 2.98e-14 | 2.98e-14 | 5.9076e-07
18.2 | 2.61e-14 | 2.61e-14 | 5.1741e-07
18.3 | 1.93e-14 | 1.93e-14 | 3.826e-07
18.4 | 2.26e-15 | 2.2601e-15 | 4.4802e-08
18.5 | 2.38e-16 | 2.3813e-16 | 4.7181e-09
18.6 | 2.57e-16 | 2.5713e-16 | 5.0948e-09
18.7 | 2.95e-16 | 2.9513e-16 | 5.8481e-09
18.8 | 3.44e-16 | 3.4413e-16 | 6.8195e-09
18.9 | 3.74e-16 | 3.7413e-16 | 7.4142e-09
19 | 3.82e-16 | 3.8213e-16 | 7.5728e-09
19.1 | 3.57e-16 | 3.5713e-16 | 7.0772e-09
19.2 | 2.77e-16 | 2.7713e-16 | 5.4913e-09
19.3 | 1.94e-16 | 1.9413e-16 | 3.8459e-09
19.4 | 3.14e-17 | 3.1528e-17 | 6.2248e-10
19.5 | 3.52e-17 | 3.5328e-17 | 6.9781e-10
19.6 | 5.39e-17 | 5.4028e-17 | 1.0685e-09
19.7 | 6.62e-17 | 6.6328e-17 | 1.3124e-09
19.8 | 7.66e-17 | 7.6728e-17 | 1.5185e-09
19.9 | 7.74e-17 | 7.7528e-17 | 1.5344e-09
20 | 8.12e-17 | 8.1328e-17 | 1.6097e-09
20.1 | 7.33e-17 | 7.3428e-17 | 1.4531e-09
20.2 | 6.48e-17 | 6.4928e-17 | 1.2846e-09
20.3 | 4.83e-17 | 4.8428e-17 | 9.575e-10
20.4 | 1.11e-17 | 1.1228e-17 | 2.2005e-10
20.5 | 9.74e-19 | 1.1023e-18 | 1.9309e-11
20.6 | 5.76e-20 | 1.8587e-19 | 1.1419e-12
20.7 | -4.6e-20 | 8.227e-20 | -9.1191e-13
20.8 | 4.3e-20 | 1.7127e-19 | 8.5244e-13
20.9 | 1.27e-19 | 2.5527e-19 | 2.5177e-12
21 | 4.19e-19 | 5.4727e-19 | 8.3063e-12
21.1 | 8.08e-19 | 9.3627e-19 | 1.6018e-11
21.2 | 1.79e-18 | 1.9183e-18 | 3.5485e-11
21.3 | 4.73e-18 | 4.8583e-18 | 9.3768e-11
21.4 | 1.57e-17 | 1.5828e-17 | 3.1124e-10
21.5 | 4.17e-17 | 4.1828e-17 | 8.2666e-10
21.6 | 5.7e-17 | 5.7128e-17 | 1.13e-09
21.7 | 6.98e-17 | 6.9928e-17 | 1.3837e-09
21.8 | 7.78e-17 | 7.7928e-17 | 1.5423e-09
21.9 | 7.78e-17 | 7.7928e-17 | 1.5423e-09
22 | 8.13e-17 | 8.1428e-17 | 1.6117e-09
22.1 | 7.19e-17 | 7.2028e-17 | 1.4254e-09
22.2 | 6.3e-17 | 6.3128e-17 | 1.2489e-09
22.3 | 4.5e-17 | 4.5128e-17 | 8.9208e-10
22.4 | 6.22e-18 | 6.3483e-18 | 1.2331e-10
22.5 | 1.42e-18 | 1.5483e-18 | 2.815e-11
22.6 | 1.17e-18 | 1.2983e-18 | 2.3194e-11
22.7 | 1.17e-18 | 1.2983e-18 | 2.3194e-11
22.8 | 1.03e-18 | 1.1583e-18 | 2.0419e-11
22.9 | 6.43e-19 | 7.7127e-19 | 1.2747e-11
23 | 4.54e-19 | 5.8227e-19 | 9.0001e-12
23.1 | 4.35e-19 | 5.6327e-19 | 8.6235e-12
23.2 | 5.04e-19 | 6.3227e-19 | 9.9913e-12
23.3 | 3.35e-19 | 4.6327e-19 | 6.6411e-12
23.4 | 9.64e-20 | 2.2467e-19 | 1.911e-12
23.5 | 2.57e-20 | 1.5397e-19 | 5.0948e-13
23.6 | 1.88e-20 | 1.4707e-19 | 3.7269e-13
23.7 | 9.79e-20 | 2.2617e-19 | 1.9408e-12
23.8 | -1.16e-20 | 1.1667e-19 | -2.2996e-13
23.9 | 4.15e-20 | 1.6977e-19 | 8.227e-13
24 | -6.48e-20 | 6.347e-20 | -1.2846e-12
24.1 | -8.09e-20 | 4.737e-20 | -1.6038e-12
24.2 | 2.98e-21 | 1.3125e-19 | 5.9076e-14
24.3 | -2.33e-20 | 1.0497e-19 | -4.619e-13
24.4 | 4.18e-20 | 1.7007e-19 | 8.2865e-13
24.5 | 7.97e-20 | 2.0797e-19 | 1.58e-12
24.6 | 9.82e-20 | 2.2647e-19 | 1.9467e-12
24.7 | 6.66e-20 | 1.9487e-19 | 1.3203e-12
24.8 | 4.12e-20 | 1.6947e-19 | 8.1675e-13
24.9 | 3.64e-20 | 1.6467e-19 | 7.216e-13
25 | 3.31e-20 | 1.6137e-19 | 6.5618e-13
25.1 | 6.39e-20 | 1.9217e-19 | 1.2668e-12
25.2 | 1.23e-19 | 2.5127e-19 | 2.4384e-12
25.3 | 6.09e-20 | 1.8917e-19 | 1.2073e-12
25.4 | 1.39e-19 | 2.6727e-19 | 2.7555e-12
25.5 | 3.38e-19 | 4.6627e-19 | 6.7005e-12
25.6 | 4.32e-19 | 5.6027e-19 | 8.564e-12
25.7 | 4.32e-19 | 5.6027e-19 | 8.564e-12
25.8 | 5.41e-19 | 6.6927e-19 | 1.0725e-11
25.9 | 5.74e-19 | 7.0227e-19 | 1.1379e-11
26 | 8.04e-19 | 9.3227e-19 | 1.5939e-11
26.1 | 1.02e-18 | 1.1483e-18 | 2.0221e-11
26.2 | 1.22e-18 | 1.3483e-18 | 2.4185e-11
26.3 | 1.32e-18 | 1.4483e-18 | 2.6168e-11
26.4 | 1.68e-18 | 1.8083e-18 | 3.3304e-11
26.5 | 2.68e-18 | 2.8083e-18 | 5.3129e-11
26.6 | 3.7e-18 | 3.8283e-18 | 7.3349e-11
26.7 | 6e-18 | 6.1283e-18 | 1.1894e-10
26.8 | 1.09e-17 | 1.1028e-17 | 2.1608e-10
26.9 | 1.94e-17 | 1.9528e-17 | 3.8459e-10
27 | 3.9e-17 | 3.9128e-17 | 7.7314e-10
27.1 | 7.3e-17 | 7.3128e-17 | 1.4472e-09
27.2 | 1.49e-16 | 1.4913e-16 | 2.9538e-09
27.3 | 3.48e-16 | 3.4813e-16 | 6.8988e-09
27.4 | 9.55e-16 | 9.5513e-16 | 1.8932e-08
27.5 | 2.25e-15 | 2.2501e-15 | 4.4604e-08
27.6 | 3.06e-15 | 3.0601e-15 | 6.0662e-08
27.7 | 3.7e-15 | 3.7001e-15 | 7.3349e-08
27.8 | 4.14e-15 | 4.1401e-15 | 8.2072e-08
27.9 | 4.13e-15 | 4.1301e-15 | 8.1873e-08
28 | 4.3e-15 | 4.3001e-15 | 8.5244e-08 | 1.0929e-07 | effective air leak rate (T-L/s) for the load (7 AOSEMs)
28.1 | 3.88e-15 | 3.8801e-15 | 7.6917e-08 | 1.5612e-08 | effective air leak rate per AOSEM (T-L/s)
28.2 | 3.33e-15 | 3.3301e-15 | 6.6014e-08
28.3 | 2.43e-15 | 2.4301e-15 | 4.8173e-08
28.4 | 2.75e-16 | 2.7513e-16 | 5.4516e-09
28.5 | 2.59e-17 | 2.6028e-17 | 5.1344e-10
28.6 | 2.4e-17 | 2.4128e-17 | 4.7578e-10
28.7 | 2.82e-17 | 2.8328e-17 | 5.5904e-10
28.8 | 3.22e-17 | 3.2328e-17 | 6.3834e-10
28.9 | 3.3e-17 | 3.3128e-17 | 6.5419e-10
29 | 3.5e-17 | 3.5128e-17 | 6.9384e-10
29.1 | 3.22e-17 | 3.2328e-17 | 6.3834e-10
29.2 | 2.85e-17 | 2.8628e-17 | 5.6499e-10
29.3 | 2.1e-17 | 2.1128e-17 | 4.1631e-10
29.4 | 5.31e-18 | 5.4383e-18 | 1.0527e-10
29.5 | 1.78e-18 | 1.9083e-18 | 3.5287e-11
29.6 | 2.34e-18 | 2.4683e-18 | 4.6388e-11
29.7 | 3.23e-18 | 3.3583e-18 | 6.4032e-11
29.8 | 3.73e-18 | 3.8583e-18 | 7.3944e-11
29.9 | 4.04e-18 | 4.1683e-18 | 8.0089e-11
30 | 4.26e-18 | 4.3883e-18 | 8.4451e-11
30.1 | 4.08e-18 | 4.2083e-18 | 8.0882e-11
30.2 | 3.51e-18 | 3.6383e-18 | 6.9583e-11
30.3 | 2.71e-18 | 2.8383e-18 | 5.3723e-11
30.4 | 1.02e-18 | 1.1483e-18 | 2.0221e-11
30.5 | 2.49e-19 | 3.7727e-19 | 4.9362e-12
30.6 | 3.22e-19 | 4.5027e-19 | 6.3834e-12
30.7 | 4.16e-19 | 5.4427e-19 | 8.2468e-12
30.8 | 4.45e-19 | 5.7327e-19 | 8.8217e-12
30.9 | 6.11e-19 | 7.3927e-19 | 1.2113e-11
31 | 8.23e-19 | 9.5127e-19 | 1.6315e-11
31.1 | 1.11e-18 | 1.2383e-18 | 2.2005e-11
31.2 | 1.8e-18 | 1.9283e-18 | 3.5683e-11
31.3 | 3.57e-18 | 3.6983e-18 | 7.0772e-11
31.4 | 8.79e-18 | 8.9183e-18 | 1.7425e-10
31.5 | 2.14e-17 | 2.1528e-17 | 4.2424e-10
31.6 | 2.95e-17 | 2.9628e-17 | 5.8481e-10
31.7 | 3.59e-17 | 3.6028e-17 | 7.1168e-10
31.8 | 4.23e-17 | 4.2428e-17 | 8.3856e-10
31.9 | 4.17e-17 | 4.1828e-17 | 8.2666e-10
32 | 4.58e-17 | 4.5928e-17 | 9.0794e-10
32.1 | 4.13e-17 | 4.1428e-17 | 8.1873e-10
32.2 | 3.39e-17 | 3.4028e-17 | 6.7204e-10
32.3 | 2.38e-17 | 2.3928e-17 | 4.7181e-10
32.4 | 4.56e-18 | 4.6883e-18 | 9.0398e-11
32.5 | 4.58e-19 | 5.8627e-19 | 9.0794e-12
32.6 | -4.27e-20 | 8.557e-20 | -8.4649e-13
32.7 | -2.15e-20 | 1.0677e-19 | -4.2622e-13
32.8 | -1.12e-19 | 1.627e-20 | -2.2203e-12
32.9 | -7.25e-20 | 5.577e-20 | -1.4372e-12
33 | 1.01e-20 | 1.3837e-19 | 2.0022e-13
33.1 | 1.37e-20 | 1.4197e-19 | 2.7159e-13
33.2 | 4.3e-20 | 1.7127e-19 | 8.5244e-13
33.3 | 5.61e-20 | 1.8437e-19 | 1.1121e-12
33.4 | 1.07e-20 | 1.3897e-19 | 2.1212e-13
33.5 | 4.8e-20 | 1.7627e-19 | 9.5156e-13
33.6 | 6.74e-20 | 1.9567e-19 | 1.3361e-12
33.7 | 1.29e-19 | 2.5727e-19 | 2.5573e-12
33.8 | 1.64e-19 | 2.9227e-19 | 3.2512e-12
33.9 | 2.46e-19 | 3.7427e-19 | 4.8767e-12
34 | 1.82e-19 | 3.1027e-19 | 3.608e-12
34.1 | 2.77e-19 | 4.0527e-19 | 5.4913e-12
34.2 | 2.28e-19 | 3.5627e-19 | 4.5199e-12
34.3 | 1.78e-19 | 3.0627e-19 | 3.5287e-12
34.4 | 3.28e-19 | 4.5627e-19 | 6.5023e-12
34.5 | 5.89e-19 | 7.1727e-19 | 1.1676e-11
34.6 | 9.18e-19 | 1.0463e-18 | 1.8199e-11
34.7 | 1.2e-18 | 1.3283e-18 | 2.3789e-11
34.8 | 1.39e-18 | 1.5183e-18 | 2.7555e-11
34.9 | 1.66e-18 | 1.7883e-18 | 3.2908e-11
35 | 1.48e-18 | 1.6083e-18 | 2.934e-11
35.1 | 1.44e-18 | 1.5683e-18 | 2.8547e-11
35.2 | 1.43e-18 | 1.5583e-18 | 2.8348e-11
35.3 | 1.09e-18 | 1.2183e-18 | 2.1608e-11
35.4 | 4.65e-19 | 5.9327e-19 | 9.2182e-12
35.5 | 7.41e-19 | 8.6927e-19 | 1.469e-11
35.6 | 8.45e-19 | 9.7327e-19 | 1.6751e-11
35.7 | 1.21e-18 | 1.3383e-18 | 2.3987e-11
35.8 | 1.64e-18 | 1.7683e-18 | 3.2512e-11
35.9 | 1.57e-18 | 1.6983e-18 | 3.1124e-11
36 | 1.83e-18 | 1.9583e-18 | 3.6278e-11
36.1 | 1.88e-18 | 2.0083e-18 | 3.7269e-11
36.2 | 1.6e-18 | 1.7283e-18 | 3.1719e-11
36.3 | 1.18e-18 | 1.3083e-18 | 2.3392e-11
36.4 | 4.03e-19 | 5.3127e-19 | 7.9891e-12
36.5 | 1.75e-19 | 3.0327e-19 | 3.4692e-12
36.6 | 2.28e-19 | 3.5627e-19 | 4.5199e-12
36.7 | 3.26e-19 | 4.5427e-19 | 6.4627e-12
36.8 | 3.47e-19 | 4.7527e-19 | 6.879e-12
36.9 | 4.15e-19 | 5.4327e-19 | 8.227e-12
37 | 3.82e-19 | 5.1027e-19 | 7.5728e-12
37.1 | 3.43e-19 | 4.7127e-19 | 6.7997e-12
37.2 | 3.31e-19 | 4.5927e-19 | 6.5618e-12
37.3 | 3.76e-19 | 5.0427e-19 | 7.4539e-12
37.4 | 2.93e-19 | 4.2127e-19 | 5.8085e-12
37.5 | 2.65e-19 | 3.9327e-19 | 5.2534e-12
37.6 | 3.7e-19 | 4.9827e-19 | 7.3349e-12
37.7 | 3.33e-19 | 4.6127e-19 | 6.6014e-12
37.8 | 4.37e-19 | 5.6527e-19 | 8.6631e-12
37.9 | 4.42e-19 | 5.7027e-19 | 8.7622e-12
38 | 5.62e-19 | 6.9027e-19 | 1.1141e-11
38.1 | 4.28e-19 | 5.5627e-19 | 8.4847e-12
38.2 | 3.2e-19 | 4.4827e-19 | 6.3437e-12
38.3 | 3.61e-19 | 4.8927e-19 | 7.1565e-12
38.4 | 2.42e-19 | 3.7027e-19 | 4.7974e-12
38.5 | 1.53e-19 | 2.8127e-19 | 3.0331e-12
38.6 | 1.47e-19 | 2.7527e-19 | 2.9141e-12
38.7 | 2.91e-19 | 4.1927e-19 | 5.7688e-12
38.8 | 2.92e-19 | 4.2027e-19 | 5.7886e-12
38.9 | 5.06e-19 | 6.3427e-19 | 1.0031e-11
39 | 7.96e-19 | 9.2427e-19 | 1.578e-11
39.1 | 1.37e-18 | 1.4983e-18 | 2.7159e-11
39.2 | 2.61e-18 | 2.7383e-18 | 5.1741e-11
39.3 | 5.68e-18 | 5.8083e-18 | 1.126e-10
39.4 | 1.51e-17 | 1.5228e-17 | 2.9934e-10
39.5 | 2.76e-17 | 2.7728e-17 | 5.4714e-10
39.6 | 3.84e-17 | 3.8528e-17 | 7.6124e-10
39.7 | 4.75e-17 | 4.7628e-17 | 9.4164e-10
39.8 | 5.17e-17 | 5.1828e-17 | 1.0249e-09
39.9 | 5.34e-17 | 5.3528e-17 | 1.0586e-09
40 | 5.44e-17 | 5.4528e-17 | 1.0784e-09
40.1 | 5.12e-17 | 5.1328e-17 | 1.015e-09
40.2 | 4.06e-17 | 4.0728e-17 | 8.0486e-10
40.3 | 2.34e-17 | 2.3528e-17 | 4.6388e-10
40.4 | 3.44e-18 | 3.5683e-18 | 6.8195e-11
40.5 | 4.24e-19 | 5.5227e-19 | 8.4054e-12
40.6 | -1.31e-20 | 1.1517e-19 | -2.597e-13
40.7 | 6.57e-20 | 1.9397e-19 | 1.3024e-12
40.8 | 2.45e-19 | 3.7327e-19 | 4.8569e-12
40.9 | 8.95e-20 | 2.1777e-19 | 1.7743e-12
41 | 1.35e-19 | 2.6327e-19 | 2.6763e-12
41.1 | 2.37e-19 | 3.6527e-19 | 4.6983e-12
41.2 | 2e-19 | 3.2827e-19 | 3.9648e-12
41.3 | 1.92e-19 | 3.2027e-19 | 3.8062e-12
41.4 | 3.31e-19 | 4.5927e-19 | 6.5618e-12
41.5 | 3.24e-19 | 4.5227e-19 | 6.423e-12
41.6 | 4.52e-19 | 5.8027e-19 | 8.9605e-12
41.7 | 5.49e-19 | 6.7727e-19 | 1.0883e-11
41.8 | 7.5e-19 | 8.7827e-19 | 1.4868e-11
41.9 | 9.05e-19 | 1.0333e-18 | 1.7941e-11
42 | 1.07e-18 | 1.1983e-18 | 2.1212e-11
42.1 | 1.64e-18 | 1.7683e-18 | 3.2512e-11
42.2 | 2.71e-18 | 2.8383e-18 | 5.3723e-11
42.3 | 4.01e-18 | 4.1383e-18 | 7.9495e-11
42.4 | 6.37e-18 | 6.4983e-18 | 1.2628e-10
42.5 | 1.18e-17 | 1.1928e-17 | 2.3392e-10
42.6 | 2.01e-17 | 2.0228e-17 | 3.9846e-10
42.7 | 3.64e-17 | 3.6528e-17 | 7.216e-10
42.8 | 6.17e-17 | 6.1828e-17 | 1.2231e-09
42.9 | 9.45e-17 | 9.4628e-17 | 1.8734e-09
43 | 1.49e-16 | 1.4913e-16 | 2.9538e-09
43.1 | 2.4e-16 | 2.4013e-16 | 4.7578e-09
43.2 | 4.4e-16 | 4.4013e-16 | 8.7226e-09
43.3 | 9.24e-16 | 9.2413e-16 | 1.8317e-08
43.4 | 2.25e-15 | 2.2501e-15 | 4.4604e-08
43.5 | 4.69e-15 | 4.6901e-15 | 9.2975e-08
43.6 | 6.17e-15 | 6.1701e-15 | 1.2231e-07
43.7 | 7.77e-15 | 7.7701e-15 | 1.5403e-07
43.8 | 8.86e-15 | 8.8601e-15 | 1.7564e-07
43.9 | 9.23e-15 | 9.2301e-15 | 1.8298e-07
44 | 9.67e-15 | 9.6701e-15 | 1.917e-07
44.1 | 9.17e-15 | 9.1701e-15 | 1.8179e-07
44.2 | 7.45e-15 | 7.4501e-15 | 1.4769e-07
44.3 | 4.6e-15 | 4.6001e-15 | 9.1191e-08
44.4 | 6.25e-16 | 6.2513e-16 | 1.239e-08
44.5 | 6.54e-17 | 6.5528e-17 | 1.2965e-09
44.6 | 6.26e-17 | 6.2728e-17 | 1.241e-09
44.7 | 8.12e-17 | 8.1328e-17 | 1.6097e-09
44.8 | 9.22e-17 | 9.2328e-17 | 1.8278e-09
44.9 | 9.75e-17 | 9.7628e-17 | 1.9328e-09
45 | 1.05e-16 | 1.0513e-16 | 2.0815e-09
45.1 | 1.01e-16 | 1.0113e-16 | 2.0022e-09
45.2 | 8.69e-17 | 8.7028e-17 | 1.7227e-09
45.3 | 5.84e-17 | 5.8528e-17 | 1.1577e-09
45.4 | 1.83e-17 | 1.8428e-17 | 3.6278e-10
45.5 | 1.42e-17 | 1.4328e-17 | 2.815e-10
45.6 | 2.03e-17 | 2.0428e-17 | 4.0243e-10
45.7 | 2.73e-17 | 2.7428e-17 | 5.412e-10
45.8 | 3.17e-17 | 3.1828e-17 | 6.2842e-10
45.9 | 3.43e-17 | 3.4428e-17 | 6.7997e-10
46 | 3.64e-17 | 3.6528e-17 | 7.216e-10
46.1 | 3.53e-17 | 3.5428e-17 | 6.9979e-10
46.2 | 3.07e-17 | 3.0828e-17 | 6.086e-10
46.3 | 1.96e-17 | 1.9728e-17 | 3.8855e-10
46.4 | 5.94e-18 | 6.0683e-18 | 1.1776e-10
46.5 | 8.57e-19 | 9.8527e-19 | 1.6989e-11
46.6 | 2.36e-19 | 3.6427e-19 | 4.6785e-12
46.7 | 1.02e-19 | 2.3027e-19 | 2.0221e-12
46.8 | 2.62e-19 | 3.9027e-19 | 5.1939e-12
46.9 | 3.78e-19 | 5.0627e-19 | 7.4935e-12
47 | 4.57e-19 | 5.8527e-19 | 9.0596e-12
47.1 | 4.66e-19 | 5.9427e-19 | 9.238e-12
47.2 | 4.66e-19 | 5.9427e-19 | 9.238e-12
47.3 | 2.42e-19 | 3.7027e-19 | 4.7974e-12
47.4 | 9.22e-20 | 2.2047e-19 | 1.8278e-12
47.5 | 6.15e-20 | 1.8977e-19 | 1.2192e-12
47.6 | 1.01e-19 | 2.2927e-19 | 2.0022e-12
47.7 | 1.06e-19 | 2.3427e-19 | 2.1014e-12
47.8 | 1.52e-19 | 2.8027e-19 | 3.0133e-12
47.9 | 1.74e-19 | 3.0227e-19 | 3.4494e-12
48 | 2.06e-19 | 3.3427e-19 | 4.0838e-12
48.1 | 1.57e-19 | 2.8527e-19 | 3.1124e-12
48.2 | 1.01e-19 | 2.2927e-19 | 2.0022e-12
48.3 | 1.4e-19 | 2.6827e-19 | 2.7754e-12
48.4 | -2.87e-20 | 9.957e-20 | -5.6895e-13
48.5 | -3.58e-21 | 1.2469e-19 | -7.097e-14
48.6 | 1.76e-20 | 1.4587e-19 | 3.489e-13
48.7 | 2.03e-20 | 1.4857e-19 | 4.0243e-13
48.8 | -3.01e-20 | 9.817e-20 | -5.9671e-13
48.9 | -8.06e-20 | 4.767e-20 | -1.5978e-12
49 | 2.03e-20 | 1.4857e-19 | 4.0243e-13
49.1 | -1.61e-20 | 1.1217e-19 | -3.1917e-13
49.2 | 2.39e-21 | 1.3066e-19 | 4.738e-14
49.3 | -3.28e-21 | 1.2499e-19 | -6.5023e-14
49.4 | -1.13e-19 | 1.527e-20 | -2.2401e-12
49.5 | 1.94e-20 | 1.4767e-19 | 3.8459e-13
49.6 | 8.21e-20 | 2.1037e-19 | 1.6276e-12
49.7 | 3.97e-20 | 1.6797e-19 | 7.8702e-13
49.8 | 1.35e-19 | 2.6327e-19 | 2.6763e-12
49.9 | 1.9e-19 | 3.1827e-19 | 3.7666e-12
50 | 1.52e-19 | 2.8027e-19 | 3.0133e-12
50.1 | 1.67e-19 | 2.9527e-19 | 3.3106e-12
50.2 | -2.92e-20 | 9.907e-20 | -5.7886e-13
50.3 | 7.34e-20 | 2.0167e-19 | 1.4551e-12
50.4 | 4.72e-20 | 1.7547e-19 | 9.357e-13
50.5 | 1.07e-19 | 2.3527e-19 | 2.1212e-12
50.6 | 3.13e-20 | 1.5957e-19 | 6.2049e-13
50.7 | 3.31e-20 | 1.6137e-19 | 6.5618e-13
50.8 | -5.67e-21 | 1.226e-19 | -1.124e-13
50.9 | 9.61e-20 | 2.2437e-19 | 1.9051e-12
51 | 1.84e-19 | 3.1227e-19 | 3.6476e-12
51.1 | 1.86e-19 | 3.1427e-19 | 3.6873e-12
51.2 | 1.91e-19 | 3.1927e-19 | 3.7864e-12
51.3 | 1.05e-19 | 2.3327e-19 | 2.0815e-12
51.4 | 9.82e-20 | 2.2647e-19 | 1.9467e-12
51.5 | 1.07e-19 | 2.3527e-19 | 2.1212e-12
51.6 | 1.2e-19 | 2.4827e-19 | 2.3789e-12
51.7 | 2.48e-20 | 1.5307e-19 | 4.9164e-13
51.8 | 7.61e-20 | 2.0437e-19 | 1.5086e-12
51.9 | -6.09e-20 | 6.737e-20 | -1.2073e-12
52 | -7.04e-20 | 5.787e-20 | -1.3956e-12
52.1 | -2.45e-20 | 1.0377e-19 | -4.8569e-13
52.2 | 1.04e-19 | 2.3227e-19 | 2.0617e-12
52.3 | 3.97e-20 | 1.6797e-19 | 7.8702e-13
52.4 | 4.48e-21 | 1.3275e-19 | 8.8812e-14
52.5 | 7.91e-20 | 2.0737e-19 | 1.5681e-12
52.6 | 2.81e-20 | 1.5637e-19 | 5.5706e-13
52.7 | 1.21e-19 | 2.4927e-19 | 2.3987e-12
52.8 | 1.58e-19 | 2.8627e-19 | 3.1322e-12
52.9 | -2.09e-20 | 1.0737e-19 | -4.1432e-13
53 | -3.1e-20 | 9.727e-20 | -6.1455e-13
53.1 | 6.06e-20 | 1.8887e-19 | 1.2013e-12
53.2 | -2.21e-20 | 1.0617e-19 | -4.3811e-13
53.3 | 3.85e-20 | 1.6677e-19 | 7.6323e-13
53.4 | 2.33e-20 | 1.5157e-19 | 4.619e-13
53.5 | 2.78e-20 | 1.5607e-19 | 5.5111e-13
53.6 | 3.58e-20 | 1.6407e-19 | 7.097e-13
53.7 | -4.18e-21 | 1.2409e-19 | -8.2865e-14
53.8 | -5.22e-20 | 7.607e-20 | -1.0348e-12
53.9 | -1.55e-20 | 1.1277e-19 | -3.0727e-13
54 | -1.61e-20 | 1.1217e-19 | -3.1917e-13
54.1 | -9.01e-20 | 3.817e-20 | -1.7862e-12
54.2 | -1.07e-20 | 1.1757e-19 | -2.1212e-13
54.3 | -5.67e-20 | 7.157e-20 | -1.124e-12
54.4 | -6.57e-20 | 6.257e-20 | -1.3024e-12
54.5 | -4.48e-21 | 1.2379e-19 | -8.8812e-14
54.6 | 3.52e-20 | 1.6347e-19 | 6.9781e-13
54.7 | -3.52e-20 | 9.307e-20 | -6.9781e-13
54.8 | 1.4e-20 | 1.4227e-19 | 2.7754e-13
54.9 | -4.48e-20 | 8.347e-20 | -8.8812e-13
55 | 7.34e-20 | 2.0167e-19 | 1.4551e-12
55.1 | 1.31e-20 | 1.4137e-19 | 2.597e-13
55.2 | 8.06e-21 | 1.3633e-19 | 1.5978e-13
55.3 | 1.43e-20 | 1.4257e-19 | 2.8348e-13
55.4 | -2.18e-20 | 1.0647e-19 | -4.3217e-13
55.5 | 1.19e-19 | 2.4727e-19 | 2.3591e-12
55.6 | 1.47e-19 | 2.7527e-19 | 2.9141e-12
55.7 | 1.35e-19 | 2.6327e-19 | 2.6763e-12
55.8 | 2.02e-19 | 3.3027e-19 | 4.0045e-12
55.9 | 1.07e-19 | 2.3527e-19 | 2.1212e-12
56 | -9.85e-21 | 1.1842e-19 | -1.9527e-13
56.1 | -5.79e-20 | 7.037e-20 | -1.1478e-12
56.2 | 6.18e-20 | 1.9007e-19 | 1.2251e-12
56.3 | 1.17e-19 | 2.4527e-19 | 2.3194e-12
56.4 | 9.55e-21 | 1.3782e-19 | 1.8932e-13
56.5 | 8.39e-20 | 2.1217e-19 | 1.6632e-12
56.6 | 1.11e-19 | 2.3927e-19 | 2.2005e-12
56.7 | 4.33e-20 | 1.7157e-19 | 8.5838e-13
56.8 | 2.69e-21 | 1.3096e-19 | 5.3327e-14
56.9 | 6.39e-20 | 1.9217e-19 | 1.2668e-12
57 | 2.39e-21 | 1.3066e-19 | 4.738e-14
57.1 | 2.09e-21 | 1.3036e-19 | 4.1432e-14
57.2 | 1.67e-20 | 1.4497e-19 | 3.3106e-13
57.3 | 1.97e-20 | 1.4797e-19 | 3.9053e-13
57.4 | -1.07e-19 | 2.127e-20 | -2.1212e-12
57.5 | -6.95e-20 | 5.877e-20 | -1.3778e-12
57.6 | 9.73e-20 | 2.2557e-19 | 1.9289e-12
57.7 | 6.6e-20 | 1.9427e-19 | 1.3084e-12
57.8 | -1e-19 | 2.827e-20 | -1.9824e-12
57.9 | -9.61e-20 | 3.217e-20 | -1.9051e-12
58 | 1.79e-21 | 1.3006e-19 | 3.5485e-14
58.1 | -7.52e-20 | 5.307e-20 | -1.4908e-12
58.2 | -4.8e-20 | 8.027e-20 | -9.5156e-13
58.3 | 1.79e-20 | 1.4617e-19 | 3.5485e-13
58.4 | -5.73e-20 | 7.097e-20 | -1.1359e-12
58.5 | -2.89e-20 | 9.937e-20 | -5.7292e-13
58.6 | 3.22e-20 | 1.6047e-19 | 6.3834e-13
58.7 | -6.21e-20 | 6.617e-20 | -1.2311e-12
58.8 | -8.51e-20 | 4.317e-20 | -1.687e-12
58.9 | -4.69e-20 | 8.137e-20 | -9.2975e-13
59 | 2.89e-20 | 1.5717e-19 | 5.7292e-13
59.1 | 3.88e-21 | 1.3215e-19 | 7.6917e-14
59.2 | -2.48e-20 | 1.0347e-19 | -4.9164e-13
59.3 | 1.49e-20 | 1.4317e-19 | 2.9538e-13
59.4 | 1.49e-19 | 2.7727e-19 | 2.9538e-12
59.5 | 8.39e-20 | 2.1217e-19 | 1.6632e-12
59.6 | 1.95e-19 | 3.2327e-19 | 3.8657e-12
59.7 | 1.35e-19 | 2.6327e-19 | 2.6763e-12
59.8 | 1.26e-19 | 2.5427e-19 | 2.4978e-12
59.9 | 1.89e-19 | 3.1727e-19 | 3.7468e-12
60 | 3.31e-19 | 4.5927e-19 | 6.5618e-12
60.1 | 3.17e-19 | 4.4527e-19 | 6.2842e-12
60.2 | 1.95e-19 | 3.2327e-19 | 3.8657e-12
60.3 | 2.3e-19 | 3.5827e-19 | 4.5595e-12
60.4 | 2.36e-19 | 3.6427e-19 | 4.6785e-12
60.5 | 1.76e-19 | 3.0427e-19 | 3.489e-12
60.6 | 4.89e-20 | 1.7717e-19 | 9.694e-13
60.7 | 1.12e-19 | 2.4027e-19 | 2.2203e-12
60.8 | 1.54e-19 | 2.8227e-19 | 3.0529e-12
60.9 | 5.91e-20 | 1.8737e-19 | 1.1716e-12
61 | -1.91e-20 | 1.0917e-19 | -3.7864e-13
61.1 | 1.19e-20 | 1.4017e-19 | 2.3591e-13
61.2 | 3.01e-20 | 1.5837e-19 | 5.9671e-13
61.3 | 1.19e-20 | 1.4017e-19 | 2.3591e-13
61.4 | 5.07e-21 | 1.3334e-19 | 1.0051e-13
61.5 | 1.73e-19 | 3.0127e-19 | 3.4296e-12
61.6 | -2.98e-20 | 9.847e-20 | -5.9076e-13
61.7 | 8.6e-20 | 2.1427e-19 | 1.7049e-12
61.8 | 6.74e-20 | 1.9567e-19 | 1.3361e-12
61.9 | 8.71e-20 | 2.1537e-19 | 1.7267e-12
62 | 6.8e-20 | 1.9627e-19 | 1.348e-12
62.1 | 2.98e-21 | 1.3125e-19 | 5.9076e-14
62.2 | 6.77e-20 | 1.9597e-19 | 1.3421e-12
62.3 | 8.42e-20 | 2.1247e-19 | 1.6692e-12
62.4 | 5.1e-20 | 1.7927e-19 | 1.011e-12
62.5 | 6.71e-20 | 1.9537e-19 | 1.3302e-12
62.6 | 1.5e-19 | 2.7827e-19 | 2.9736e-12
62.7 | 3.31e-20 | 1.6137e-19 | 6.5618e-13
62.8 | 1.43e-19 | 2.7127e-19 | 2.8348e-12
62.9 | 8.54e-20 | 2.1367e-19 | 1.693e-12
63 | 6.57e-20 | 1.9397e-19 | 1.3024e-12
63.1 | 5.94e-20 | 1.8767e-19 | 1.1776e-12
63.2 | 8.39e-20 | 2.1217e-19 | 1.6632e-12
63.3 | 7.49e-20 | 2.0317e-19 | 1.4848e-12
63.4 | -1.49e-20 | 1.1337e-19 | -2.9538e-13
63.5 | 1.41e-19 | 2.6927e-19 | 2.7952e-12
63.6 | 6.66e-20 | 1.9487e-19 | 1.3203e-12
63.7 | 1.01e-20 | 1.3837e-19 | 2.0022e-13
63.8 | 5.16e-20 | 1.7987e-19 | 1.0229e-12
63.9 | 1.5e-19 | 2.7827e-19 | 2.9736e-12
64 | 1.86e-19 | 3.1427e-19 | 3.6873e-12
64.1 | 1.08e-19 | 2.3627e-19 | 2.141e-12
64.2 | 1.25e-19 | 2.5327e-19 | 2.478e-12
64.3 | 1.54e-19 | 2.8227e-19 | 3.0529e-12
64.4 | 1.64e-19 | 2.9227e-19 | 3.2512e-12
64.5 | -6.86e-21 | 1.2141e-19 | -1.3599e-13
64.6 | 6.18e-20 | 1.9007e-19 | 1.2251e-12
64.7 | 4.57e-20 | 1.7397e-19 | 9.0596e-13
64.8 | -5.25e-20 | 7.577e-20 | -1.0408e-12
64.9 | 3.19e-20 | 1.6017e-19 | 6.3239e-13
65 | -4.89e-20 | 7.937e-20 | -9.694e-13
65.1 | -5.82e-20 | 7.007e-20 | -1.1538e-12
65.2 | -7.16e-21 | 1.2111e-19 | -1.4194e-13
65.3 | -4.92e-20 | 7.907e-20 | -9.7535e-13
65.4 | -2.6e-20 | 1.0227e-19 | -5.1543e-13
65.5 | -2.69e-20 | 1.0137e-19 | -5.3327e-13
65.6 | 2.66e-20 | 1.5487e-19 | 5.2732e-13
65.7 | -3.07e-20 | 9.757e-20 | -6.086e-13
65.8 | 6.27e-21 | 1.3454e-19 | 1.243e-13
65.9 | -2.6e-20 | 1.0227e-19 | -5.1543e-13
66 | -4.3e-20 | 8.527e-20 | -8.5244e-13
66.1 | 1.3e-19 | 2.5827e-19 | 2.5771e-12
66.2 | 3.25e-20 | 1.6077e-19 | 6.4428e-13
66.3 | 4.18e-20 | 1.7007e-19 | 8.2865e-13
66.4 | -2.36e-20 | 1.0467e-19 | -4.6785e-13
66.5 | 4.63e-20 | 1.7457e-19 | 9.1786e-13
66.6 | 3.16e-20 | 1.5987e-19 | 6.2644e-13
66.7 | -5.76e-20 | 7.067e-20 | -1.1419e-12
66.8 | -4.06e-20 | 8.767e-20 | -8.0486e-13
66.9 | -5.97e-22 | 1.2767e-19 | -1.1835e-14
67 | 5.97e-20 | 1.8797e-19 | 1.1835e-12
67.1 | 2.95e-20 | 1.5777e-19 | 5.8481e-13
67.2 | 2.54e-20 | 1.5367e-19 | 5.0353e-13
67.3 | 8.98e-20 | 2.1807e-19 | 1.7802e-12
67.4 | -9.25e-21 | 1.1902e-19 | -1.8337e-13
67.5 | 2.63e-20 | 1.5457e-19 | 5.2137e-13
67.6 | 7.46e-20 | 2.0287e-19 | 1.4789e-12
67.7 | 5.85e-20 | 1.8677e-19 | 1.1597e-12
67.8 | 3.64e-20 | 1.6467e-19 | 7.216e-13
67.9 | 2.35e-19 | 3.6327e-19 | 4.6587e-12
68 | 2e-19 | 3.2827e-19 | 3.9648e-12
68.1 | 4.6e-20 | 1.7427e-19 | 9.1191e-13
68.2 | -6.27e-21 | 1.22e-19 | -1.243e-13
68.3 | 5.58e-20 | 1.8407e-19 | 1.1062e-12
68.4 | 7.4e-20 | 2.0227e-19 | 1.467e-12
68.5 | 5.16e-20 | 1.7987e-19 | 1.0229e-12
68.6 | 6.74e-20 | 1.9567e-19 | 1.3361e-12
68.7 | 3.28e-21 | 1.3155e-19 | 6.5023e-14
68.8 | 2.72e-20 | 1.5547e-19 | 5.3922e-13
68.9 | -5.28e-20 | 7.547e-20 | -1.0467e-12
69 | 2.45e-20 | 1.5277e-19 | 4.8569e-13
69.1 | -8.06e-21 | 1.2021e-19 | -1.5978e-13
69.2 | -2.12e-20 | 1.0707e-19 | -4.2027e-13
69.3 | 2.95e-20 | 1.5777e-19 | 5.8481e-13
69.4 | 4.78e-21 | 1.3305e-19 | 9.4759e-14
69.5 | 2.21e-20 | 1.5037e-19 | 4.3811e-13
69.6 | -7.01e-20 | 5.817e-20 | -1.3897e-12
69.7 | -5.22e-20 | 7.607e-20 | -1.0348e-12
69.8 | 8.06e-21 | 1.3633e-19 | 1.5978e-13
69.9 | 1.01e-20 | 1.3837e-19 | 2.0022e-13
70 | -7.79e-20 | 5.037e-20 | -1.5443e-12
70.1 | -5.13e-20 | 7.697e-20 | -1.017e-12
70.2 | -5.31e-20 | 7.517e-20 | -1.0527e-12
70.3 | -5.19e-20 | 7.637e-20 | -1.0289e-12
70.4 | 6.86e-21 | 1.3513e-19 | 1.3599e-13
70.5 | 7.52e-20 | 2.0347e-19 | 1.4908e-12
70.6 | 1.07e-20 | 1.3897e-19 | 2.1212e-13
70.7 | -8.71e-20 | 4.117e-20 | -1.7267e-12
70.8 | 5.07e-21 | 1.3334e-19 | 1.0051e-13
70.9 | -4.18e-20 | 8.647e-20 | -8.2865e-13
71 | -4.66e-20 | 8.167e-20 | -9.238e-13
71.1 | -7.97e-20 | 4.857e-20 | -1.58e-12
71.2 | 5.76e-20 | 1.8587e-19 | 1.1419e-12
71.3 | 6.77e-20 | 1.9597e-19 | 1.3421e-12
71.4 | 8.77e-20 | 2.1597e-19 | 1.7386e-12
71.5 | -6.36e-20 | 6.467e-20 | -1.2608e-12
71.6 | -1.16e-19 | 1.227e-20 | -2.2996e-12
71.7 | 2.54e-20 | 1.5367e-19 | 5.0353e-13
71.8 | -1.28e-20 | 1.1547e-19 | -2.5375e-13
71.9 | 2.09e-20 | 1.4917e-19 | 4.1432e-13
72 | 3.61e-20 | 1.6437e-19 | 7.1565e-13
72.1 | 1.23e-19 | 2.5127e-19 | 2.4384e-12
72.2 | 1.22e-19 | 2.5027e-19 | 2.4185e-12
72.3 | 5.1e-20 | 1.7927e-19 | 1.011e-12
72.4 | -2.48e-20 | 1.0347e-19 | -4.9164e-13
72.5 | 5.37e-21 | 1.3364e-19 | 1.0646e-13
72.6 | 1.49e-20 | 1.4317e-19 | 2.9538e-13
72.7 | 6.33e-20 | 1.9157e-19 | 1.2549e-12
72.8 | -7.85e-20 | 4.977e-20 | -1.5562e-12
72.9 | 6.57e-20 | 1.9397e-19 | 1.3024e-12
73 | 8.36e-20 | 2.1187e-19 | 1.6573e-12
73.1 | 1.13e-20 | 1.3957e-19 | 2.2401e-13
73.2 | 4.45e-20 | 1.7277e-19 | 8.8217e-13
73.3 | 5.4e-20 | 1.8227e-19 | 1.0705e-12
73.4 | 8.71e-20 | 2.1537e-19 | 1.7267e-12
73.5 | -1.94e-20 | 1.0887e-19 | -3.8459e-13
73.6 | -8.3e-20 | 4.527e-20 | -1.6454e-12
73.7 | -8.36e-20 | 4.467e-20 | -1.6573e-12
73.8 | -1.27e-19 | 1.27e-21 | -2.5177e-12
73.9 | 5.19e-20 | 1.8017e-19 | 1.0289e-12
74 | 1.21e-19 | 2.4927e-19 | 2.3987e-12
74.1 | 3.7e-20 | 1.6527e-19 | 7.3349e-13
74.2 | 1.39e-19 | 2.6727e-19 | 2.7555e-12
74.3 | 9.19e-20 | 2.2017e-19 | 1.8218e-12
74.4 | -3.88e-21 | 1.2439e-19 | -7.6917e-14
74.5 | 6.98e-20 | 1.9807e-19 | 1.3837e-12
74.6 | 1.15e-19 | 2.4327e-19 | 2.2798e-12
74.7 | -5.13e-20 | 7.697e-20 | -1.017e-12
74.8 | 2.75e-20 | 1.5577e-19 | 5.4516e-13
74.9 | 2.6e-20 | 1.5427e-19 | 5.1543e-13
75 | 1.79e-21 | 1.3006e-19 | 3.5485e-14
75.1 | 8.21e-20 | 2.1037e-19 | 1.6276e-12
75.2 | -7.97e-20 | 4.857e-20 | -1.58e-12
75.3 | -3.79e-20 | 9.037e-20 | -7.5133e-13
75.4 | 8.03e-20 | 2.0857e-19 | 1.5919e-12
75.5 | 6.24e-20 | 1.9067e-19 | 1.237e-12
75.6 | -2.03e-20 | 1.0797e-19 | -4.0243e-13
75.7 | 1.53e-19 | 2.8127e-19 | 3.0331e-12
75.8 | -2.98e-22 | 1.2797e-19 | -5.9076e-15
75.9 | -1.1e-20 | 1.1727e-19 | -2.1806e-13
76 | 1.94e-20 | 1.4767e-19 | 3.8459e-13
76.1 | 1.36e-19 | 2.6427e-19 | 2.6961e-12
76.2 | 6.15e-20 | 1.8977e-19 | 1.2192e-12
76.3 | 3.91e-20 | 1.6737e-19 | 7.7512e-13
76.4 | 7.73e-20 | 2.0557e-19 | 1.5324e-12
76.5 | 1.46e-20 | 1.4287e-19 | 2.8943e-13
76.6 | -3.76e-20 | 9.067e-20 | -7.4539e-13
76.7 | 5.61e-20 | 1.8437e-19 | 1.1121e-12
76.8 | 2.03e-20 | 1.4857e-19 | 4.0243e-13
76.9 | 5.91e-20 | 1.8737e-19 | 1.1716e-12
77 | 7.76e-21 | 1.3603e-19 | 1.5383e-13
77.1 | 1.38e-19 | 2.6627e-19 | 2.7357e-12
77.2 | 2.1e-19 | 3.3827e-19 | 4.1631e-12
77.3 | 1.08e-19 | 2.3627e-19 | 2.141e-12
77.4 | 5.46e-20 | 1.8287e-19 | 1.0824e-12
77.5 | 6e-20 | 1.8827e-19 | 1.1894e-12
77.6 | 2.21e-19 | 3.4927e-19 | 4.3811e-12
77.7 | 1.15e-19 | 2.4327e-19 | 2.2798e-12
77.8 | 2.33e-19 | 3.6127e-19 | 4.619e-12
77.9 | 1.09e-19 | 2.3727e-19 | 2.1608e-12
78 | 1.24e-19 | 2.5227e-19 | 2.4582e-12
78.1 | 1.85e-19 | 3.1327e-19 | 3.6675e-12
78.2 | 1.87e-19 | 3.1527e-19 | 3.7071e-12
78.3 | 1.07e-19 | 2.3527e-19 | 2.1212e-12
78.4 | 2.18e-19 | 3.4627e-19 | 4.3217e-12
78.5 | 1.12e-19 | 2.4027e-19 | 2.2203e-12
78.6 | 1.08e-19 | 2.3627e-19 | 2.141e-12
78.7 | 1.32e-19 | 2.6027e-19 | 2.6168e-12
78.8 | 2.48e-20 | 1.5307e-19 | 4.9164e-13
78.9 | -2.66e-20 | 1.0167e-19 | -5.2732e-13
79 | -3.16e-20 | 9.667e-20 | -6.2644e-13
79.1 | -1.43e-20 | 1.1397e-19 | -2.8348e-13
79.2 | 1.13e-19 | 2.4127e-19 | 2.2401e-12
79.3 | 1.05e-19 | 2.3327e-19 | 2.0815e-12
79.4 | 1.54e-19 | 2.8227e-19 | 3.0529e-12
79.5 | 2.97e-19 | 4.2527e-19 | 5.8878e-12
79.6 | 2.08e-19 | 3.3627e-19 | 4.1234e-12
79.7 | -2.69e-20 | 1.0137e-19 | -5.3327e-13
79.8 | -7.67e-20 | 5.157e-20 | -1.5205e-12
79.9 | 4.89e-20 | 1.7717e-19 | 9.694e-13
80 | 2.3e-20 | 1.5127e-19 | 4.5595e-13
80.1 | -4.42e-20 | 8.407e-20 | -8.7622e-13
80.2 | -8.77e-20 | 4.057e-20 | -1.7386e-12
80.3 | 1.58e-19 | 2.8627e-19 | 3.1322e-12
80.4 | 1.18e-19 | 2.4627e-19 | 2.3392e-12
80.5 | 7.1e-20 | 1.9927e-19 | 1.4075e-12
80.6 | 7.4e-20 | 2.0227e-19 | 1.467e-12
80.7 | 5.31e-20 | 1.8137e-19 | 1.0527e-12
80.8 | 8.18e-20 | 2.1007e-19 | 1.6216e-12
80.9 | 2.12e-20 | 1.4947e-19 | 4.2027e-13
81 | -4.12e-20 | 8.707e-20 | -8.1675e-13
81.1 | -9.64e-20 | 3.187e-20 | -1.911e-12
81.2 | -2.84e-20 | 9.987e-20 | -5.63e-13
81.3 | 1.41e-19 | 2.6927e-19 | 2.7952e-12
81.4 | 3.07e-20 | 1.5897e-19 | 6.086e-13
81.5 | -3.73e-20 | 9.097e-20 | -7.3944e-13
81.6 | 1.76e-20 | 1.4587e-19 | 3.489e-13
81.7 | -2e-20 | 1.0827e-19 | -3.9648e-13
81.8 | 3.19e-20 | 1.6017e-19 | 6.3239e-13
81.9 | -2.75e-20 | 1.0077e-19 | -5.4516e-13
82 | -1.88e-20 | 1.0947e-19 | -3.7269e-13
82.1 | 2.63e-20 | 1.5457e-19 | 5.2137e-13
82.2 | 6.48e-20 | 1.9307e-19 | 1.2846e-12
82.3 | 1.61e-19 | 2.8927e-19 | 3.1917e-12
82.4 | 7.79e-20 | 2.0617e-19 | 1.5443e-12
82.5 | 2.18e-20 | 1.5007e-19 | 4.3217e-13
82.6 | 2.39e-21 | 1.3066e-19 | 4.738e-14
82.7 | 3.88e-20 | 1.6707e-19 | 7.6917e-13
82.8 | -5.52e-20 | 7.307e-20 | -1.0943e-12
82.9 | 1.28e-20 | 1.4107e-19 | 2.5375e-13
83 | -1.07e-20 | 1.1757e-19 | -2.1212e-13
83.1 | 5.46e-20 | 1.8287e-19 | 1.0824e-12
83.2 | 4.18e-21 | 1.3245e-19 | 8.2865e-14
83.3 | 3.28e-21 | 1.3155e-19 | 6.5023e-14
83.4 | -2.24e-20 | 1.0587e-19 | -4.4406e-13
83.5 | -6.71e-20 | 6.117e-20 | -1.3302e-12
83.6 | -4.48e-20 | 8.347e-20 | -8.8812e-13
83.7 | -3.94e-20 | 8.887e-20 | -7.8107e-13
83.8 | -2.27e-20 | 1.0557e-19 | -4.5001e-13
83.9 | 2.98e-22 | 1.2857e-19 | 5.9076e-15
84 | 3.85e-20 | 1.6677e-19 | 7.6323e-13
84.1 | 9.58e-20 | 2.2407e-19 | 1.8991e-12
84.2 | 1.16e-19 | 2.4427e-19 | 2.2996e-12
84.3 | 4.06e-20 | 1.6887e-19 | 8.0486e-13
84.4 | 7.97e-20 | 2.0797e-19 | 1.58e-12
84.5 | 1.18e-19 | 2.4627e-19 | 2.3392e-12
84.6 | 3.49e-20 | 1.6317e-19 | 6.9186e-13
84.7 | 2.48e-20 | 1.5307e-19 | 4.9164e-13
84.8 | 7.04e-20 | 1.9867e-19 | 1.3956e-12
84.9 | 1.17e-19 | 2.4527e-19 | 2.3194e-12
85 | 2.39e-20 | 1.5217e-19 | 4.738e-13
85.1 | -4.09e-20 | 8.737e-20 | -8.1081e-13
85.2 | -2.6e-20 | 1.0227e-19 | -5.1543e-13
85.3 | 1.25e-20 | 1.4077e-19 | 2.478e-13
85.4 | 4.72e-20 | 1.7547e-19 | 9.357e-13
85.5 | 1.26e-19 | 2.5427e-19 | 2.4978e-12
85.6 | 1.35e-19 | 2.6327e-19 | 2.6763e-12
85.7 | 1.25e-19 | 2.5327e-19 | 2.478e-12
85.8 | 1.04e-20 | 1.3867e-19 | 2.0617e-13
85.9 | 4.66e-20 | 1.7487e-19 | 9.238e-13
86 | 1.28e-19 | 2.5627e-19 | 2.5375e-12
86.1 | 2.04e-19 | 3.3227e-19 | 4.0441e-12
86.2 | 9.58e-20 | 2.2407e-19 | 1.8991e-12
86.3 | 8.71e-20 | 2.1537e-19 | 1.7267e-12
86.4 | 1.37e-20 | 1.4197e-19 | 2.7159e-13
86.5 | 6.69e-20 | 1.9517e-19 | 1.3262e-12
86.6 | 1.53e-19 | 2.8127e-19 | 3.0331e-12
86.7 | 7.49e-20 | 2.0317e-19 | 1.4848e-12
86.8 | 1.11e-19 | 2.3927e-19 | 2.2005e-12
86.9 | 8.95e-20 | 2.1777e-19 | 1.7743e-12
87 | 1.33e-19 | 2.6127e-19 | 2.6366e-12
87.1 | 1.36e-19 | 2.6427e-19 | 2.6961e-12
87.2 | 2.51e-20 | 1.5337e-19 | 4.9758e-13
87.3 | -2.51e-20 | 1.0317e-19 | -4.9758e-13
87.4 | -6.69e-20 | 6.137e-20 | -1.3262e-12
87.5 | -4.18e-21 | 1.2409e-19 | -8.2865e-14
87.6 | 1.21e-19 | 2.4927e-19 | 2.3987e-12
87.7 | 1.06e-19 | 2.3427e-19 | 2.1014e-12
87.8 | 1.19e-19 | 2.4727e-19 | 2.3591e-12
87.9 | 4.98e-20 | 1.7807e-19 | 9.8724e-13
88 | 2.13e-19 | 3.4127e-19 | 4.2225e-12
88.1 | 2.98e-19 | 4.2627e-19 | 5.9076e-12
88.2 | 2.38e-19 | 3.6627e-19 | 4.7181e-12
88.3 | 1.67e-19 | 2.9527e-19 | 3.3106e-12
88.4 | 4.06e-20 | 1.6887e-19 | 8.0486e-13
88.5 | 3.88e-21 | 1.3215e-19 | 7.6917e-14
88.6 | 9.19e-20 | 2.2017e-19 | 1.8218e-12
88.7 | 9.73e-20 | 2.2557e-19 | 1.9289e-12
88.8 | -1.79e-21 | 1.2648e-19 | -3.5485e-14
88.9 | 4.03e-20 | 1.6857e-19 | 7.9891e-13
89 | 7.49e-20 | 2.0317e-19 | 1.4848e-12
89.1 | -6e-20 | 6.827e-20 | -1.1894e-12
89.2 | 4.83e-20 | 1.7657e-19 | 9.575e-13
89.3 | 3.73e-20 | 1.6557e-19 | 7.3944e-13
89.4 | 4.18e-20 | 1.7007e-19 | 8.2865e-13
89.5 | 1.37e-20 | 1.4197e-19 | 2.7159e-13
89.6 | -2.51e-20 | 1.0317e-19 | -4.9758e-13
89.7 | 7.55e-20 | 2.0377e-19 | 1.4967e-12
89.8 | 4.86e-20 | 1.7687e-19 | 9.6345e-13
89.9 | -5.97e-21 | 1.223e-19 | -1.1835e-13
90 | 1.08e-19 | 2.3627e-19 | 2.141e-12
90.1 | 5.13e-20 | 1.7957e-19 | 1.017e-12
90.2 | 2e-20 | 1.4827e-19 | 3.9648e-13
90.3 | 3.88e-21 | 1.3215e-19 | 7.6917e-14
90.4 | 4.18e-20 | 1.7007e-19 | 8.2865e-13
90.5 | 1.41e-19 | 2.6927e-19 | 2.7952e-12
90.6 | 2.87e-20 | 1.5697e-19 | 5.6895e-13
90.7 | 2.87e-20 | 1.5697e-19 | 5.6895e-13
90.8 | 1.82e-19 | 3.1027e-19 | 3.608e-12
90.9 | 4.36e-20 | 1.7187e-19 | 8.6433e-13
91 | 8.86e-20 | 2.1687e-19 | 1.7564e-12
91.1 | 9.73e-20 | 2.2557e-19 | 1.9289e-12
91.2 | 1.56e-19 | 2.8427e-19 | 3.0926e-12
91.3 | 6.69e-20 | 1.9517e-19 | 1.3262e-12
91.4 | 1.33e-19 | 2.6127e-19 | 2.6366e-12
91.5 | 2.12e-19 | 3.4027e-19 | 4.2027e-12
91.6 | 1.83e-19 | 3.1127e-19 | 3.6278e-12
91.7 | 1.17e-19 | 2.4527e-19 | 2.3194e-12
91.8 | 1.27e-19 | 2.5527e-19 | 2.5177e-12
91.9 | 8.06e-20 | 2.0887e-19 | 1.5978e-12
92 | 9.73e-20 | 2.2557e-19 | 1.9289e-12
92.1 | 9.4e-20 | 2.2227e-19 | 1.8635e-12
92.2 | 1.35e-19 | 2.6327e-19 | 2.6763e-12
92.3 | 1.14e-19 | 2.4227e-19 | 2.2599e-12
92.4 | 1.28e-19 | 2.5627e-19 | 2.5375e-12
92.5 | 1e-19 | 2.2827e-19 | 1.9824e-12
92.6 | 1.02e-19 | 2.3027e-19 | 2.0221e-12
92.7 | 9.22e-20 | 2.2047e-19 | 1.8278e-12
92.8 | -8.15e-20 | 4.677e-20 | -1.6157e-12
92.9 | -4.36e-20 | 8.467e-20 | -8.6433e-13
93 | -4.86e-20 | 7.967e-20 | -9.6345e-13
93.1 | 5.28e-20 | 1.8107e-19 | 1.0467e-12
93.2 | 9.04e-20 | 2.1867e-19 | 1.7921e-12
93.3 | 1.04e-20 | 1.3867e-19 | 2.0617e-13
93.4 | -4.63e-20 | 8.197e-20 | -9.1786e-13
93.5 | 3.28e-20 | 1.6107e-19 | 6.5023e-13
93.6 | 1.96e-19 | 3.2427e-19 | 3.8855e-12
93.7 | 3.34e-20 | 1.6167e-19 | 6.6212e-13
93.8 | -1.4e-20 | 1.1427e-19 | -2.7754e-13
93.9 | 1.46e-20 | 1.4287e-19 | 2.8943e-13
94 | 1.04e-19 | 2.3227e-19 | 2.0617e-12
94.1 | 8.6e-20 | 2.1427e-19 | 1.7049e-12
94.2 | 5.91e-20 | 1.8737e-19 | 1.1716e-12
94.3 | 4.66e-20 | 1.7487e-19 | 9.238e-13
94.4 | 1.16e-19 | 2.4427e-19 | 2.2996e-12
94.5 | -4.78e-21 | 1.2349e-19 | -9.4759e-14
94.6 | 6.51e-20 | 1.9337e-19 | 1.2905e-12
94.7 | 1.42e-19 | 2.7027e-19 | 2.815e-12
94.8 | 1.47e-19 | 2.7527e-19 | 2.9141e-12
94.9 | 8.3e-20 | 2.1127e-19 | 1.6454e-12
95 | 4.95e-20 | 1.7777e-19 | 9.8129e-13
95.1 | 6.57e-21 | 1.3484e-19 | 1.3024e-13
95.2 | 7.16e-21 | 1.3543e-19 | 1.4194e-13
95.3 | 1.55e-19 | 2.8327e-19 | 3.0727e-12
95.4 | 2.18e-20 | 1.5007e-19 | 4.3217e-13
95.5 | 8.03e-20 | 2.0857e-19 | 1.5919e-12
95.6 | 1.66e-19 | 2.9427e-19 | 3.2908e-12
95.7 | 2.12e-19 | 3.4027e-19 | 4.2027e-12
95.8 | 9.01e-20 | 2.1837e-19 | 1.7862e-12
95.9 | 8.6e-20 | 2.1427e-19 | 1.7049e-12
96 | 8.3e-20 | 2.1127e-19 | 1.6454e-12
96.1 | 8.57e-20 | 2.1397e-19 | 1.6989e-12
96.2 | 5.94e-20 | 1.8767e-19 | 1.1776e-12
96.3 | -1.13e-20 | 1.1697e-19 | -2.2401e-13
96.4 | -5.97e-22 | 1.2767e-19 | -1.1835e-14
96.5 | -5.79e-20 | 7.037e-20 | -1.1478e-12
96.6 | -3.1e-20 | 9.727e-20 | -6.1455e-13
96.7 | 6.8e-20 | 1.9627e-19 | 1.348e-12
96.8 | 7.19e-20 | 2.0017e-19 | 1.4254e-12
96.9 | 9.97e-20 | 2.2797e-19 | 1.9765e-12
97 | 1.45e-19 | 2.7327e-19 | 2.8745e-12
97.1 | 1.07e-19 | 2.3527e-19 | 2.1212e-12
97.2 | 1.66e-19 | 2.9427e-19 | 3.2908e-12
97.3 | 1.36e-19 | 2.6427e-19 | 2.6961e-12
97.4 | 3.52e-20 | 1.6347e-19 | 6.9781e-13
97.5 | 2.39e-21 | 1.3066e-19 | 4.738e-14
97.6 | 2e-19 | 3.2827e-19 | 3.9648e-12
97.7 | 1.15e-19 | 2.4327e-19 | 2.2798e-12
97.8 | -4.3e-20 | 8.527e-20 | -8.5244e-13
97.9 | 9.01e-20 | 2.1837e-19 | 1.7862e-12
98 | 2.12e-19 | 3.4027e-19 | 4.2027e-12
98.1 | 2.2e-19 | 3.4827e-19 | 4.3613e-12
98.2 | 6.8e-20 | 1.9627e-19 | 1.348e-12
98.3 | 7.88e-20 | 2.0707e-19 | 1.5621e-12
98.4 | 6.24e-20 | 1.9067e-19 | 1.237e-12
98.5 | 1.91e-19 | 3.1927e-19 | 3.7864e-12
98.6 | 8.48e-20 | 2.1307e-19 | 1.6811e-12
98.7 | 7.91e-20 | 2.0737e-19 | 1.5681e-12
98.8 | 3.24e-19 | 4.5227e-19 | 6.423e-12
98.9 | 8.03e-20 | 2.0857e-19 | 1.5919e-12
99 | 3.58e-21 | 1.3185e-19 | 7.097e-14
99.1 | 1.1e-19 | 2.3827e-19 | 2.1806e-12
99.2 | 2.36e-20 | 1.5187e-19 | 4.6785e-13
99.3 | 1.38e-19 | 2.6627e-19 | 2.7357e-12
99.4 | 2.18e-20 | 1.5007e-19 | 4.3217e-13
99.5 | 6.74e-20 | 1.9567e-19 | 1.3361e-12
99.6 | 8.92e-20 | 2.1747e-19 | 1.7683e-12
99.7 | 3.4e-20 | 1.6227e-19 | 6.7402e-13
99.8 | 1.2e-19 | 2.4827e-19 | 2.3789e-12
99.9 | 1.17e-19 | 2.4527e-19 | 2.3194e-12
100 | 1.66e-19 | 2.9427e-19 | 3.2908e-12

## VBOC+Bake+324+RGA+Oven+Leak
2010-07-14 | 2010 02:54:44 PM
Residual Gas Analyzer Software
Software Version | 3.205.003
4 | 7.89e-17
4.1 | 7.02e-17
4.2 | 6.41e-17
4.3 | 4.19e-17
4.4 | 1.31e-17
4.5 | 1.68e-18
4.6 | 5.69e-19
4.7 | 3.14e-19
4.8 | 2.12e-19
4.9 | 1.88e-19
5 | 2.98e-19
5.1 | 2.19e-19
5.2 | 2.6e-19
5.3 | 2.43e-19
5.4 | 1.36e-19
5.5 | 9.31e-20
5.6 | 2.33e-19
5.7 | 2.5e-19
5.8 | 2.76e-19
5.9 | 4.33e-19
6 | 4.02e-19
6.1 | 3.45e-19
6.2 | 2.94e-19
6.3 | 3.67e-19
6.4 | 2.51e-19
6.5 | 2.51e-19
6.6 | 3.37e-19
6.7 | 4.64e-19
6.8 | 3.13e-19
6.9 | 4.45e-19
7 | 2.22e-19
7.1 | 1.6e-19
7.2 | 2.37e-19
7.3 | 1.07e-19
7.4 | 2.16e-19
7.5 | 3.19e-19
7.6 | 3.86e-19
7.7 | 4.2e-19
7.8 | 6.28e-19
7.9 | 6.58e-19
8 | 6.17e-19
8.1 | 4.87e-19
8.2 | 6.19e-19
8.3 | 5.62e-19
8.4 | 4.06e-19
8.5 | 4.4e-19
8.6 | 2.26e-19
8.7 | 2.15e-19
8.8 | 2.61e-19
8.9 | 2.08e-19
9 | 2.62e-19
9.1 | 3.12e-19
9.2 | 2.2e-19
9.3 | 1.78e-19
9.4 | 2.64e-19
9.5 | 3.39e-19
9.6 | 3.71e-19
9.7 | 2.84e-19
9.8 | 3.95e-19
9.9 | 3.42e-19
10 | 3.83e-19
10.1 | 4.27e-19
10.2 | 4.22e-19
10.3 | 4.79e-19
10.4 | 5.94e-19
10.5 | 8.92e-19
10.6 | 9.03e-19
10.7 | 1.08e-18
10.8 | 1.43e-18
10.9 | 2.59e-18
11 | 4.73e-18
11.1 | 9.76e-18
11.2 | 1.85e-17
11.3 | 5.38e-17
11.4 | 1.77e-16
11.5 | 4.3e-16
11.6 | 5.93e-16
11.7 | 7.21e-16
11.8 | 8.52e-16
11.9 | 9.4e-16
12 | 9.43e-16
12.1 | 8.2e-16
12.2 | 6.33e-16
12.3 | 3.85e-16
12.4 | 3.18e-17
12.5 | 6.66e-18
12.6 | 6.87e-18
12.7 | 8e-18
12.8 | 9.65e-18
12.9 | 1.11e-17
13 | 1.21e-17
13.1 | 1.2e-17
13.2 | 1.32e-17
13.3 | 1.72e-17
13.4 | 4.16e-17
13.5 | 1.08e-16
13.6 | 1.51e-16
13.7 | 1.76e-16
13.8 | 2.03e-16
13.9 | 2.19e-16
14 | 2.21e-16
14.1 | 1.92e-16
14.2 | 1.57e-16
14.3 | 1.07e-16
14.4 | 1.46e-17
14.5 | 8.69e-18
14.6 | 1.15e-17
14.7 | 1.59e-17
14.8 | 1.9e-17
14.9 | 2.61e-17
15 | 3.9e-17
15.1 | 6.54e-17
15.2 | 1.31e-16
15.3 | 3.2e-16
15.4 | 9.98e-16
15.5 | 2.45e-15
15.6 | 3.13e-15
15.7 | 3.61e-15
15.8 | 4e-15
15.9 | 4.27e-15
16 | 4.31e-15
16.1 | 3.97e-15
16.2 | 3.67e-15
16.3 | 3.8e-15
16.4 | 6.58e-15
16.5 | 1.7e-14
16.6 | 2.24e-14
16.7 | 2.61e-14
16.8 | 2.88e-14
16.9 | 3.02e-14
17 | 3.04e-14
17.1 | 2.78e-14
17.2 | 2.5e-14
17.3 | 2.23e-14
17.4 | 2.25e-14
17.5 | 5.81e-14
17.6 | 8.05e-14
17.7 | 9.45e-14
17.8 | 1.06e-13
17.9 | 1.08e-13
18 | 1.1e-13
18.1 | 9.78e-14
18.2 | 8.33e-14
18.3 | 5.86e-14
18.4 | 3.7e-15
18.5 | 3.33e-16
18.6 | 2e-16
18.7 | 2.03e-16
18.8 | 2.24e-16
18.9 | 2.4e-16
19 | 2.48e-16
19.1 | 2.25e-16
19.2 | 1.97e-16
19.3 | 1.65e-16
19.4 | 1.49e-16
19.5 | 3.21e-16
19.6 | 4.1e-16
19.7 | 5.04e-16
19.8 | 5.55e-16
19.9 | 5.71e-16
20 | 5.73e-16
20.1 | 5.02e-16
20.2 | 4.19e-16
20.3 | 2.61e-16
20.4 | 1.87e-17
20.5 | 2.03e-18
20.6 | 5.06e-19
20.7 | 4.27e-19
20.8 | 3.44e-19
20.9 | 3.67e-19
21 | 7.21e-19
21.1 | 1.45e-18
21.2 | 2.99e-18
21.3 | 7.8e-18
21.4 | 2.62e-17
21.5 | 6.22e-17
21.6 | 7.96e-17
21.7 | 9.36e-17
21.8 | 1.03e-16
21.9 | 1.04e-16
22 | 1.05e-16
22.1 | 9.48e-17
22.2 | 7.92e-17
22.3 | 5.08e-17
22.4 | 5.59e-18
22.5 | 1.92e-18
22.6 | 1.87e-18
22.7 | 1.79e-18
22.8 | 1.84e-18
22.9 | 1.33e-18
23 | 1.12e-18
23.1 | 9.89e-19
23.2 | 1.02e-18
23.3 | 8.45e-19
23.4 | 4.32e-19
23.5 | 1.82e-19
23.6 | 2.46e-19
23.7 | 4.45e-19
23.8 | 3.91e-19
23.9 | 3.23e-19
24 | 3e-19
24.1 | 3.79e-19
24.2 | 3.46e-19
24.3 | 3.42e-19
24.4 | 3.47e-19
24.5 | 3.68e-19
24.6 | 4.68e-19
24.7 | 5.65e-19
24.8 | 7.97e-19
24.9 | 8.76e-19
25 | 8.23e-19
25.1 | 7.96e-19
25.2 | 7.55e-19
25.3 | 8.98e-19
25.4 | 9.1e-19
25.5 | 1.6e-18
25.6 | 2.26e-18
25.7 | 2.88e-18
25.8 | 3.65e-18
25.9 | 3.67e-18
26 | 3.75e-18
26.1 | 3.71e-18
26.2 | 3.63e-18
26.3 | 3.76e-18
26.4 | 3.59e-18
26.5 | 4.98e-18
26.6 | 8.06e-18
26.7 | 1.2e-17
26.8 | 1.94e-17
26.9 | 3.69e-17
27 | 6.26e-17
27.1 | 1.18e-16
27.2 | 2.41e-16
27.3 | 5.72e-16
27.4 | 1.6e-15
27.5 | 3.43e-15
27.6 | 4.32e-15
27.7 | 5.16e-15
27.8 | 5.57e-15
27.9 | 5.67e-15
28 | 5.67e-15
28.1 | 5.12e-15
28.2 | 4.27e-15
28.3 | 2.8e-15
28.4 | 1.81e-16
28.5 | 3.76e-17
28.6 | 3.97e-17
28.7 | 4.64e-17
28.8 | 5.29e-17
28.9 | 5.33e-17
29 | 5.78e-17
29.1 | 5.19e-17
29.2 | 4.48e-17
29.3 | 3e-17
29.4 | 6.17e-18
29.5 | 4.95e-18
29.6 | 7.32e-18
29.7 | 8.97e-18
29.8 | 1.01e-17
29.9 | 1.07e-17
30 | 1.05e-17
30.1 | 9.7e-18
30.2 | 9.33e-18
30.3 | 6.93e-18
30.4 | 1.99e-18
30.5 | 1.26e-18
30.6 | 1.98e-18
30.7 | 2.72e-18
30.8 | 3.79e-18
30.9 | 4.01e-18
31 | 4.28e-18
31.1 | 5.52e-18
31.2 | 8.01e-18
31.3 | 1.42e-17
31.4 | 3.56e-17
31.5 | 7.68e-17
31.6 | 9.86e-17
31.7 | 1.19e-16
31.8 | 1.28e-16
31.9 | 1.31e-16
32 | 1.29e-16
32.1 | 1.2e-16
32.2 | 9.49e-17
32.3 | 6.02e-17
32.4 | 5.81e-18
32.5 | 7.84e-19
32.6 | 3.08e-19
32.7 | 1.09e-19
32.8 | 1.48e-19
32.9 | 2.35e-19
33 | 2.69e-19
33.1 | 3.09e-19
33.2 | 2.08e-19
33.3 | 3.89e-19
33.4 | 2.63e-19
33.5 | 5.29e-19
33.6 | 7.4e-19
33.7 | 7.79e-19
33.8 | 9.15e-19
33.9 | 8.68e-19
34 | 8.51e-19
34.1 | 8.02e-19
34.2 | 9e-19
34.3 | 8.73e-19
34.4 | 7.93e-19
34.5 | 1.15e-18
34.6 | 1.24e-18
34.7 | 1.75e-18
34.8 | 2.04e-18
34.9 | 2.02e-18
35 | 2.1e-18
35.1 | 2.22e-18
35.2 | 2.18e-18
35.3 | 2.5e-18
35.4 | 4.39e-18
35.5 | 8.7e-18
35.6 | 1.09e-17
35.7 | 1.42e-17
35.8 | 1.49e-17
35.9 | 1.42e-17
36 | 1.54e-17
36.1 | 1.47e-17
36.2 | 1.13e-17
36.3 | 6.81e-18
36.4 | 1.36e-18
36.5 | 5.49e-19
36.6 | 5.16e-19
36.7 | 7.36e-19
36.8 | 8.06e-19
36.9 | 9.36e-19
37 | 8.77e-19
37.1 | 9.49e-19
37.2 | 8.62e-19
37.3 | 9.57e-19
37.4 | 1.18e-18
37.5 | 1.99e-18
37.6 | 2.62e-18
37.7 | 3.24e-18
37.8 | 3.75e-18
37.9 | 4.21e-18
38 | 3.85e-18
38.1 | 4.09e-18
38.2 | 3.73e-18
38.3 | 3.77e-18
38.4 | 4.36e-18
38.5 | 5.53e-18
38.6 | 9.19e-18
38.7 | 1.63e-17
38.8 | 2.93e-17
38.9 | 4.96e-17
39 | 7.83e-17
39.1 | 1.39e-16
39.2 | 2.68e-16
39.3 | 6.3e-16
39.4 | 1.58e-15
39.5 | 2.54e-15
39.6 | 3.1e-15
39.7 | 3.68e-15
39.8 | 3.88e-15
39.9 | 4.02e-15
40 | 3.97e-15
40.1 | 3.72e-15
40.2 | 2.66e-15
40.3 | 1.12e-15
40.4 | 4.78e-17
40.5 | 6.7e-18
40.6 | 3.74e-18
40.7 | 3.36e-18
40.8 | 3.36e-18
40.9 | 4.24e-18
41 | 4.8e-18
41.1 | 5.23e-18
41.2 | 5.33e-18
41.3 | 5.45e-18
41.4 | 7.26e-18
41.5 | 1e-17
41.6 | 1.17e-17
41.7 | 1.36e-17
41.8 | 1.25e-17
41.9 | 1.34e-17
42 | 1.43e-17
42.1 | 1.37e-17
42.2 | 1.32e-17
42.3 | 1.19e-17
42.4 | 1.12e-17
42.5 | 1.84e-17
42.6 | 3.11e-17
42.7 | 5.35e-17
42.8 | 8.63e-17
42.9 | 1.29e-16
43 | 2.06e-16
43.1 | 3.43e-16
43.2 | 6.29e-16
43.3 | 1.37e-15
43.4 | 3.44e-15
43.5 | 6.36e-15
43.6 | 8.02e-15
43.7 | 9.84e-15
43.8 | 1.08e-14
43.9 | 1.12e-14
44 | 1.15e-14
44.1 | 1.11e-14
44.2 | 8.41e-15
44.3 | 4.52e-15
44.4 | 3.76e-16
44.5 | 7.51e-17
44.6 | 8.01e-17
44.7 | 1.01e-16
44.8 | 1.15e-16
44.9 | 1.19e-16
45 | 1.24e-16
45.1 | 1.23e-16
45.2 | 9.75e-17
45.3 | 6e-17
45.4 | 1.71e-17
45.5 | 2.03e-17
45.6 | 2.68e-17
45.7 | 3.51e-17
45.8 | 3.92e-17
45.9 | 4.14e-17
46 | 4.35e-17
46.1 | 4.4e-17
46.2 | 3.42e-17
46.3 | 2.02e-17
46.4 | 4.56e-18
46.5 | 7.13e-19
46.6 | 4.52e-19
46.7 | 3.97e-19
46.8 | 4.13e-19
46.9 | 5.66e-19
47 | 5.95e-19
47.1 | 5.61e-19
47.2 | 5.58e-19
47.3 | 3.32e-19
47.4 | 2.9e-19
47.5 | 2.93e-19
47.6 | 2.1e-19
47.7 | 3.59e-19
47.8 | 4.05e-19
47.9 | 3.47e-19
48 | 3.93e-19
48.1 | 3.67e-19
48.2 | 2.29e-19
48.3 | 1.98e-19
48.4 | 2.48e-19
48.5 | 1.23e-19
48.6 | 2.39e-20
48.7 | 3.7e-19
48.8 | 2.87e-19
48.9 | 3.06e-19
49 | 2.83e-19
49.1 | 2.09e-19
49.2 | 1.75e-19
49.3 | 3.2e-19
49.4 | 1.68e-19
49.5 | 2.4e-19
49.6 | 3.24e-19
49.7 | 2.25e-19
49.8 | 3.42e-19
49.9 | 6.09e-19
50 | 5.12e-19
50.1 | 5.18e-19
50.2 | 5e-19
50.3 | 4.89e-19
50.4 | 2.98e-19
50.5 | 4e-19
50.6 | 2.92e-19
50.7 | 4.18e-19
50.8 | 6.01e-19
50.9 | 7.4e-19
51 | 6.92e-19
51.1 | 5.79e-19
51.2 | 6.13e-19
51.3 | 4.63e-19
51.4 | 4.1e-19
51.5 | 3.83e-19
51.6 | 4.1e-19
51.7 | 3.88e-19
51.8 | 3.7e-19
51.9 | 5.86e-19
52 | 4.28e-19
52.1 | 4.32e-19
52.2 | 3.29e-19
52.3 | 3.82e-19
52.4 | 3.21e-19
52.5 | 2.45e-19
52.6 | 1.09e-19
52.7 | 1.3e-19
52.8 | 8.27e-20
52.9 | 1.03e-19
53 | 2.19e-19
53.1 | 1.4e-19
53.2 | 1.75e-19
53.3 | 3.59e-19
53.4 | 1.7e-19
53.5 | 2.46e-19
53.6 | 1.92e-19
53.7 | 2.21e-19
53.8 | 3.89e-19
53.9 | 3.38e-19
54 | 1.76e-19
54.1 | 1.96e-19
54.2 | 1.28e-19
54.3 | 1.29e-19
54.4 | 2.79e-19
54.5 | 2.78e-19
54.6 | 2.77e-19
54.7 | 2.12e-19
54.8 | 2.44e-19
54.9 | 2.81e-19
55 | 2.2e-19
55.1 | 2.65e-19
55.2 | 2.07e-19
55.3 | 1.93e-19
55.4 | 1.63e-19
55.5 | 1.63e-19
55.6 | 1.65e-19
55.7 | 2.1e-19
55.8 | 2.77e-19
55.9 | 3.07e-19
56 | 2.6e-19
56.1 | 3.07e-19
56.2 | 2.21e-19
56.3 | 1.26e-19
56.4 | 2.06e-19
56.5 | 1.65e-19
56.6 | 1.32e-19
56.7 | 1.75e-19
56.8 | 2.87e-20
56.9 | 2.81e-19
57 | 1.79e-19
57.1 | 1.03e-19
57.2 | 5.61e-20
57.3 | 1.02e-19
57.4 | 2.22e-19
57.5 | 2.66e-19
57.6 | 2.01e-19
57.7 | 2.66e-19
57.8 | 2.24e-19
57.9 | 2.02e-19
58 | 4.71e-19
58.1 | 3.46e-19
58.2 | 2.47e-19
58.3 | 2.47e-19
58.4 | 1.98e-19
58.5 | 1.86e-19
58.6 | 2.22e-19
58.7 | 2.1e-19
58.8 | 1.29e-19
58.9 | 1.27e-19
59 | 1.36e-19
59.1 | 1.33e-19
59.2 | 3.05e-19
59.3 | 2.41e-19
59.4 | 3.26e-19
59.5 | 4.11e-19
59.6 | 5.4e-19
59.7 | 7.8e-19
59.8 | 9.32e-19
59.9 | 8.99e-19
60 | 8.82e-19
60.1 | 7.52e-19
60.2 | 7.15e-19
60.3 | 4.07e-19
60.4 | 1.4e-19
60.5 | 1.84e-19
60.6 | 2.36e-19
60.7 | 3.11e-19
60.8 | 1.3e-19
60.9 | 6.27e-20
61 | 2.48e-20
61.1 | 5.25e-20
61.2 | 2.72e-20
61.3 | 1.38e-19
61.4 | 1.28e-19
61.5 | 1.87e-19
61.6 | 1.32e-19
61.7 | 2.41e-19
61.8 | 2.1e-19
61.9 | 2.75e-19
62 | 4.43e-19
62.1 | 2.47e-19
62.2 | 2.08e-19
62.3 | 2.46e-19
62.4 | 2.75e-19
62.5 | 1.56e-19
62.6 | 1.07e-19
62.7 | 2.84e-19
62.8 | 3.69e-19
62.9 | 3.39e-19
63 | 3.72e-19
63.1 | 4.38e-19
63.2 | 3.41e-19
63.3 | 3.38e-19
63.4 | 4.49e-19
63.5 | 7.11e-19
63.6 | 9.14e-19
63.7 | 1.23e-18
63.8 | 2.32e-18
63.9 | 4.07e-18
64 | 6.04e-18
64.1 | 7.57e-18
64.2 | 8.7e-18
64.3 | 1e-17
64.4 | 1.06e-17
64.5 | 1.22e-17
64.6 | 1.27e-17
64.7 | 1.01e-17
64.8 | 6.41e-18
64.9 | 5.43e-18
65 | 6.19e-18
65.1 | 7.33e-18
65.2 | 8.52e-18
65.3 | 9.53e-18
65.4 | 1.11e-17
65.5 | 1.33e-17
65.6 | 1.47e-17
65.7 | 1.48e-17
65.8 | 1.17e-17
65.9 | 1.1e-17
66 | 1.02e-17
66.1 | 1.05e-17
66.2 | 8.54e-18
66.3 | 4.09e-18
66.4 | 2.31e-18
66.5 | 2.13e-18
66.6 | 2.52e-18
66.7 | 3.13e-18
66.8 | 3.79e-18
66.9 | 4.03e-18
67 | 4.55e-18
67.1 | 4.34e-18
67.2 | 3.68e-18
67.3 | 1.99e-18
67.4 | 1.56e-18
67.5 | 1.89e-18
67.6 | 2.05e-18
67.7 | 2.86e-18
67.8 | 3.35e-18
67.9 | 3.45e-18
68 | 3.78e-18
68.1 | 3.75e-18
68.2 | 3.2e-18
68.3 | 1.95e-18
68.4 | 6.39e-19
68.5 | 4.18e-19
68.6 | 2.7e-19
68.7 | 2.93e-19
68.8 | 2.35e-19
68.9 | 1.59e-19
69 | 6.45e-20
69.1 | 8.45e-20
69.2 | 2.38e-19
69.3 | 2.97e-19
69.4 | 2.14e-19
69.5 | 1.6e-19
69.6 | 2.54e-19
69.7 | 3.32e-19
69.8 | 6.12e-20
69.9 | 1.03e-19
70 | 2.62e-19
70.1 | 1.27e-19
70.2 | 1.59e-19
70.3 | 1.38e-19
70.4 | 2.09e-19
70.5 | 1.79e-19
70.6 | 1.53e-19
70.7 | 2.45e-19
70.8 | 1.32e-19
70.9 | 8.83e-20
71 | 2.61e-19
71.1 | 1.71e-19
71.2 | 2.34e-19
71.3 | 2.3e-19
71.4 | 1.6e-19
71.5 | 8.09e-20
71.6 | 1.65e-19
71.7 | 2.2e-19
71.8 | 2.45e-19
71.9 | 1.96e-19
72 | 1.66e-19
72.1 | 1.84e-19
72.2 | 2.35e-19
72.3 | 2.72e-19
72.4 | 3.04e-19
72.5 | 1.85e-19
72.6 | 1.33e-19
72.7 | 2.55e-19
72.8 | 3.58e-19
72.9 | 2.92e-19
73 | 2.36e-19
73.1 | 1.25e-19
73.2 | 2.31e-19
73.3 | 3.8e-19
73.4 | 4.12e-19
73.5 | 3.78e-19
73.6 | 2.74e-19
73.7 | 2.78e-19
73.8 | 3.23e-19
73.9 | 2.44e-19
74 | 2.57e-19
74.1 | 2.98e-19
74.2 | 1.76e-19
74.3 | 2.8e-19
74.4 | 1.27e-19
74.5 | 2.84e-19
74.6 | 2.9e-19
74.7 | 1.78e-19
74.8 | 2.84e-19
74.9 | 3.45e-19
75 | 1.99e-19
75.1 | 3.31e-19
75.2 | 1.9e-19
75.3 | 2.46e-19
75.4 | 1.92e-19
75.5 | 3.07e-19
75.6 | 3.13e-19
75.7 | 1.98e-19
75.8 | 8.71e-20
75.9 | 3.19e-19
76 | 3.08e-19
76.1 | 2.51e-19
76.2 | 3.5e-19
76.3 | 3.35e-19
76.4 | 1.85e-19
76.5 | 2.4e-19
76.6 | 2.41e-19
76.7 | 3.24e-19
76.8 | 2.74e-19
76.9 | 4.41e-19
77 | 4.44e-19
77.1 | 3.76e-19
77.2 | 4.18e-19
77.3 | 5.61e-19
77.4 | 7.07e-19
77.5 | 6.35e-19
77.6 | 1.11e-18
77.7 | 1.28e-18
77.8 | 1.45e-18
77.9 | 1.79e-18
78 | 1.75e-18
78.1 | 1.85e-18
78.2 | 1.73e-18
78.3 | 1.51e-18
78.4 | 8.03e-19
78.5 | 4.24e-19
78.6 | 2.49e-19
78.7 | 2.2e-19
78.8 | 1.06e-19
78.9 | 2.15e-19
79 | 2.4e-19
79.1 | 2.32e-19
79.2 | 5.57e-19
79.3 | 6.4e-19
79.4 | 1.1e-18
79.5 | 1.67e-18
79.6 | 1.84e-18
79.7 | 2.67e-18
79.8 | 3.11e-18
79.9 | 3.55e-18
80 | 3.55e-18
80.1 | 3.52e-18
80.2 | 2.88e-18
80.3 | 1.32e-18
80.4 | 4.84e-19
80.5 | 1.98e-19
80.6 | 2.3e-19
80.7 | 2.42e-19
80.8 | 3.35e-19
80.9 | 4.82e-19
81 | 6.79e-19
81.1 | 6.88e-19
81.2 | 1.26e-18
81.3 | 2.47e-18
81.4 | 5.01e-18
81.5 | 7.84e-18
81.6 | 9.84e-18
81.7 | 1.25e-17
81.8 | 1.49e-17
81.9 | 1.63e-17
82 | 1.74e-17
82.1 | 1.76e-17
82.2 | 1.29e-17
82.3 | 6e-18
82.4 | 6e-18
82.5 | 8.15e-18
82.6 | 9.47e-18
82.7 | 1.36e-17
82.8 | 1.46e-17
82.9 | 1.75e-17
83 | 1.83e-17
83.1 | 1.95e-17
83.2 | 1.63e-17
83.3 | 1.51e-17
83.4 | 2.49e-17
83.5 | 3.56e-17
83.6 | 4.8e-17
83.7 | 6.29e-17
83.8 | 7.06e-17
83.9 | 7.9e-17
84 | 8.1e-17
84.1 | 7.96e-17
84.2 | 5.18e-17
84.3 | 1.37e-17
84.4 | 3.26e-18
84.5 | 1.18e-18
84.6 | 5.45e-19
84.7 | 2.41e-19
84.8 | 4.49e-19
84.9 | 5.46e-19
85 | 8.86e-19
85.1 | 7.84e-19
85.2 | 1.3e-18
85.3 | 3.03e-18
85.4 | 6.95e-18
85.5 | 1.06e-17
85.6 | 1.41e-17
85.7 | 1.82e-17
85.8 | 2.06e-17
85.9 | 2.34e-17
86 | 2.45e-17
86.1 | 2.3e-17
86.2 | 1.65e-17
86.3 | 5.39e-18
86.4 | 1.16e-18
86.5 | 5.2e-19
86.6 | 2.67e-19
86.7 | 2.43e-19
86.8 | 1.44e-19
86.9 | 1.55e-19
87 | 3.76e-20
87.1 | 9.16e-20
87.2 | 2.25e-19
87.3 | 3.13e-19
87.4 | 2.37e-19
87.5 | 1.15e-19
87.6 | 2.25e-19
87.7 | 1.57e-19
87.8 | 2.43e-19
87.9 | 4.08e-19
88 | 3.75e-19
88.1 | 3.91e-19
88.2 | 1.94e-19
88.3 | 1.68e-19
88.4 | 2.62e-19
88.5 | 2.96e-19
88.6 | 3.18e-19
88.7 | 1.15e-19
88.8 | 1.13e-19
88.9 | 2.01e-19
89 | 1.93e-19
89.1 | 5.34e-20
89.2 | 5.49e-20
89.3 | 2.47e-19
89.4 | 3.34e-19
89.5 | 3.04e-19
89.6 | 1.8e-19
89.7 | 1.24e-19
89.8 | 2.04e-19
89.9 | 5.58e-20
90 | 1.96e-19
90.1 | 2.8e-19
90.2 | 1.89e-19
90.3 | 2.07e-19
90.4 | 3.17e-19
90.5 | 2.07e-19
90.6 | 2.78e-19
90.7 | 2.86e-19
90.8 | 2.77e-19
90.9 | 2.74e-19
91 | 1.48e-19
91.1 | 2.06e-19
91.2 | 2.08e-19
91.3 | 3.77e-19
91.4 | 2.19e-19
91.5 | 1.14e-19
91.6 | 3.46e-19
91.7 | 2.62e-19
91.8 | 2.77e-19
91.9 | 2.03e-19
92 | 2.85e-19
92.1 | 2.19e-19
92.2 | 1.77e-19
92.3 | 2.02e-19
92.4 | 2.34e-19
92.5 | 2.29e-19
92.6 | 2.6e-19
92.7 | 2.65e-19
92.8 | 2.12e-19
92.9 | 2.41e-19
93 | 1.92e-19
93.1 | 1.87e-19
93.2 | 1.14e-19
93.3 | 6.54e-20
93.4 | 1.59e-19
93.5 | 4.18e-19
93.6 | 3.36e-19
93.7 | 2.48e-19
93.8 | 2.4e-19
93.9 | 2.95e-19
94 | 2.4e-19
94.1 | 1.71e-19
94.2 | 2.58e-19
94.3 | 1.55e-19
94.4 | 1.81e-19
94.5 | 1.87e-19
94.6 | 9.67e-20
94.7 | 2.32e-19
94.8 | 3.07e-19
94.9 | 3.36e-19
95 | 1.81e-19
95.1 | 2.08e-19
95.2 | 2.69e-19
95.3 | 3.12e-19
95.4 | 2.68e-19
95.5 | 1.66e-19
95.6 | 2.36e-19
95.7 | 1.91e-19
95.8 | 3.75e-19
95.9 | 3.35e-19
96 | 1.93e-19
96.1 | 2.13e-19
96.2 | 3.19e-19
96.3 | 2.67e-19
96.4 | 1.22e-19
96.5 | 2.41e-19
96.6 | 3.8e-19
96.7 | 4.73e-19
96.8 | 3.04e-19
96.9 | 2.48e-19
97 | 3.22e-19
97.1 | 1.85e-19
97.2 | 1.84e-19
97.3 | 1.02e-19
97.4 | 1.16e-19
97.5 | 1.88e-19
97.6 | 2.16e-19
97.7 | 1.85e-19
97.8 | 3.22e-19
97.9 | 2.47e-19
98 | 3.11e-19
98.1 | 3.23e-19
98.2 | 2.38e-19
98.3 | 3.51e-19
98.4 | 2.58e-19
98.5 | 2.5e-19
98.6 | 3.22e-19
98.7 | 1.81e-19
98.8 | 3.09e-19
98.9 | 2.01e-19
99 | 1.29e-19
99.1 | 1.83e-19
99.2 | 2.42e-19
99.3 | 2.99e-19
99.4 | 3.52e-19
99.5 | 2.65e-19
99.6 | 2.57e-19
99.7 | 2.37e-19
99.8 | 3.03e-19
99.9 | 1.64e-19
100 | 1.84e-19

## Cal
VBOC Calibration Gas
Supplier=Vacuum Technologies Inc.
Model=CL-7-StdMix-MCFF-300DOT-MFV-SR15
Serial #=013463
Calibration gas=3%Kr/Xe/He/H2, 1.8%N2 in Argon
LR=9.12x10-8 torr*L/sec into Vacuum
Depletion Rate=1.7%/yr
Temp Coeff.=0.2%/deg C
Calibration Date=Aug. 7, 2009
Fill Pressure=570 Torr
86.2% | Argon (cracking pattern peak at AMU 40)
7.8614e-08 | Torr-Liter/sec Argon
5.44e-17 | RGA at AMU 40 without leak (A)
4.02e-15 | RGA at AMU 40 with leak (A)
3.9656e-15 | Increase at AMU 40 due to Cal leak gas (A)
19824087.1495 | Torr-L/s/A calibration
7 | qty of A-OSEMS in oven
2832012.4499 | Leak_Rate_Cal_factor_per_AOSEM | Torr-L/s/A/AOSEM