| AUTHOR(S) | DATE | Document Change Notice, Release or Approval |
| --- | --- | --- |
| Kiwamu Izumi | 28/9/2013 | see LIGO DCC record Status |

- Introduction

This document covers reflectivity measurement performed in the lab with a 90% P-polarizing beam splitter (E040512-B3) and a 1064 nm free space laser at S-polarization.

- Background

To attenuate the interferometer reflected light we were seeking a beam splitter, whose reflectivity is in a range of 90-95 %, in our stock. One candidate is this 90% beam splitter (E040512-B3) although this is designed for P-polarization. Because the reflected light is S-polarizing we needed to measure the reflectivity of this beam splitter for S-polarization.

3. Set up

Figure.1 Measurement setup

As for the laser source we used the existing Lightwave NPRO which had been setup for a reference cavity and SURF squeezer experiment at the Hanford OSB optics lab. A polarizing beam splitter (PBS), which was already setup, was used to generate S-polarizing beams as shown in figure 1.

Items used :

- 1064 nm laser source (which had been already setup)

- Polarizing cube beam splitter, 1064nm (which had been already setup)

- Half wave plate, 1064 nm (which had been already setup)

- ND01 neutral density filter

- Steering mirror, 1064 nm

- Ophir laser power meter, Vega , with filter removed

- Measurements

Below are the results of the measurements. We flipped the beam splitter so that we obtain the reflectivity for both high reflective and anti-reflective surface.

HR surface R = 92.5 % for S-pol at 1064 nm, 45 deg

AR surface R = 2.8 % for S-pol at 1064 nm, 45 deg

- Links

- Beam splitter specification : E040512-x0

https://dcc.ligo.org/DocDB/0023/E040512/000/E040512-01.pdf

- LHO alog 7879

https://alog.ligo-wa.caltech.edu/aLOG/index.php?callRep=7879