# LIGO-T040068-00-V

# Checklist for Opening Calibration Gas into RGA Volume

- Obtain approved work permit

- Notify operator on duty of activity on day of activity and verify GV6 and GV8 “AUTOCLOSE” is enabled if opened

- Isolate RGA volume from VE

- Reduce RGA filament current to minimum value

- Connect auxilary turbo pump cart to RGA volume pump port

- Start pump cart and verify turbo at full speed

- Open RGA Calibration Gas into RGA volume while monitoring ion current.  If pressure remains lower than ~10(-4) torr it is not necessary to open the RGA to the pump cart. Proceed to step 9.

- Open RGA volume to pump cart

- Setup RGA in MID mode using AMUs 14 and 28 and sampling every 3 seconds

- Start scanning and monitor values of AMUs 14 and 28

- Isolate pump cart from RGA volume

- Monitor rate of rise of AMU 14 and 28

- Stop pump cart and confirm turbo inlet vented

- Is rate of rise the same as in step 12?

- If yes, disconnect pump cart and blank off RGA volume’s pump port

- Open valve connecting RGA volume to VE

- Restore RGA setup to the “as found” state