LIGO Laboratory / LIGO Scientific Collaboration

LIGO-T060153-00-E	Advanced LIGO	6/27/06

Degenerate RC Core Optics Position & Orientation

(Some revisions to T010076)

Michael Smith

Distribution of this document:

LIGO Science Collaboration

This is an internal working note

of the LIGO Project.

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW17-161 175 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 1970 Mail Stop S9-02 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

http://www.ligo.caltech.edu/

Table of Contents

1	Introduction	4

1.1	Background	4

1.2	Scope	4

1.3	Referenced Documents	4

2	Non-folded IFO	5

2.1	RC Optics Position & Orientation, Degenerate Cavity	5

2.2	RC Core Optics Ray Parameters, Degenerate Cavity	11

2.3	RC Core Optics Wedge Angles, Degenerate Cavity	12

3	Folded IFO	12

3.1	RC Optics Position & Orientation, Degenerate Cavity	12

3.2	RC Core Optics Ray Parameters, Degenerate Cavity	12

3.3	RC Core Optics Wedge Angles, Degenerate Cavity	12

Abstract

The positions and orientations of the Core Optical components for Advanced LIGO non-folded and folded degenerate recycling cavity interferometers have been determined using ZEMAX optical design program. For the non-folded IFO, a symmetric wedge angle of 2.0000 deg was placed on the BS optic in order to achieve the required clearance between the closest extent of the BS PO mirror assembly and the 100ppm radius of the main interferometer beam within the recycling cavity.

# Introduction

The positions and orientations of the Core Optical components for Advanced LIGO degenerate recycling cavity have been determined using ZEMAX optical design program.

## Background

Coyne has previously established the positions and orientations of the optical elements in the optical train of the Advanced LIGO interferometer. The driving parameter that determines the required wedge angle on the BS optic is the clearance of the BS PO beam above the main beam at the location of the BS PO mirror on BSC3 chamber. A symmetric wedge angle of 2.0000 deg was placed on the BS optic In order to achieve a required clearance > 30mm measured between the closest extent of the BS PO mirror assembly and the 100ppm radius of the main interferometer beam within the recycling cavity, which is 128.8 mm. The prior BS wedge angle assumed by Coyne was 1.3000 deg.

## Scope

The wedge angles, positions, and orientations of the optical elements listed in this document are based on the following assumptions:

Degenerate recycling cavity

Minimum clearance between BS PO assembly and main RC beam @ 100 ppm radius	30 mm

Global height of x-arm beam									80 mm

Wedge of compensation plate							                       negligible

Height of MMT3 and PRM 								 nominally equal

Length of PRC										8341 mm

Length of SRC										8327 mm

Schnoop asymmetry										400 mm

## Referenced Documents

T010076-01 Optical Layout for Advanced LIGO, D. Coyne

# Non-folded IFO

## RC Optics Position & Orientation, Degenerate Cavity

| Table 5: RC Core Optics Position & Orientation, Degenerate Cavity |
| --- |
|  |  |  | CENTERLINE COORDINATE | SURFACE ORIENTATION |
| CHAMBER | Optic | SURFACE | Pt | X (mm) | Y (mm) | Z (mm) | n | i | j | k |
| non-fold #1 |  |  |  |  |  |  |  |  |  |  |
| HAM1 | DLC2 OPTLEV SM1 | HR | p18 | -22691.9 | -77.8 | -100.7 | n18 | 0.769400 | -0.638768 | 0.000000 |
| HAM2 | DLC1 | HR | p01 | -20356.9 | -413.5 | -100.2 | n01 | 0.707042 | 0.707172 | 0.000000 |
|  | DLC2 | HR | p02 | -20356.8 | -905.1 | -100.2 | n02 | 0.707107 | 0.707107 | 0.000000 |
|  | DLC3 | HR | p03 | -20085.1 | -905.1 | -100.2 | n03 | 0.693398 | -0.720555 | 0.000061 |
|  | MC1 | AR | p04 | -20079.1 | -749.3 | -100.2 | n04 | -0.685864 | -0.727730 | -0.000436 |
|  | MC1 | HR | p05 | -20055.4 | -668.5 | -100.2 | n05 | -0.710844 | -0.703350 | -0.000436 |
|  | MC3 | HR | p06 | -20050.0 | -219.3 | -100.2 | n06 | 0.693590 | -0.720369 | 0.000436 |
|  | MC3 | AR | p07 | -20071.7 | -138.0 | -100.2 | n07 | 0.668027 | -0.744137 | 0.000436 |
|  | SM1 | HR | p08 | -20084.3 | 764.4 | -100.3 | n08 | -0.378487 | 0.925607 | 0.000000 |
|  | SM2 | HR | p09 | -19645.3 | 304.8 | -100.4 | n09 | 0.870402 | -0.492342 | 0.000000 |
|  | MMT1 | HR | p010 | -20740.9 | 551.1 | -100.5 | n010 | -0.994461 | 0.105110 | -0.000367 |
|  | MMT3 | HR | p011 | -19338.7 | 221.5 | -156.5 | n011 | -0.999920 | -0.012430 | -0.002147 |
|  | MMT3 | AR | p012 | -19238.7 | 221.6 | -156.4 | n012 | -0.999991 | -0.003744 | -0.002147 |
|  | DLC1 OPTLEV SM1 | HR | p013 | -19890.5 | 442.4 | -100.7 | n013 | 0.913545 | -0.406737 | 0.000000 |
|  | DLC OPTLEV MC1 | HR | p014 | -20634.8 | -761.7 | -95.0 | n014 | 0.719340 | -0.694658 | 0.000438 |
|  | DLC MCR | HR | p015 | -20559.5 | -680.8 | -100.5 | n015 | 0.850811 | 0.525472 | 0.000000 |
|  | DLC1 OPTLEV MC3 | HR | p016 | -19703.8 | -294.3 | -94.0 | n016 | 0.882948 | -0.469472 | 0.000438 |
|  | DLC2 OPTLEV MC3 | HR | p017 | -19614.3 | -414.3 | -93.8 | n017 | 0.770513 | -0.637424 | 0.000000 |
|  | DLC3 OPTLEV MC3 | HR | p018 | -19904.7 | -512.8 | -92.6 | n018 | -0.543942 | -0.837597 | -0.050593 |
|  | DLC REFL | HR | p019 | -19813.8 | 238.4 | -100.4 | n019 | 0.767165 | -0.641450 | 0.000438 |
|  | DLC1 MON1 | HR | p020 | -19926.2 | 903.5 | -100.4 | n020 | 0.642788 | -0.766044 | 0.000438 |
|  | DLC2 MON1 | HR | p021 | -19635.9 | 797.5 | -100.2 | n021 | 0.972776 | 0.231748 | 0.000438 |
|  | DLC3 MON1 | HR | p022 | -20598.5 | -230.1 | -99.3 | n022 | 0.398749 | -0.917060 | 0.000000 |
|  | DLC1 MON2 | HR | p023 | -20488.9 | 385.1 | -100.5 | n023 | 0.461749 | -0.887011 | 0.000438 |
|  | DLC1 MON3 | HR | p024 | -20233.7 | 755.1 | -100.4 | n024 | 0.887815 | 0.460200 | 0.000438 |
|  | DLC2 MON3 | HR | p025 | -20588.9 | 754.3 | -100.6 | n025 | 0.707107 | -0.707107 | 0.000438 |
|  | DLC3 MON3 | HR | p026 | -20590.3 | 149.4 | -100.7 | n026 | 0.707107 | -0.707107 | 0.000438 |
|  | DLC1 OPTLEV MMT1 | HR | p027 | -19448.3 | 219.7 | -161.9 | n027 | 0.614873 | -0.787002 | -0.050593 |
|  | DLC2 OPTLEV MMT1 | HR | p028 | -19443.6 | 549.3 | -146.5 | n028 | 0.707973 | -0.704030 | -0.055822 |
|  | DLC HEAT MMT2 | HR | p029 | -20480.1 | 656.9 | -91.3 | n029 | 0.620819 | 0.783954 | 0.000070 |
| HAM3 | MC2 | HR | p030 | -3588.5 | -643.0 | -88.4 | n030 | 0.999927 | -0.012098 | 0.000611 |
|  | MMT2 | HR | p031 | -3893.0 | 728.0 | -90.0 | n031 | 0.999762 | 0.021658 | 0.002443 |
|  | PRM1 | AR | p2 | -3809.0 | 219.3 | -156.7 | n2 | -0.999979 | 0.000152 | 0.006418 |
|  | PRM1 | HR | p3 | -3709.0 | 219.3 | -156.9 | n3 | -0.999998 | 0.000152 | 0.002003 |
|  | DLC1 OPTLEV MC2 | HR | p032 | -4289.9 | -716.4 | -100.9 | n032 | 0.999048 | 0.043619 | 0.000438 |
|  | DLC2 OPTLEV MC2 | HR | p033 | -3845.3 | -716.6 | -101.0 | n033 | 0.998365 | -0.054594 | -0.016929 |
|  | DLC1 OPTLEV SM2 | HR | p034 | -19725.6 | 471.3 | -94.2 | n034 | 0.172961 | -0.980909 | 0.088894 |
|  | DLC2 OPTLEV SM2 | HR | p035 | -22382.7 | -74.3 | -98.1 | n035 | 0.804894 | -0.593419 | 0.000000 |
|  | DLC1 OPTLEV MMT2 | HR | p036 | -4382.3 | 769.3 | -110.4 | n036 | 0.698767 | -0.713558 | -0.050593 |
|  | DLC2 OPTLEV MMT2 | HR | p037 | -4376.0 | -500.8 | -241.7 | n037 | 0.717542 | -0.694133 | -0.057564 |
|  | DLC HEAT MMT3 | HR | p038 | -3554.7 | 551.9 | -91.3 | n038 | 0.804894 | -0.593419 | 0.000070 |
|  | ITMY HARTMANN M1 | HR | p039 | -2685.2 | -234.5 | 148.8 | n039 | 0.700534 | -0.695660 | 0.159089 |
|  | ITMY HARTMANN M2 | HR | p040 | -2677.6 | 162.1 | 58.1 | n040 | 0.706161 | -0.701542 | 0.095794 |
|  | ITMY HART TELE M1 | HR | p041 | -4550.0 | -248.7 | -101.2 | n041 | -1.000000 | 0.000000 | 0.000000 |
|  | ITMY HART TELE M2 | HR | p042 | -4607.1 | -248.9 | -101.3 | n042 | -0.706922 | -0.707292 | -0.000436 |
|  | ITMY HART STEER 1 | HR | p043 | -4366.4 | 60.1 | -106.6 | n043 | 0.707107 | -0.707107 | 0.000000 |
|  | ITMY HART STEER 2 | HR | p044 | -4708.5 | -252.2 | -101.5 | n044 | -0.865919 | 0.499938 | 0.015707 |
|  | ITMY HART STEER 3 | HR | p045 | -4366.5 | 60.0 | -100.8 | n045 | 0.707107 | -0.707107 | 0.000000 |
|  | BD BSHR3P MIRROR | HR | p046 | -4705.8 | -247.6 | -100.8 | n046 | -0.865919 | 0.499938 | 0.015707 |
| HAM4 | SRM1 | HR | p10 | -180.9 | -3536.0 | -157.8 | n10 | 0.000151 | -0.999998 | 0.001982 |
|  | SRM1 | AR | p11 | -180.9 | -3636.0 | -157.6 | n11 | 0.000152 | -0.999980 | 0.006398 |
|  | OMMT2 | HR | p21 | 204.1 | -3670.0 | -156.3 | n21 | -0.015800 | -0.999874 | 0.001586 |
|  | OMMT3 | HR | p22 | 80.0 | -20600.0 | -101.1 | n22 | 0.047455 | -0.998872 | 0.001693 |
|  | OMMT3 | AR | p23 | 80.5 | -20650.0 | -100.9 | n23 | 0.047455 | -0.998872 | 0.001693 |
|  | BS PO MIRROR2 | HR | p047 | 598.3 | -4316.0 | -101.3 | n047 | -0.043500 | 0.998813 | 0.021900 |
|  | BSPO STEER M1 | HR | p048 | 597.6 | -3214.6 | -99.7 | n048 | 0.707107 | 0.707107 | 0.000000 |
|  | BSPO STEER M2 | HR | p049 | 448.9 | -3214.5 | -99.7 | n049 | 0.703395 | -0.710799 | 0.000000 |
|  | ITMX HARTMANN M1 | HR | p050 | -597.1 | -2695.1 | 148.1 | n050 | -0.000111 | -0.707107 | -0.707107 |
|  | ITMX HARTMANN M2 | HR | p051 | -597.2 | -2695.2 | -416.8 | n051 | 0.707107 | 0.000000 | 0.707107 |
|  | ITMX HARTMANN M3 | HR | p052 | -144.2 | -2695.2 | -416.7 | n052 | 0.705828 | -0.705754 | -0.060979 |
|  | ITMXPO TELE M1 | HR | p053 | -598.5 | -4288.2 | 148.5 | n053 | -0.000098 | 0.996149 | -0.087680 |
|  | ITMXPO TELE M2 | HR | p054 | -598.6 | -2889.3 | -99.8 | n054 | -0.000724 | 0.996188 | -0.087228 |
|  | ITMXPO STEER M1 | HR | p055 | -598.1 | -4514.5 | -100.8 | n055 | 0.706693 | 0.707520 | 0.000000 |
|  | ITMX HART STEER M1 | HR | p056 | -346.8 | -3833.2 | -101.0 | n056 | 0.707107 | -0.707107 | 0.000000 |
|  | ITMX HART STEER M2 | HR | p057 | -345.3 | -4514.3 | -101.0 | n057 | 0.708155 | -0.706057 | -0.000454 |
|  | ITMXPO STEER M2 | HR | p058 | 491.7 | -4516.5 | -101.6 | n058 | 0.712026 | 0.702153 | 0.000000 |
| HAM5 | OMMT1 | HR | p20 | -178.5 | -19430.0 | -157.7 | n20 | 0.012061 | 0.999927 | 0.000047 |
| BSC1 | BS PO MIRROR1 | HR | p059 | -131.5 | 4123.1 | 251.1 | n059 | 0.000000 | 0.000000 | 0.000000 |
|  | CPy | front | p18 | -200.2 | 4610.4 | -86.9 | n18 | 0.000000 | -0.999840 | -0.017889 |
|  | CPy | back | p19 | -200.2 | 4675.4 | -85.8 | n19 | 0.000000 | -0.999813 | -0.019342 |
|  | ITMY | AR | p8 | -200.2 | 4731.1 | -84.8 | n8 | 0.000000 | -0.999298 | 0.037465 |
|  | ITMY | HR | p9 | -200.2 | 4931.2 | -84.8 | n9 | 0.000000 | -1.000000 | 0.000000 |
| BSC2 | BS1 | BS | p4 | -200.2 | 218.8 | -163.9 | n4 | -0.707039 | 0.707039 | 0.013812 |
|  | BS1 | AR-x | p5 | -134.1 | 200.0 | -164.4 | n5 | -0.706268 | 0.706268 | 0.048700 |
|  | BS1 | AR-aps | p5p | -134.1 | 200.0 | -164.4 | n5n | -0.706268 | 0.706268 | 0.048700 |
| BSC3 | CPx | front | p16 | 4558.0 | 200.0 | -82.1 | n16 | -0.999840 | 0.000000 | -0.017889 |
|  | CPx | back | p17 | 4623.0 | 200.0 | -80.9 | n17 | -0.999813 | 0.000000 | -0.019342 |
|  | ITMX | AR | p6 | 4678.7 | 200.0 | -80.0 | n6 | -0.999298 | 0.000000 | 0.037465 |
|  | ITMX | HR | p7 | 4878.8 | 200.0 | -80.0 | n7 | -1.000000 | 0.000000 | 0.000000 |
| BSC9 | ETMX | HR | p12 | 18100.0 | 200.0 | -80.0 | n12 | -1.000000 | 0.000000 | 0.000000 |
|  | ETMX | AR | p13 | 18200.0 | 200.0 | -80.0 | n13 | -0.999962 | 0.000000 | 0.008727 |
| BSC10 | ETMY | HR | p14 | -200.2 | 18000.0 | -84.8 | n14 | 0.000000 | -1.000000 | 0.000000 |
|  | ETMY | AR | p15 | -200.2 | 18200.0 | -84.8 | n15 | 0.000000 | -0.999962 | 0.008727 |

## RC Core Optics Ray Parameters, Degenerate Cavity

| Table 6: RC Core Optics Ray Parameters |
| --- |
|  |  |  |  | RAY UNIT VECTOR |
| CHAMBER | Descripton | Surface | Ray | (direction cosines) |
|  |  |  |  | i | j | k |
| HAM2 | MC1 | AR | u04 | 0.038215 | 0.999270 | -0.000084 |
|  | MC1 | HR | u05 | 0.012137 | 0.999926 | -0.000092 |
|  | MC3 | HR | u06 | 0.012137 | 0.999926 | -0.000092 |
|  | MC3 | AR | u07 | -0.013952 | 0.999903 | -0.000100 |
|  | MMT3 | HR | u011 | -0.999452 | -0.032828 | -0.004309 |
|  | MMT3 | AR | u012 | 1.000000 | -0.000151 | -0.000017 |
| HAM3 | MMT2 | HR | u031 | -0.999562 | -0.029361 | -0.003730 |
|  | PRM1 | AR | u2 | 1.000000 | -0.000151 | -0.000017 |
|  | PRM1 | HR | u3 | 0.999998 | -0.000151 | -0.002004 |
| HAM4 | SRM1 | HR | u10 | 0.000152 | -0.999998 | 0.001983 |
|  | SRM1 | AR | u11 | 0.000152 | -1.000000 | -0.000004 |
| HAM5 | CPy | front | u18 | 0.000000 | 0.999846 | 0.017531 |
|  | CPy | back | u19 | 0.000000 | 0.999858 | 0.016877 |
|  | ITMY | AR | u8 | 0.000000 | 0.999858 | 0.016877 |
|  | ITMY | HR | u9 | 0.000000 | 1.000000 | -0.000001 |
| BSC2 | BS1 | BS | u4 | 0.999998 | -0.000151 | -0.002004 |
|  | BS1 | AR-x | u5 | 0.999846 | 0.000000 | 0.017532 |
|  | BS1 | AR-aps | u5u | 0.000152 | -0.999998 | 0.001983 |
| BSC3 | CPx | front | u16 | 0.999846 | 0.000000 | 0.017532 |
|  | CPx | back | u17 | 0.999858 | 0.000000 | 0.016877 |
|  | ITMX | AR | u6 | 0.999858 | 0.000000 | 0.016877 |
|  | ITMX | HR | u7 | 1.000000 | 0.000000 | 0.000000 |
| BSC9 | ETMX | HR | u12 | 1.000000 | 0.000000 | 0.000000 |
|  | ETMX | AR | u13 | 0.999992 | 0.000000 | 0.003927 |
| BSC10 | ETMY | HR | u14 | 0.000000 | 1.000000 | -0.000001 |
|  | ETMY | AR | u15 | 0.000000 | 0.999992 | 0.003927 |

## RC Core Optics Wedge Angles, Degenerate Cavity

| Table 6: RC Core Optics Wedge Angles |
| --- |
|  |  | WEDGE |
| Interferometer | Optic |  |  |
|  |  | Angle (Deg) | Orientation of thick side |
| Non-folded, unstable (#1) |  |  |  |
|  | MC1 | 2.0000 | away from MC2 |
|  |  | symmetric |  |
|  | MC3 | 2.0000 | away from MC2 |
|  |  | symmetric |  |
|  | MMT3 | 0.5000 | +y direction |
|  |  | symmetric |  |
|  | PRM1 | 0.2530 | down |
|  |  | symmetric |  |
|  | SRM1 | 0.2530 | down |
|  |  | symmetric |  |
|  | BS1 | 2.0000 | up |
|  |  | symmetric |  |
|  | ITM1 | 2.14708 | down |
|  |  | symmetric |  |
|  | CP | 0.0833 | down |
|  |  | single sided |  |
|  | ETM | 0.5000 | up |
|  |  | single sided |  |

# Folded IFO

## RC Optics Position & Orientation, Degenerate Cavity

## RC Core Optics Ray Parameters, Degenerate Cavity

## RC Core Optics Wedge Angles, Degenerate Cavity