## Sheet1
Pressure Contribution from Flag Hydrocarbons
40M Lab RGA Scan Results
F21611 |  | Description: |  | Electro polished test parts | Date: | 2011-02-16
|  |  | Oven Used: | F
AMU 41 | 7.08e-16 | amps | from RGA scan listing
AMU 43 | 4.83e-16 | amps | from RGA scan listing
AMU 53 | 1.74e-16 | amps | from RGA scan listing
AMU 55 | 2.63e-17 | amps | from RGA scan listing
AMU 57 | 1.2e-17 | amps | from RGA scan listing
Sum Flag H/C AMUs | 1.4033e-15 | amps
Calib leak rate | 2.36e-10 | torr l/s | (Argon)
AMU 40 (w/leak open) | 8.14e-13 | amps
AMU 40 (background) | 7.68e-16 | amps
Calib leak contributes | 8.1323e-13 | amps | = (w/leak open) - (background)
Flag H/C Outgassing | 4.0724e-13 | torr l/s | = (Sum Flag H/C AMUs) x (Calib leak rate)/(Calib leak contrib.)
Test item surf area | 135700 | cm^2
Normalized outgassing | 3.001e-18 | torr l/s-cm2 |  | = Flag H/C Outgassing/Test item surf area
Full description: | Assorted aluminum electroplated parts with holes, taped holes and welds
Pre-scan bake: | 120C for 48Hrs.