LASER INTERFEROMETER GRAVITATIONAL WAVE OBSERVATORY

-LIGO-

CALIFORNIA INSTITUTE OF TECHNOLOGY

MASSACHUSETTS INSTITUTE OF TECHNOLOGY

| Document Type | DCC Number T1100171-v1 | March 28, 2011 |
| --- | --- | --- |
| aLIGO BSC ISI Interface Chassis Test Plan |
| B. Abbott, S. Abbott |

Distribution of this draft:

This is an internal working note of the LIGO Laboratory

California Institute of Technology	Massachusetts Institute of Technology

LIGO Project – MS 18-33	LIGO Project – MS 20B-145

Pasadena, CA 91125	Cambridge, MA 01239

Phone (626) 395-2129	Phone (617) 253-4824

Fax (626) 304-9834	Fax (617) 253-7014

E-mail: info@ligo.caltech.edu	E-mail: info@ligo.mit.edu

http://www.ligo.caltech.edu/

# Introduction

The tests described below are required to verify the correct operation of the aLIGO BSC ISI Interface Chassis electronics.  These tests should verify the correct operation of all four boards internal to this chassis.  There is an L4C interface board, a GS-13 Interface board, a Capacitive Position Sensor Interface board, and a Back Board in each chassis.

# Test Equipment

Network Analyzer

Oscilloscope

Digital Multimeter (DMM)

Triple Power Supply (+/- 18V, 5V Fixed)

Two 25-pin Dsub Breakouts

# Serial Numbers

Input the internal board serial numbers below:

| L4C Interface Board | GS-13 Interface Board | Cap. Pos. Sens. Int. | Back Board |
| --- | --- | --- | --- |

# Tests

## Input Power

Hook up power to the back panel of the chassis with +18V going to pin 1, -18V to pin 3, and GND going to pin 2.  Turn on all of the power switches, and record the current being drawn.  The nominal current is ~320mA +/- 10mA.

+18V_______mA   -18V_______mA

At this point, all of the front and back panel power LEDs should be lit.

LEDs Lit?____________

## Voltage Output

The following tests make sure that the chassis is supplying the correct voltages to the geophones and Position sensor.  On the appropriate connector, verify that the correct voltages are present in the table below.  Use the appropriate switch on the back panel to turn on and off the power to the specific device to verify switch operation.  Watch the front panel LEDs to verify the correct switch operation.

| Connector | Pin | Voltage expected | Present and switchable? |
| --- | --- | --- | --- |
| Geophones 1&2 | Pin 3(+) to Pin 16(GND) | +15V |  |
| Geophones 1&2 | Pin 4(-) to Pin 17(GND) | -15V |  |
| Geophones 1&2 | Pin11(+) to Pin 24(GND) | +15V |  |
| Geophones 1&2 | Pin 12(-) to Pin 25(GND) | -15V |  |
| Geophones 3&4 | Pin 3(+) to Pin 16(GND) | +15V |  |
| Geophones 3&4 | Pin 4(-) to Pin 17(GND) | -15V |  |
| Geophones 3&4 | Pin11(+) to Pin 24(GND) | +15V |  |
| Geophones 3&4 | Pin 12(-) to Pin 25(GND) | -15V |  |
| Capacitive Position Sensor Power | Pin 1(+) to Pin 2(GND) | +18V |  |
| Capacitive Position Sensor Power | Pin3 (-) to Pin 2 (GND) | -18V |  |

## Channel Tests

### Geophone Pressure Read backs

Using a network analyzer, float channel 1 and 2 and input a 50mVp-p, signal into the pins below, and verify that the transfer function in the table is correct.

| Input Connector | Signal Out | Expected transfer function | Function correct? |
| --- | --- | --- | --- |
| Geophones 1&2 pins 2(+) and 15(-) | To AA Chassis pins 4(+) and 17(-) | DC Gain of 4 with an 8Hz pole |  |
| Geophones 1&2 pins 2(+) and 10(-) | To AA Chassis pins 3(+) and 16(-) | DC Gain of 20 with an 8Hz pole |  |
| Geophones 3&4 pins 2(+) and 15(-) | To AA Chassis pins 8(+) and 21(-) | DC Gain of 4 with an 8Hz pole |  |
| Geophones 3&4 pins 2(+) and 10(-) | To AA Chassis pins 7(+) and 10(-) | DC Gain of 20 with an 8Hz pole |  |

### Geophone Channel Tests

To test the whitening and gain of the geophone channels, input the same signal as above, from 1Hz to 10KHz in the pins specified below, and read the appropriate outputs.

4.3.2.1 No Binary In pins Grounded

| Input Connector | Signal Out | Expected transfer function | Function correct? |
| --- | --- | --- | --- |
| Geophones 1&2 pins 1(+) and 14(-) | To AA Chassis pins 1(+) and 14(-) | DC Gain of 14. |  |
| Geophones 1&2 pins 9(+) and 22(-) | To AA Chassis pins 2(+) and 15(-) | DC Gain of 14. |  |
| Geophones 3&4 pins 1(+) and 14(-) | To AA Chassis pins 5(+) and 18(-) | DC Gain of 20 |  |
| Geophones 3&4 pins 9(+) and 22(-) | To AA Chassis pins 6(+) and 19(-) | DC Gain of 20 |  |

4.3.2.2  Ground “From Binary Out” connector pins 1, 2, 3, and 4 to pin 5.

| Input Connector | Signal Out | Expected transfer function | Function correct? |
| --- | --- | --- | --- |
| Geophones 1&2 pins 1(+) and 14(-) | To AA Chassis pins 1(+) and 14(-) | DC Gain of 2 |  |
| Geophones 1&2 pins 9(+) and 22(-) | To AA Chassis pins 2(+) and 15(-) | DC Gain of 2 |  |
| Geophones 3&4 pins 1(+) and 14(-) | To AA Chassis pins 5(+) and 18(-) | DC Gain of 2 |  |
| Geophones 3&4 pins 9(+) and 22(-) | To AA Chassis pins 6(+) and 19(-) | DC Gain of 2 |  |

Unground  “From Binary Out” connector pins 1, 2, 3, and 4 to pin 5.

4.3.2.3  Ground “From Binary Out” connector pins 8 and 9 to pin 5.

| Input Connector | Signal Out | Expected transfer function | Function correct? |
| --- | --- | --- | --- |
| Geophones 1&2 pins 1(+) and 14(-) | To AA Chassis pins 1(+) and 14(-) | DC Gain of 14 |  |
| Geophones 1&2 pins 9(+) and 22(-) | To AA Chassis pins 2(+) and 15(-) | DC Gain of 14 |  |
| Geophones 3&4 pins 1(+) and 14(-) | To AA Chassis pins 5(+) and 18(-) | DC Gain of 20, a 10Hz zero, and poles at 50Hz, and 5KHz |  |
| Geophones 3&4 pins 9(+) and 22(-) | To AA Chassis pins 6(+) and 19(-) | DC Gain of 20, a 10Hz zero, and poles at 50Hz, and 5KHz |  |

Unground  “From Binary Out” connector pins 8 and 9 from pin 5.

### Geophone Channel Noise Tests

To test the noise of the geophone channels, ground the appropriate pins on the input connector, and do a noise measurement from 0.1 Hz to 10KHz on the output connector pins specified below. Read the noise at 1Hz and 10Hz.  Ground the appropriate pins on the From Binary Out connector.  The observed numbers should be blow the specification to pass this test.

| Geophone | Ground the following | Expected Noise @1Hz | Actual Noise @1Hz | Expected Noise @10Hz | Actual Noise @10Hz |
| --- | --- | --- | --- | --- | --- |
| Geophone 1 GND pins 1 and 14 Read Out Pins 1&14 | Nothing | 270nV |  | 157nV |  |
|  | Pin 1 (Gain) | 190nV |  | 88nV |  |
| Geophone 2 GND pins 9 and 22 Read Out pins 2 and 15 | Nothing | 270nV |  | 157nV |  |
|  | Pin 2 (Gain) | 190nV |  | 88nV |  |
| Geophone 3 GND pins 1 and 14 Read Out pins 5 and 18 | Nothing | 326nV |  | 197nV |  |
|  | Pin 3 (Gain) | 190nV |  | 88nV |  |
|  | Pin 8 (White) | 190nV |  | 107nV |  |
|  | Pin 3 & Pin 8 | 328nV |  | 269nV |  |
| Geophone 4 GND pins 9 and 22 GND pins 6and 19 | Nothing | 326nV |  | 197nV |  |
|  | Pin 4 (Gain) | 190nV |  | 88nV |  |
|  | Pin 9 (White) | 190nV |  | 107nV |  |
|  | Pin 4 & Pin 9 | 328nV |  | 269nV |  |

### Capacitive Position Sensor Channel Tests

First, using a voltage calibrator, put 1V DC (single-ended to Ground) into each channel in the table below and verify that each output has a gain of 1V/V.  Then, using an SR785 put a 50mV signal from 700mHz to 102KHz into each channel, and read the output at the frequencies in the table.

| Input Pin (Cap Pos Interface) | Output Pin (To AA Chassis) | DC Output (Nom. 1V +/- 0.05V) | Output @ 700mHZ (Nom. 13dB +/- 0.5dB) | Output @ 1.4HZ (Nom. 17dB +/- 0.5dB) | Output @ 1KHZ (Nom. 20dB +/- 0.5dB) | Output @ 23.4KHZ (Nom. 17dB +/- 0.5dB) |
| --- | --- | --- | --- | --- | --- | --- |
| Pin1 | Pin 9 |  |  |  |  |  |
| Pin14 | Pin 22 |  |  |  |  |  |
| Pin2 | Pin 10 |  |  |  |  |  |
| Pin15 | Pin 23 |  |  |  |  |  |
| Pin3 | Pin 11 |  |  |  |  |  |
| Pin16 | Pin 24 |  |  |  |  |  |
| Pin4 | Pin 12 |  |  |  |  |  |
| Pin17 | Pin 25 |  |  |  |  |  |

### Capacitive Position Sensor Noise Tests

Set up a SR785 to measure noise between 100mHz and 12.4Hz.  Ground the appropriate input pin in the table below, and note the noise seen between the output pin and Ground on an SR785 at the frequencies requested.

| Ground Pin (Cap Pos Interface) | Output Pin (To AA Chassis) | Noise @ 312mHZ Nominal is under 3.2uV/√Hz | Noise @ 1HZ Nominal is under 1.5uV/√Hz | Noise @ 5HZ Nominal is under 525nV/√Hz | Noise @ 10HZ Nominal is under 375nV/√Hz |
| --- | --- | --- | --- | --- | --- |
| Pin1 | Pin 9 |  |  |  |  |
| Pin14 | Pin 22 |  |  |  |  |
| Pin2 | Pin 10 |  |  |  |  |
| Pin15 | Pin 23 |  |  |  |  |
| Pin3 | Pin 11 |  |  |  |  |
| Pin16 | Pin 24 |  |  |  |  |
| Pin4 | Pin 12 |  |  |  |  |
| Pin17 | Pin 25 |  |  |  |  |