## Staging building
BSC-ISI - Unit 4
Electronics Inventory
Hardware |  | Ligo reference | S/N
Interface Chassis - Corner 1 |  | D1002432 | S110223
Interface Chassis - Corner 2 |  |  | S1102224
Interface Chassis - Corner 3 |  |  | S1102218
Anti-Alliasing Chassis - Corner 1 |  | D1002693 | S1102693
Anti-Alliasing Chassis - Corner 2 |  |  | S1102694
Anti-Alliasing Chassis - Corner 3 |  |  | S1102679
Anti-image Chassis |  | D070081 | S1000250
Binary Input Chassis |  | D1001726 | S1101309
Binary Input Chassis |  |  | S11031308
Binary Output Chassis |  | D1001728 | S1101347
T240 Interface - Corner 1 |  | D1002694 | S1101040
T240 Interface - Corner 2 |  |  | S1101838
T240 Interface - Corner 3 |  |  | S1101839
I/O Chassis |  | n/a
Coil driver Pod 1 |  | D0902744 | S1000266
Coil driver Pod 2 |  |  | S1000269
Coil driver Pod 3 |  |  | S110692
[7 empty] | Actuator | Coil driver name | Resistance (Ω)
Stage 1 |  | Stage 2 |  |  |  |  | ST1 H1 | Coil1 Coarse 1 | 6.4
Actuator | Actuator S/N | Actuator | Actuator S/N |  |  |  | ST2 H1 | Coil 1 Fine 1 | 10.1
ST1 - H1 | 04 | ST2 - H1 | 42 |  |  |  | ST2 V1 | Coil 1 Fine 2 | 10.1
ST1 - H2 | 18 | ST2 - H2 | 41 |  |  |  | ST1 V1 | Coil 1 Coarse 2 | 6.3
ST1 - H3 | 186 | ST2 - H3 | 44 |  |  |  | ST1 H2 | Coil 2 Coarse 2 | 6.4
ST1 - V1 | 27 | ST2 - V1 | 46 |  |  |  | ST2 H2 | Coil 2 Fine 1 | 10.1
ST1 - V2 | 177 | ST2 - V2 | 48 |  |  |  | ST2 V2 | Coil 2 Fine 2 | 10.1
ST1 - V3 | 31 | ST2 - V3 | 45 |  |  |  | ST1 V2 | Coil 2 Coarse 2 | 6.4
[7 empty] | ST1 H3 | Coil 3 Coarse 1 | 6.4
Step 2 - Sensors inventory | [6 empty] | ST2 H3 | Coil 3 Fine 1 | 10.1
[7 empty] | ST2 V3 | Coil 3 Fine 2 | 10
CPS Stage 1 | CPS S/N | ADE board serial # |  |  |  |  | ST1 V3 | Coil 3 Coarse 2 | 6.3
H1 | 12951 | 12339
H2 | 13235 | 12410
H3 | 12959 | 12384 |  |  |  |  | Actuator | Resistance (Ω)
V1 | 12958 | 12341 |  |  |  |  | ST1 H1
V2 | 13444 | 12424 |  |  |  |  | ST1 H2
V3 | 12895 | 12826 |  |  |  |  | ST1 H3
ST1 V1
CPS Stage 2 | CPS S/N | ADE board serial # |  |  |  |  | ST1 V2
H1 | 12948 | 13076 |  |  |  |  | ST1 V3
H2 | 12961 | 12430 |  |  |  |  | ST2 H1
H3 | 13430 | 12576 |  |  |  |  | ST2 H2
V1 | 12904 | 12335 |  |  |  |  | ST2 H3
V2 | 12892 | 13064 |  |  |  |  | ST2 V1
V3 | 12945 | 12451 |  |  |  |  | ST2 V2
ST2 V3
Geophones GS13 | Serial Number | POD |  | E1000058
H1 | 865 | 67
H2 | 816 | 92
H3 | 862 | 43
V1 | 668 | 51 | HAM
V2 | 768 | 58
V3 | 750 | 72
Geophones L4C | Serial Number | POD |  | E1000136
H1 | 1104 | 96
H2 | 1120 | 94
H3 | 961 | 46
V1 | 925 | 35
V2 | 917 | 125
V3 | 801 | 42
Geophones T240 | Serial Number | POD
1 | 107 | 37
2 | 143 | 34
3 | 103 | 4
Stage 0-1 |  | Stage 1-2
Lockers | Shim thickness (mil) | Lockers | Shim thickness (mil)
Corner 1 | 122 | Corner 1 | 122
Corner 2 | 122 | Corner 2 | 120
Corner 3 | 127 | Corner 3 | 123
[8 empty] | table locked | 1041792635 | X1_ISI_ITMX_CPS_Read_Back_ISI_Locked_2013_01_09_105019
| Table locked |  | Table unlocked |  | Difference locked - unlocked |  |  | table unlocked | 1041793513 | X1_ISI_ITMX_CPS_Read_Back_ISI_Unlocked_2013_01_09_110457
Sensors | Offset (Mean) | Std deviation | Offset (Mean) | Std deviation | Offset (Mean) | mil
ST1 - H1 | 178.3797 | 7.6995 | 341.4621 | 39.952 | -163.0825 | -0.19415
ST1 - H2 | 222.815 | 6.7987 | 487.8664 | 28.0854 | -265.0515 | -0.31554
ST1 - H3 | -337.0075 | 7.4263 | 11.5771 | 26.3107 | -348.5846 | -0.41498
ST1 - V1 | 30.8487 | 5.4917 | -439.137 | 30.0965 | 469.9857 | 0.55951
ST1 - V2 | -174.3022 | 5.5901 | -274.9284 | 32.7771 | 100.6263 | 0.11979
ST1 - V3 | 408.3173 | 8.2157 | 249.5116 | 24.8957 | 158.8057 | 0.18905
ST2 - H1 | -386.7871 | 23.2287 | -319.3198 | 35.4632 | -67.4674 | -0.02008
ST2 - H2 | -182.2927 | 42.4882 | -257.0659 | 37.6491 | 74.7732 | 0.022254
ST2 - H3 | -592.0015 | 43.0472 | -265.6748 | 55.8587 | -326.3267 | -0.097121
ST2 - V1 | 16.7041 | 31.1201 | -576.2715 | 72.9337 | 592.9756 | 0.17648
ST2 - V2 | -125.558 | 40.3902 | 96.8629 | 83.2432 | -222.4209 | -0.066197
ST2 - V3 | -227.4643 | 30.8758 | -1139.7025 | 61.2256 | 912.2382 | 0.2715
Sensors | Push in positive direction | Push in negative direction | Mil | Mil | Railing | Actuator Gap Check
ST1 - H1 | 20000 | -17000 | 23.8095 | -20.2381 |  | X
ST1 - H2 | 18000 | -16000 | 21.4286 | -19.0476 |  | X
ST1 - H3 | 18500 | -18000 | 22.0238 | -21.4286 |  | X
ST1 - V1 | 25000 | -21000 | 29.7619 | -25 |  | X
ST1 - V2 | 21000 | -21000 | 25 | -25 |  | X
ST1 - V3 | 20000 | -19000 | 23.8095 | -22.619 |  | X
ST2 - H1 |  |  |  |  | X | X
ST2 - H2 |  |  |  |  | X | X
ST2 - H3 |  |  |  |  | X | X
ST2 - V1 |  |  |  |  | X | X
ST2 - V2 |  |  |  |  | X | X
ST2 - V3 |  |  |  |  | X | X
Sensor readout (counts) | Negative drive | no drive | Positive drive | Amplitude count | mil
ST1 - H1 | -16071.106 | 423 | 16157.988 | 32229.094 | 38.368
ST1 - H2 | -16485.42 | 248 | 15918.246 | 32403.666 | 38.5758
ST1 - H3 | -16772.354 | -136 | 15417.602 | 32189.956 | 38.3214
ST1 - V1 | -15171.804 | -142 | 14975.982 | 30147.786 | 35.8902
ST1 - V2 | -14630.65 | -79 | 14610.022 | 29240.672 | 34.8103
ST1 - V3 | -14048.082 | 746 | 15547.776 | 29595.858 | 35.2332
ST2 - H1 | -10631.464 | -436 | 9755.3936 | 20386.8576 | 6.0675
ST2 - H2 | -10014.3936 | 71 | 10172.094 | 20186.4876 | 6.0079
ST2 - H3 | -9734.4358 | 207 | 10156.554 | 19890.9898 | 5.9199
ST2 - V1 | -12559.684 | -187 | 12190.006 | 24749.69 | 7.366
ST2 - V2 | -11799.792 | 491 | 12782.014 | 24581.806 | 7.316
ST2 - V3 | -12891.918 | -844 | 11208.248 | 24100.166 | 7.1727
Calibration check on stage 2 + first blade stiffnes measurements
ST1 locked - ST2 unlocked |  | No load | No load 2 | Mean No load | Load 1 | Load 2 | Load 3 | Load 4 | Mean Load | Diff | [7 empty] | ST1 locked - ST2 unlocked |  | No load | No load 2 | Mean No load | Load 1 | Load 2 | Load 3 | Load 4 | Mean Load | Diff
3x5Kg | V1 | -547 | -553 | -550 | -25893 | -26059 | -26051 |  | -26001 | -25451 | [5 empty] | -1428.0497 | 0.025222 | 3x5Kg | V1 | 6418 |  | 6418 | -20163 |  |  |  | -20163 | -26581
| V2 | 11 | -20 | -4.5 | -26496 | -26364 | -26556 |  | -26472 | -26467.5 | [5 empty] | -1373.2046 | 0.062659 | 2012-10-12 | V2 | 7746 |  | 7746 | -18467 |  |  |  | -18467 | -26213
| V3 | -1162 | -1158 | -1160 | -26747 | -26691 | -26581 |  | -26673 | -25513 | [5 empty] | -1424.5793 | 0.027591 |  | V3 | -1624 |  | -1624 | -28867 |  |  |  | -28867 | -27243
[10 empty] | -25810.5 | count | 3329.2288 | Cout/mil | [14 empty] | -26679 | count | 0 | Cout/mil
[10 empty] | -7.7527 | mil | [16 empty] | #DIV/0! | mil
[10 empty] | -1408.1592 | lb/in | 131072 | Ct/m | [14 empty] | #DIV/0! | lb/in | 131072 | Ct/m
[10 empty] | -3.8799 | % | 1465 | lb/in | Reference value stage 2 | [13 empty] | #DIV/0! | % | 1465 | lb/in | Reference value stage 2
[12 empty] | 3329.2288 | Ct/mil
ST1 unlocked - ST2 locked |  | No load | No load 2 | Mean no load | Load 15Kg 1 | Load 15Kg 2 | Load 15Kg 3 | Mean load | Load 30Kg | Diff 1 | Diff 2
| V1 | -445 | -448 | -446.5 | -8033 | -8034 | -8099 | -8055.3333 | -15592 | -7610.3333 | -15144
| V2 | -264 | -258 | -261 | -7968 | -7888 | -7804 | -7886.6667 | -15707 | -7622.6667 | -15449
| V3 | 237 | 226 | 231.5 | -7414 | -7592 | -7424 | -7476.6667 | -15082 | -7713.6667 | -15308
[10 empty] | -7648.8889 | -15300.3333 | count | 832.3072 | Ct/mil
[10 empty] | -9.19 | -18.383 | mil | 1241 | lb/in | Reference value stage 1
[10 empty] | -1198.3935 | -1198.1934 | lb/in | 107617.3298 | KN/m
[10 empty] | -3.4332 | -3.4494 | %
| Sensors | Mean diff (counts) | Mean diff m | K (N/m)
Stage 1 | V1 | #REF! | #REF! | #REF!
| V2 | #REF! | #REF! | #REF!
| V3 | #REF! | #REF! | #REF!
|  |  | Average (N/m) | #REF!
| Sensors | Mean diff (counts) | Mean diff m | K (N/m)
Stage 2 | V1 | 0 | 0 | #REF!
| V2 | 3.4332 | 1.0368e-07 | #REF!
| V3 | 0 | 0 | #REF!
|  |  | Average (N/m) | #REF!
Second Measurements
|  | Unloaded |  |  |  | Loaded
| Stage 2 | Meas 1 | Meas 2 | Meas 3 | Average | Meas 1 | Meas 2 | Meas 3 | Average | Diff
| V1 |  |  |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0!
| V2 |  |  |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0!
| V3 |  |  |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0!
[10 empty] | #DIV/0! | count
[10 empty] | #DIV/0! | mil
Mass | 15006 | g
| 33.0826 | lb
K | #DIV/0! | lb/in
Error with nominal | #DIV/0! | %
|  | Unloaded |  |  |  | Loaded
| Stage1 | Meas 1 | Meas 2 | Meas 3 | Average | Meas 1 | Meas 2 | Meas 3 | Average | Diff
| V1 |  |  |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0!
| V2 |  |  |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0!
| V3 |  |  |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0!
[10 empty] | #DIV/0! | count
[10 empty] | #DIV/0! | mil
Mass | 21133 | g
| 46.5903 | lb
K | #DIV/0! | lb/in
Error with nominal | #DIV/0! | %
Stage 1 - Platform move for 32K counts | 12.6263 | mil | coef*U/(K*R)
Stage 2 - Platform move for 32K counts | 3.5938 | mil | [7 empty] | 2012-10-18
|  | Slope | Offset | Average slope | Variation from average(%)
Stage 1 | ST1 - H1 | 0.61347 | 107.8038 | 0.60338 | 1.6721
| ST1 - H2 | 0.6038 | 203.5664 |  | 0.070008
| ST1 - H3 | 0.59287 | -107.4789 |  | -1.7421
| ST1 - V1 | 0.50367 | -487.2087 | 0.49536 | 1.6778
| ST1 - V2 | 0.48779 | -253.32 |  | -1.5268
| ST1 - V3 | 0.49461 | 535.5041 |  | -0.15105
Stage 2 | ST2 - H1 | 0.34094 | -401.3207 | 0.33798 | 0.8747
| ST2 - H2 | 0.33908 | -280.8061 |  | 0.32531
| ST2 - H3 | 0.33393 | 411.3281 |  | -1.2
| ST2 - V1 | 0.41409 | -32.2154 | 0.41002 | 0.99375
| ST2 - V2 | 0.41171 | 412.7267 |  | 0.41378
| ST2 - V3 | 0.40425 | -409.9497 |  | -1.4075
Pressure (count)
Sensors | Day 1 | Day 2 | Day 3 | Day 4
L4C - H1
L4C - H2
L4C - H3
L4C - V1
L4C - V2
L4C - V3
GS13 - H1
GS13 - H2
GS13 - H3
GS13 - V1
GS13 - V2
GS13 - V3
T240 - 1
T240 - 2
T240 - 3
Test Local drive | [9 empty] | Refrence table
|  |  | Sensors | [8 empty] | Sensors
|  |  | ST1 - H1 | ST1 - H2 | ST1 - H3 | ST1 - V1 | ST1 - V2 | ST1 - V3 |  |  |  | ST1 - H1 | ST1 - H2 | ST1 - H3 | ST1 - V1 | ST1 - V2 | ST1 - V3
| Actuators | ST1 - H1 | 4286.8815 | 1722.6672 | 1736.0657 | -6.0389 | 5.0252 | -0.96622 |  | Actuators | ST1 - H1
|  | ST1 - H2 | 1690.6259 | 4205.5838 | 1706.3581 | 4.8478 | 20.0828 | 3.2952 |  |  | ST1 - H2
|  | ST1 - H3 | 1673.4167 | 1664.1708 | 4161.3581 | -4.8835 | 6.0912 | -6.3961 |  |  | ST1 - H3
|  | ST1 - V1 | 42.2893 | -176.6083 | 90.0933 | 3526.8168 | -650.0924 | -623.4187 |  |  | ST1 - V1
|  | ST1 - V2 | 97.8011 | 36.2319 | -146.2882 | -605.7878 | 3424.6199 | -626.2662 |  |  | ST1 - V2
|  | ST1 - V3 | -156.5899 | 98.1034 | 40.292 | -644.5636 | -597.3492 | 3457.7736 |  |  | ST1 - V3
|  |  | Sensors | [8 empty] | Sensors
|  |  | ST2 - H1 | ST2 - H2 | ST2 - H3 | ST2 - V1 | ST2 - V2 | ST2 - V3 |  |  |  | ST2 - H1 | ST2 - H2 | ST2 - H3 | ST2 - V1 | ST2 - V2 | ST2 - V3
| Actuators | ST2 - H1 | 2390.7763 | 373.3628 | 356.246 | -2.4047 | -1.0828 | -5.9785 |  | Actuators | ST2 - H1
|  | ST2 - H2 | 364.0813 | 2386.5705 | 368.2 | 13.4564 | 20.4942 | -10.0139 |  |  | ST2 - H2
|  | ST2 - H3 | 359.2166 | 356.5214 | 2337.392 | -11.7581 | 21.9981 | 22.8252 |  |  | ST2 - H3
|  | ST2 - V1 | 79.9388 | 144.218 | -210.702 | 2895.3813 | 325.5312 | -51.4628 |  |  | ST2 - V1
|  | ST2 - V2 | -198.8554 | 86.8455 | 118.4 | -6.8844 | 2901.2387 | 326.6872 |  |  | ST2 - V2
|  | ST2 - V3 | 117.3749 | -188.722 | 89.17 | 274.6003 | 13.7352 | 2826.7333 |  |  | ST2 - V3
Test Cartesian drive |  |  | Sensors | [8 empty] | Sensors
|  |  | ST1 - X | ST1 - Y | ST1 - Z | ST1 - RY | ST1 - RY | ST1 - RZ |  |  |  | ST1 - X | ST1 - Y | ST1 - Z | ST1 - RY | ST1 - RY | ST1 - RZ
| Actuators | ST1 - X | 1714.6537 | -8.9596 | -4.8011 | -3.878 | 6.8667 | 52.8905 |  | Actuators | ST1 - X | 1750 | 0 | 0 | 0 | 0 | 0
|  | ST1 - Y | 3.2673 | 1673.6504 | -3.2469 | 4.4119 | -0.1104 | 31.5296 |  |  | ST1 - Y | 0 | 1750 | 0 | 0 | 0 | 0
|  | ST1 - Z | 7.6566 | -3.9443 | 730.8936 | -24.5206 | -1.1309 | -9.6162 |  |  | ST1 - Z | 0 | 0 | 750 | 0 | 0 | 0
|  | ST1 - RX | -2.0256 | 362.0571 | -16.8932 | 3005.5068 | 15.2494 | 9.2669 |  |  | ST1 - RX | 0 | 375 | 0 | 3000 | 0 | 0
|  | ST1 - RY | -360.0138 | 18.9069 | -6.6637 | 25.6158 | 2974.4027 | 6.2319 |  |  | ST1 - RY | -375 | 0 | 0 | 0 | 3000 | 0
|  | ST1 - RZ | 30.9041 | 7.4585 | -8.1644 | 8.2384 | -7.8133 | 3235.8087 |  |  | ST1 - RZ | 0 | 0 | 0 | 0 | 0 | 3300
|  |  | Sensors | [8 empty] | Sensors
|  |  | ST2 - X | ST2 - Y | ST2 - Z | ST2 - RY | ST2 - RY | ST2 - RZ |  |  |  | ST2 - X | ST2 - Y | ST2 - Z | ST2 - RY | ST2 - RY | ST2 - RZ
| Actuators | ST2 - X | 1349.1951 | 12.2718 | -22.0052 | -7.01 | 21.0525 | -2.3874 |  | Actuators | ST2 - X | 1350 | 10 | 30 | 0 | 25 | 20
|  | ST2 - Y | 5.4856 | 1336.967 | -42.5815 | -30.7098 | 5.5892 | 24.8887 |  |  | ST2 - Y | -9.697 | 1350 | 20.1471 | -25 | 0 | 20
|  | ST2 - Z | 6.9915 | -12.7759 | 1028.0339 | 9.8844 | -28.4252 | 7.6055 |  |  | ST2 - Z | -0.3159 | 1.7254 | 1100 | -10 | -30 | 20
|  | ST2 - RX | 1.4273 | -28.3186 | -24.9316 | 4184.8632 | 67.3584 | 2.1299 |  |  | ST2 - RX | 10 | -15 | 20 | 4300 | 30 | 20
|  | ST2 - RY | 19.3486 | -9.7891 | -76.6886 | 61.804 | 4244.3268 | 14.4499 |  |  | ST2 - RY | 30 | 0 | 30.4223 | 40 | 4300 | 20
|  | ST2 - RZ | -6.6497 | 26.631 | -34.2202 | -15.5123 | 3.2673 | 2563.5044 |  |  | ST2 - RZ | 0 | 10 | 30 | -25 | -15 | 2600
|  |  | Sensors | [8 empty] | Sensors
|  |  | ST1 - H1 | ST1 - H2 | ST1 - H3 | ST1 - V1 | ST1 - V2 | ST1 - V3 |  |  |  | ST1 - H1 | ST1 - H2 | ST1 - H3 | ST1 - V1 | ST1 - V2 | ST1 - V3
| Actuators | ST1 - X | 1740.072 | -802.0519 | -797.2482 | -0.14616 | 0.35642 | 6.8585 |  | Actuators | ST1 - X | 1800 | -820 | -820 | 0 | 0 | 0
|  | ST1 - Y | 9.7789 | 1490.6957 | -1418.5628 | -6.7751 | -1.2511 | 2.7532 |  |  | ST1 - Y | 0 | 1500 | -1500 | 0 | 0 | 0
|  | ST1 - Z | -14.4202 | -0.007596 | -9.7907 | 764.2466 | 719.4722 | 729.3223 |  |  | ST1 - Z | 0 | 0 | 0 | 771.8322 | 750 | 700
|  | ST1 - RX | 16.3013 | 168.1051 | -147.9794 | -2914.8076 | 2418.3091 | 443.7886 |  |  | ST1 - RX | 0 | 160 | -160 | -2950 | 2450 | 450
|  | ST1 - RY | -178.78 | 114.6528 | 74.2867 | -1170.887 | -1894.6791 | 3049.3565 |  |  | ST1 - RY | -200 | 110 | 70 | -1150 | -2000 | 3050
|  | ST1 - RZ | 3112.6098 | 3111.5769 | 3087.7762 | -7.4836 | 7.1527 | -9.5989 |  |  | ST1 - RZ | 3200 | 3200 | 3200 | 0 | 0 | 0
|  |  | Sensors | [8 empty] | Sensors
|  |  | ST2 - H1 | ST2 - H2 | ST2 - H3 | ST2 - V1 | ST2 - V2 | ST2 - V3 |  |  |  | ST2 - H1 | ST2 - H2 | ST2 - H3 | ST2 - V1 | ST2 - V2 | ST2 - V3
| Actuators | ST2 - X | 667.897 | -1325.0048 | 644.5548 | -57.5616 | -27.2386 | 7.041 |  | Actuators | ST2 - X | 700 | -1350 | 650 | 0 | 0 | 0
|  | ST2 - Y | 1178.9839 | 35.0599 | -1150.2199 | -19.5845 | -41.7454 | -31.3595 |  |  | ST2 - Y | 1200 | 0 | -1150 | 0 | 0 | 0
|  | ST2 - Z | -6.0521 | 28.3702 | -1.9493 | 1011.352 | 1055.0368 | 1010.6261 |  |  | ST2 - Z | 0 | 0 | 0 | 1100 | 1100 | 1100
|  | ST2 - RX | -284.4308 | 11.7247 | 267.9133 | -2478.6058 | 2451.9098 | 0.984 |  |  | ST2 - RX | -300 | 0 | 300 | -2500 | 2500 | -50
|  | ST2 - RY | 149.3741 | -300.5887 | 181.4931 | -1618.7234 | -1417.5907 | 2812.7278 |  |  | ST2 - RY | 200 | -300 | 200 | -1500 | -1400 | 3000
|  | ST2 - RZ | 1776.1608 | 1796.992 | 1744.9444 | -45.0564 | -54.3321 | -31.6668 |  |  | ST2 - RZ | 1800 | 1800 | 1800 | 40 | 40 | 40
Stage 2 - - 2612lb - 5% = 2481lb
Mass Budget - Mass type | Quantity | Weight | Unit | Weight (lb) |  |  | Corner 1 | 2 type 6 + 1 type 5 + 2 type 4 + 2 type 1 + 1 type 0
| 3 | 610 | lb | 1830 |  |  | Corner 2 | 2 type 6 + 3 type 4 + 1 type 3 + 1 type 1 + 1 type 0
| 2 | 233 | lb | 466 |  |  | Corner 3 | 4 type 6 + 2 type 3 + 6 type 1 + 1 type 0
type 0 | 2 | 0.5 | lb | 1
tupe 1 | 9 | 1.1 | lb | 9.9 |  |  | top | 3 x 610lb + 2 x 233lb
tupe 2 | 0 | 2.2 | lb | 0
tupe 3 | 3 | 4.5 | lb | 13.5
tupe 4 | 5 | 7.9 | lb | 39.5
tupe 5 | 1 | 15.6 | lb | 15.6
tupe 6 | 8 | 27.2 | lb | 217.6
|  |  |  | 2593.1 |  |  | 0.76179
Stage 2
| Quantity | Weight (Kg) | Total Weight(lb) | Total Weight(Kg)
|  | 276.6913 | 0 | 0
|  | 105.687 | 0 | 0
|  | 1.1 | 0 | 0
|  | 7.8 | 0 | 0
|  | 4.5 | 0 | 0
|  | 2.1 | 0 | 0
| Total : |  | 0 | 0
Stage 1
Mass Budget Nominal - 240lb - 5% = 228lb |  |  |  |  | 64.2222
D0902612 | 12 | lb
D0902613 | 15 | lb
D0902614 | 10 | lb
D0902615 | 10 | lb
[7 empty] | Corner 1 | 2 vibration absorbers (15lb) + D0902613 + 1 type4
[7 empty] | Corner 2 | 2 vibration absorbers (15lb) + D0902613 + 1 type 4 + 1 type 1 + 1 type 3 + 2 typ5 + 1 type2
[7 empty] | Corner 3 | 2 vibration absorbers + 1 type 3 + 2 type 1
Stage 1
Location | Weight (lb)
Corner 1 | 50
Corner 2 | 55
Corner 3 | 34
Total | 139
Density (ASI 304) | 0.289 | lb/in3
Surface
Thickness | Mass(lb)
0.806 | 5.5972
0.76 | 5.2778
0.232 | 1.6111 |  |  |  | Weight (lb)
|  | Weight (lb) |  | D071200-0 | 0.6
D0902616-1 | 0.5 | 3 |  | D071200-1 | 1.1
D0902616-2 | 0.14 | 1 |  | D071200-2 | 2.2
D0902616-3 | 0.28 | 2 |  | D071200-3 | 4.5
D0902616-4 | 0.72 | 5 |  | D071200-4 | 7.9
D0902616-5 | 0.07 | 0.5 |  | D071200-5 | 15.6
D0902616-6 |  |  |  | D071200-6 | 27.2
6.9444
DCC Number | Description | Length (in) | Serial Number | Location
D1100154 | 25-pin M-to-two 9-pin F straight | 48 | S1104259 | L4C - C2
D1100154 | 25-pin M-to-two 9-pin F straight | 48 | S1104261 | L4C - C1
D1100154 | 25-pin M-to-two 9-pin F straight | 48 | S1104264 | L4C - ext - C3
D1100155 | 25-pin M-to-two 9-pin F straight | 120 | S1104123 | GS-13 - C3
D1100155 | 25-pin M-to-two 9-pin F straight | 120 | S1104247 | GS-13 - C1
D1100155 | 25-pin M-to-two 9-pin F straight | 120 | S1104248 | GS-13 - C2
D1100148 | 2-wire, 14awg 3-pin M to 3-pin F | 80 | S1106904 | ST1 - H3 - ext - C3
D1100148 | 2-wire, 14awg 3-pin M to 3-pin F | 80 | S1106914 | ST2 - V1 - ext - C1
D1100148 | 2-wire, 14awg 3-pin M to 3-pin F | 80 | S1106916 | ST2 - H3 - ext - C3
D1100148 | 2-wire, 14awg 3-pin M to 3-pin F | 80 | S1106918 | ST1 - H2 - ext
D1100148 | 2-wire, 14awg 3-pin M to 3-pin F | 80 | S1106919 | ST2 - V2 - ext
D1100148 | 2-wire, 14awg 3-pin M to 3-pin F | 80 | S1106920 | ST1 - V3 - ext
D1100148 | 2-wire, 14awg 3-pin M to 3-pin F | 80 | S1106922 | ST2 - H2 - ext
D1100148 | 2-wire, 14awg 3-pin M to 3-pin F | 80 | S1106926 | ST1 - V1 - ext
D1100148 | 2-wire, 14awg 3-pin M to 3-pin F | 80 | S1106928 | ST2 - V3 - ext
D1100148 | 2-wire, 14awg 3-pin M to 3-pin F | 80 | S1106938 | ST1 - H1 - ext
D1100148 | 2-wire, 14awg 3-pin M to 3-pin F | 80 | S1106942 | ST1 - V2 - ext
D1100150 | 2-wire, 14awg 2 pins to 3-pin F | 40 | S1107082 | ST1 - H1
D1100150 | 2-wire, 14awg 2 pins to 3-pin F | 40 | S1107089 | ST1 - V2
D1100150 | 2-wire, 14awg 2 pins to 3-pin F | 40 | S1107093 | ST1 - H3
D1100150 | 2-wire, 14awg 2 pins to 3-pin F | 40 | S1107099 | ST1 - H2 - ext
D1100150 | 2-wire, 14awg 2 pins to 3-pin F | 40 | S1107101 | ST1 - V3
D1100150 | 2-wire, 14awg 2 pins to 3-pin F | 40 | S1107102 | ST1 - V1
D1100151 | 2-wire, 14awg 2 pins to 3-pin F | 60 | S1107162 | ST2 - V2
D1100151 | 2-wire, 14awg 2 pins to 3-pin F | 60 | S1107171 | ST2 - H2
D1100151 | 2-wire, 14awg 2 pins to 3-pin F | 60 | S1107182 | ST2 - V3
D1100151 | 2-wire, 14awg 2 pins to 3-pin F | 60 | S1107194 | ST2 - V1
D1100151 | 2-wire, 14awg 2 pins to 3-pin F | 60 | S1107198 | ST2 - H1
D1100152 | 25-pin F-to-25-pin F | 110 | S1107230 | Trill - C3
D1100152 | 25-pin F-to-25-pin F | 110 | S1107236 | Trill - C2
D1100152 | 25-pin F-to-25-pin F | 110 | S1107238 | Trill - C1
D1100153 | 25-pin F-to-25-pin F | 80 | S1107309 | GS-13 - ext - C2
D1100153 | 25-pin F-to-25-pin F | 80 | S1107320 | L4C - ext - C1
D1100153 | 25-pin F-to-25-pin F | 80 | S1107321 | L4C - ext - C3
D1100153 | 25-pin F-to-25-pin F | 80 | S1107322 | L4C - ext - C2
D1100153 | 25-pin F-to-25-pin F | 80 | S1107323 | GS-13 - ext - C3
D1100153 | 25-pin F-to-25-pin F | 80 | S1107325 | GS-13 - ext - C1
D1100148 | 2-wire, 14awg 3-pin M to 3-pin F | 80 | S1106999 | ST2 - H1 - ext
D1100151 | 2-wire, 14awg 2 pins to 3-pin F | 60 | S1107208 | ST2 - H3
Pressure (counts)
Sensors | Corner 1 | Corner 2 | Corner 3
ST1-L4C-P | -279 | -254 | -256
ST1-L4C-D | 24769 | 24652 | 24652
ST1-GS13-P | 1029 | 974 | 977
ST1-GS13-D | 24540 | 24729 | 24729
ST1-T240-P | 13542 | 13140 | 13142
Table Tilted
Location | Time GPS | Mass Location
A | 1041796840 | Stage 1 Actuators
B | 1041797132 | Stage 2 Blabde
C | 1041797380 | Stage 2 Actuators
D | 1041797590 | Stage 3 Blade
E | 1041797910 | Stage 3 Actuators
F | 1041798117 | Stage 1 Blade

## LVEA
BSC-ISI | LVEA
First test without ITMY and FMY
Electronics Inventory
Hardware |  | Rack # | Ligo reference | S/N
Interface Chassis Pod 1 |  |  | D1002432 | S1102228
Interface Chassis Pod 2 |  |  | D1002433 | S1102230
Interface Chassis Pod 3 |  |  | D1002434 | S1102229
Anti-alliasing Chassis |  |  | D1002693 | S1103404
Anti-alliasing Chassis |  |  | D1002693 | S1103405
Anti-image Chassis |  |  | D1002693 | S1103402
Binary Input Chassis |  |  | D1001726 | S1101284
|  |  | D1001726 | S1101280
Binary Output Chassis |  |  | D1001728 | S1101319
T240 Interface Pod 1 |  |  | D1002694 | S1103177
T240 Interface Pod 2 |  |  | D1002694 | S1103181
T240 Interface Pod 3 |  |  | D1002694 | S1103180
Anti-image Chassis |  |  | D1000305 | S1103502
Coil driver Pod 1 |  |  | D0902744 | S1103564
Coil driver Pod 2 |  |  | D0902744 | S1103563
Coil driver Pod 3 |  |  | D0902744 | S1103356 |  |  |  | Long cables
[7 empty] | Actuator | Coil driver name | Resistance
Stage 1 |  |  |  | Stage 2 |  |  | ST1 H1 | Coil1 Coarse 1 | 7.5
Actuator | Actuator S/N |  |  | Actuator | Actuator S/N |  | ST2 H1 | Coil 1 Fine 1 | 11.3
ST1 - H1 | 24 |  |  | ST2 - H1 | 61 |  | ST2 V1 | Coil 1 Fine 2 | 11.1
ST1 - H2 | 41 |  |  | ST2 - H2 | 57 |  | ST1 V1 | Coil 1 Coarse 2 | 7.5
ST1 - H3 | 42 |  |  | ST2 - H3 | 62 |  | ST1 H2 | Coil 2 Coarse 2 | 7.4
ST1 - V1 | 43 |  |  | ST2 - V1 | 85 |  | ST2 H2 | Coil 2 Fine 1 | 11.1
ST1 - V2 | 56 |  |  | ST2 - V2 | 90 |  | ST2 V2 | Coil 2 Fine 2 | 11.1
ST1 - V3 | 63 |  |  | ST2 - V3 | 68 |  | ST1 V2 | Coil 2 Coarse 2 | 7.4
[7 empty] | ST1 H3 | Coil 3 Coarse 1
Step 2 - Sensors inventory | [6 empty] | ST2 H3 | Coil 3 Fine 1
[7 empty] | ST2 V3 | Coil 3 Fine 2
[7 empty] | ST1 V3 | Coil 3 Coarse 2
Location | Invoice | Gauge | S/N | Probe | S/N | Master/Slave | Phase
ST1-H1 | 101612 | 8800 | 12504 | 2822-V | 12949 | Slave | 180
ST1-H2 | 101406 | 8800 | 12305 | 2822-V | 12864 | Slave | 180
ST1-H3 | 101612 | 8800 | 12512 | 2822-V | 12952 | Slave | 180
ST1-V1 | 101406 | 8800 | 12312 | 2822-V | 12860 | Master | 0
ST1-V2 | 101406 | 8800 | 12307 | 2822-V | 12863 | Slave | 0
ST1-V3 | 101470 | 8800 | 12416 | 2822-V | 12898 | Slave | 0
ST2-H1 | 101406 | 8800 | 12380 | 2822-V | 12867 | Slave | 180
ST2-H2 | 101612 | 8800 | 12378 | 2822-V | 12940 | Slave | 180
ST2-H3 | 101612 | 8800 | 12423 | 2822-V | 12938 | Slave | 180
ST2-V1 | 101612 | 8800 | 12413 | 2822-V | 12939 | Slave | 0
ST2-V2 | 101612 | 8800 | 12381 | 2822-V | 12950 | Slave | 0
ST2-V3 | 101406 | 8800 | 12373 | 2822-V | 12868 | Slave | 0
Geophones GS13 | Serial Number | POD
H1 | Not noted at LLO | 33
H2 | 801 | 31
H3 | 783 | 16
V1 | 740 | 79
V2 | 739 | 94
V3 | 677 | 26
Geophones L4C | Serial Number | POD
H1 | 1105 | 76
H2 | 970 | 30
H3 | 819 | 129
V1 | 1085 | 37
V2 | 1089 | 127
V3 | 929 | 17
location T240 | Serial Number | POD | Cleaning status | Comments
Corner 1 | 112 | 32 | Class A | Initially installed - POD 12 - S/N 138 - Replacement on September 30, 2011
Corner 2 | 118 | 20 | Class A | Initially installed - POD 26 - S/N 133 - Replacement on September 30, 2011
Corner 3 | 101 | 19 | Class A | Initially installed - POD 36 - S/N 109 - Replacement on September 30, 2011
Stage 0-1 |  | Stage 1-2
Lockers | Shim thickness (mil) | Lockers | Shim thickness (mil)
Corner 1 | 0.127" | Corner 1 | 0.126"
Corner 2 | 0.125" | Corner 2 | 0.122" |  | Stage 0-1 Corner 2
Corner 3 | 0.127" | Corner 3 | 0.130"
| Stage 1 |  | Stage 2
D.I at Lockers | Dial indicators V | Dial indicators H | Dial indicators V | Dial indicators H
A
B
C
2011-06-22
| Table locked |  | Table unlocked |  | Difference locked - unlocked
Sensors | Offset (Mean) | Std deviation | Offset (Mean) | Std deviation | Offset (Mean) | mil
ST1 - H1 | -51.3074 | 3.7948 | -135.4831 | 8.9305 | 84.1758 | 0.10021
ST1 - H2 | -830.7471 | 3.8769 | -415.5642 | 9.1031 | -415.1829 | -0.49427
ST1 - H3 | -85.0436 | 5.0374 | 583.7899 | 11.1056 | -668.8335 | -0.79623
ST1 - V1 | -1298.2018 | 6.938 | -479.1497 | 38.5747 | -819.0521 | -0.97506
ST1 - V2 | -2808.0153 | 4.2321 | -2432.8742 | 46.2945 | -375.1411 | -0.4466
ST1 - V3 | 1824.8171 | 4.5407 | 2220.7182 | 34.6239 | -395.901 | -0.47131
ST2 - H1 | -397.052 | 20.4329 | 384.292 | 27.5518 | -781.344 | -0.23254
ST2 - H2 | -257.1653 | 49.9029 | -180.4716 | 21.9436 | -76.6937 | -0.022826
ST2 - H3 | 1898.0853 | 64.9378 | 3874.8316 | 19.1237 | -1976.7462 | -0.58832
ST2 - V1 | -495.6554 | 54.3823 | -2006.7141 | 105.1576 | 1511.0587 | 0.44972
ST2 - V2 | -1115.4252 | 300.798 | 794.0453 | 120.9989 | -1909.4705 | -0.56829
ST2 - V3 | 317.7375 | 456.7042 | -1316.0759 | 98.4998 | 1633.8134 | 0.48625
POD 3
Actuators | ST1 - H3 | ST2 - H3 | ST1 - V3 | ST2 - V3
Anti image pin #
Cable #
Resistance (Ohm)
CD Voltage for 1000 cts offset
Sensor readout (counts) | Negative drive | no drive | Positive drive | Amplitude count | mil
ST1 - H1 | -8198.744 | -137 | 8088.762 | 16287.506 | 19.3899
ST1 - H2 | -7783.0416 | 371 | 8238.5052 | 16021.5468 | 19.0733
ST1 - H3 | -7605.1148 | 591 | 8080.2766 | 15685.3914 | 18.6731
ST1 - V1 | -12246.054 | 932 | 14125.642 | 26371.696 | 31.3949
ST1 - V2 | -12278.238 | 831 | 13963.998 | 26242.236 | 31.2408
ST1 - V3 | -12807.87 | 236 | 13264.386 | 26072.256 | 31.0384
ST2 - H1 | -9014.7884 | 349 | 9714.9982 | 18729.7866 | 5.5743
ST2 - H2 | -9570.823 | -203 | 9176.501 | 18747.324 | 5.5796
ST2 - H3 | -5350.8822 | 3810 | 12971.91 | 18322.7922 | 5.4532
ST2 - V1 | -13391.244 | -1850 | 9687.5986 | 23078.8426 | 6.8687
ST2 - V2 | -10428.046 | 994 | 12389.126 | 22817.172 | 6.7908
ST2 - V3 | -12497.074 | -925 | 10640.88 | 23137.954 | 6.8863
|  | Slope | Offset | Average slope | Variation from average(%)
Stage 1 | ST1 - H1 | 0.54784 | -283.5847 | 0.54321 | 0.85187
| ST1 - H2 | 0.53913 | 533.6354 |  | -0.75133
| ST1 - H3 | 0.54267 | 973.5789 |  | -0.10055
| ST1 - V1 | 0.44043 | 928.5123 | 0.43801 | 0.55269
| ST1 - V2 | 0.43861 | 854.7028 |  | 0.13519
| ST1 - V3 | 0.435 | 296.161 |  | -0.68788
Stage 2 | ST2 - H1 | 0.31328 | 325.3775 | 0.311 | 0.73285
| ST2 - H2 | 0.31438 | -222.7207 |  | 1.0867
| ST2 - H3 | 0.30534 | 3722.8936 |  | -1.8196
| ST2 - V1 | 0.38633 | -1851.1648 | 0.3855 | 0.21589
| ST2 - V2 | 0.38251 | 945.065 |  | -0.77571
| ST2 - V3 | 0.38766 | -878.0218 |  | 0.55983
Pressure (count)
Sensors | Day 1 | Day 2 | Day 3 | Day 4
L4C - H1
L4C - H2
L4C - H3
L4C - V1
L4C - V2
L4C - V3
GS13 - H1
GS13 - H2
GS13 - H3
GS13 - V1
GS13 - V2
GS13 - V3
T240 - 1
T240 - 2
T240 - 3
Test Local drive
Sensors
|  |  | ST1 - H1 | ST1 - H2 | ST1 - H3 | ST1 - V1 | ST1 - V2 | ST1 - V3
| Actuators | ST1 - H1 | 3839.4015 | 1536.2304 | 1556.7577 | 19.2675 | -6.0847 | -10.6515
|  | ST1 - H2 | 1538.7769 | 3774.18 | 1532.4583 | -5.2499 | 11.1966 | -1.4934
|  | ST1 - H3 | 1543.4707 | 1534.8736 | 3810.9373 | 0.02428 | -20.8657 | 23.6317
|  | ST1 - V1 | 73.61 | -134.4073 | 106.3627 | 3095.1253 | -569.5952 | -555.0222
|  | ST1 - V2 | 84.1004 | 31.7862 | -139.615 | -553.2645 | 3076.1235 | -570.7436
|  | ST1 - V3 | -155.652 | 83.6291 | 37.9771 | -564.9577 | -537.8355 | 3050.7204
Sensors
|  |  | ST2 - H1 | ST2 - H2 | ST2 - H3 | ST2 - V1 | ST2 - V2 | ST2 - V3
| Actuators | ST2 - H1 | 2192.2833 | 327.1906 | 346.5634 | -12.3328 | -18.0349 | 21.0718
|  | ST2 - H2 | 326.4375 | 2186.7967 | 348.5802 | -9.8802 | -2.2481 | -6.3785
|  | ST2 - H3 | 326.0596 | 311.1649 | 2161.9204 | -25.2144 | 9.0671 | 5.8317
|  | ST2 - V1 | 69.0606 | 114.2405 | -181.633 | 2716.7698 | 283.4175 | -12.003
|  | ST2 - V2 | -201.8223 | 50.5711 | 123.397 | -26.7288 | 2675.8949 | 286.5231
|  | ST2 - V3 | 120.4239 | -208.6703 | 96.9326 | 294.327 | -7.1388 | 2704.8183
Test Cartesian drive
|  |  | ST1 - X | ST1 - Y | ST1 - Z | ST1 - RY | ST1 - RY | ST1 - RZ
| Actuators | ST1 - X | 1529.1023 | -1.4771 | 12.229 | -8.194 | -19.6608 | 17.8967
|  | ST1 - Y | -27.78 | 1504.4793 | 1.5824 | 19.3693 | -21.4461 | -9.3654
|  | ST1 - Z | 5.6147 | -13.1876 | 651.9252 | 1.2738 | -17.2589 | -8.5615
|  | ST1 - RX | -1.7511 | 314.9101 | -2.396 | 2656.0212 | -5.3924 | -14.5957
|  | ST1 - RY | -327.0599 | 22.7615 | 6.5228 | 4.5257 | 2635.3295 | -5.612
|  | ST1 - RZ | 13.9518 | -7.6467 | 18.0473 | -2.7298 | -2.3543 | 2900.3807
|  |  | ST2 - X | ST2 - Y | ST2 - Z | ST2 - RY | ST2 - RY | ST2 - RZ
| Actuators | ST2 - X | 1233.6156 | 7.9934 | -5.9151 | -3.0754 | 35.0252 | -14.272
|  | ST2 - Y | 11.0918 | 1229.9099 | -11.9159 | -21.3372 | -8.3907 | 19.1794
|  | ST2 - Z | 0.367 | 1.6676 | 983.549 | -8.8618 | 13.434 | -2.8976
|  | ST2 - RX | 3.2994 | -12.3468 | -20.7214 | 3939.399 | 12.2377 | -5.2722
|  | ST2 - RY | 9.4318 | 8.775 | 5.0151 | 24.479 | 3972.9052 | 4.9888
|  | ST2 - RZ | 1.3104 | 12.9644 | 2.8611 | -0.0706 | -6.315 | 2350.7464
Sensors
|  |  | ST1 - H1 | ST1 - H2 | ST1 - H3 | ST1 - V1 | ST1 - V2 | ST1 - V3
| Actuators | ST1 - X | 1547.2967 | -750.2317 | -745.3066 | 26.8776 | -0.63632 | -14.8269
|  | ST1 - Y | 40.83 | 1285.7541 | -1312.1095 | -5.2067 | 13.0678 | -11.7276
|  | ST1 - Z | -1.1149 | -18.482 | -7.9125 | 656.0633 | 644.6923 | 638.3116
|  | ST1 - RX | -15.0559 | 122.8436 | -149.177 | -2548.4121 | 2160.6831 | 393.7198
|  | ST1 - RY | -179.5587 | 74.5157 | 67.3994 | -1005.5493 | -1691.9764 | 2709.2365
|  | ST1 - RZ | 2858.1957 | 2773.3435 | 2786.9651 | 15.5698 | -5.6356 | 16.963
Sensors
|  |  | ST2 - H1 | ST2 - H2 | ST2 - H3 | ST2 - V1 | ST2 - V2 | ST2 - V3
| Actuators | ST2 - X | 656.9803 | -1254.4048 | 610.3116 | -15.4786 | -6.7226 | 64.088
|  | ST2 - Y | 1050.3739 | 3.5851 | -1048.5486 | 29.9512 | -13.9501 | 30.3799
|  | ST2 - Z | -16.5404 | 3.6732 | 4.8856 | 990.7079 | 971.4022 | 1018.2312
|  | ST2 - RX | -269.5172 | 4.0752 | 261.1008 | -2295.32 | 2315.8228 | -6.052
|  | ST2 - RY | 180.9198 | -303.8515 | 157.3736 | -1403.3786 | -1294.3411 | 2737.8032
|  | ST2 - RZ | 1627.6515 | 1646.543 | 1609.5914 | 20.4904 | 2.268 | 37.1119
ITMY and FMY installed
Measurement on November 6, 2011
Stage 2
Mass Budget | Quantity | Weight | Unit
lb
lb
Kg
1 type 4 |  |  | Kg
1type 3 |  |  | Kg
2 type 2 + 1 type 3 |  |  | Kg
| Quantity | Weight |  | Unit
|  |  | 0 | Kg
|  |  | 0 | Kg
|  |  | 0 | Kg
|  |  | 0 | Kg
|  |  | 0 | Kg
|  |  | 0 | Kg
|  |  | 0 | Kg
Stage 1
Mass Budget
D0902612 | 12 | lb
D0902614 | 10 | lb
D0902615 | 10 | lb
Stage 1
| Location | Weight (lb) | Weight (Kg)
| C1-1 |  | 0
| C1-2 |  | 0
| C2-1 |  | 0
| C2-2 |  | 0
| C3-1 |  | 0
| C3-2 |  | 0
| Total | 0 | 0
Actuator | Gain
ST1 ACT H1 | 1.0086
ST1 ACT H2 | 1
ST1 ACT H3 | 0.99159
ST1 ACT V1 | 0.98856
ST1 ACT V2 | 1.0017
ST1 ACT V3 | 1.01
ST2 ACT H1 | 0.98964
ST2 ACT H2 | 0.98953
ST2 ACT H3 | 1.0215
ST2 ACT V1 | 1.0015
ST2 ACT V2 | 1.008
ST2 ACT V3 | 0.99069
LHO_ISI_Locked/Unlocked_2011_12_06
| Table locked |  | Table unlocked |  | Difference locked - unlocked
Sensors | Offset (Mean) | Std deviation | Offset (Mean) | Std deviation | Offset (Mean) | mil
ST1 - H1 | 260.6347 | 36.6358 | 417.3063 | 62.7129 | -156.6716 | -0.18651
ST1 - H2 | -218.9695 | 22.3623 | 106.1934 | 40.9494 | -325.1629 | -0.3871
ST1 - H3 | -100.5439 | 16.4311 | 569.4021 | 73.0452 | -669.946 | -0.79755
ST1 - V1 | -160.9417 | 75.5323 | 78.7349 | 71.4276 | -239.6766 | -0.28533
ST1 - V2 | -60.9144 | 138.5216 | -312.7474 | 92.9929 | 251.8331 | 0.2998
ST1 - V3 | -112.0162 | 46.7907 | 287.2863 | 69.3474 | -399.3024 | -0.47536
ST2 - H1 | -138.1438 | 25.2875 | 590.9577 | 98.2808 | -729.1014 | -0.21699
ST2 - H2 | 857.08 | 66.9278 | 835.4156 | 97.711 | 21.6644 | 0.0064478
ST2 - H3 | -345.2713 | 48.3796 | 836.2346 | 59.0037 | -1181.5059 | -0.35164
ST2 - V1 | 923.6051 | 35.6359 | 1020.2103 | 209.1094 | -96.6052 | -0.028752
ST2 - V2 | 706.334 | 55.7364 | 468.9167 | 244.1775 | 237.4172 | 0.07066
ST2 - V3 | -470.0898 | 54.0443 | -147.2069 | 203.4305 | -322.8829 | -0.096096
LHO_ISI_BSC8_Offset_Local_Drive_20111206 |  |  | Sensors
|  |  | ST1 - H1 | ST1 - H2 | ST1 - H3 | ST1 - V1 | ST1 - V2 | ST1 - V3
| Actuators | ST1 - H1 | 3836.2426 | 1536.7107 | 1534.7291 | 19.2029 | -8.1177 | -9.3374
|  | ST1 - H2 | 1519.0842 | 3756.4487 | 1509.0049 | -0.35737 | 15.3372 | -4.2657
|  | ST1 - H3 | 1517.1078 | 1515.1597 | 3756.0223 | 2.4477 | -4.8506 | 12.6639
|  | ST1 - V1 | 58.7943 | -129.7271 | 87.0092 | 2892.0888 | -495.9877 | -468.2862
|  | ST1 - V2 | 77.7509 | 29.8422 | -143.9458 | -479.4923 | 2909.3199 | -499.6949
|  | ST1 - V3 | -145.5734 | 79.4754 | 15.9566 | -505.7681 | -467.4016 | 2917.8109
Sensors
|  |  | ST2 - H1 | ST2 - H2 | ST2 - H3 | ST2 - V1 | ST2 - V2 | ST2 - V3
| Actuators | ST2 - H1 | 2145.6685 | 317.8281 | 323.3029 | -26.9044 | -17.2834 | -16.7466
|  | ST2 - H2 | 321.39 | 2146.8099 | 318.3149 | -18.9359 | -13.4734 | -26.7086
|  | ST2 - H3 | 337.6114 | 333.1025 | 2237.2277 | -20.2364 | 23.9054 | -6.9299
|  | ST2 - V1 | 61.9415 | 109.7567 | -174.0402 | 2609.1783 | 369.9721 | 3.9191
|  | ST2 - V2 | -184.005 | 58.929 | 88.5605 | 10.1274 | 2606.2891 | 303.2543
|  | ST2 - V3 | 105.5961 | -178.2662 | 61.138 | 298.8647 | 24.7071 | 2584.691
Test Cartesian drive
LHO_ISI_BSC8_Offset_Cartesian_Drive_20111206 |  |  | ST1 - X | ST1 - Y | ST1 - Z | ST1 – RX | ST1 - RY | ST1 - RZ
| Actuators | ST1 - X | 1522.9321 | 1.9773 | -7.1407 | -12.2292 | -6.8932 | 30.9852
|  | ST1 - Y | -7.4661 | 1505.6627 | -24.8754 | 11.3241 | -9.3006 | 3.3554
|  | ST1 - Z | 26.5878 | -0.73838 | 656.4688 | 5.4675 | 4.6723 | -6.7045
|  | ST1 - RX | 5.4514 | 305.0242 | -0.031414 | 2469.0794 | -6.2807 | -10.2663
|  | ST1 - RY | -290.8652 | 18.9105 | -2.4745 | -0.52794 | 2472.2221 | -8.6027
|  | ST1 - RZ | 13.7258 | 1.7791 | -1.3828 | 1.3841 | -0.46478 | 2884.3424
|  |  | ST2 - X | ST2 - Y | ST2 - Z | ST2 – RX | ST2 - RY | ST2 - RZ
| Actuators | ST2 - X | 1228.7257 | -18.7373 | 56.8693 | 24.6945 | 5.0336 | 14.7167
|  | ST2 - Y | -0.35774 | 1233.9523 | -26.9807 | 6.5185 | 1.2663 | -41.7523
|  | ST2 - Z | 9.0591 | 2.9783 | 963.8483 | 10.2115 | 14.6222 | -3.0406
|  | ST2 - RX | 2.914 | -6.4512 | -5.4258 | 3754.7698 | -1.3767 | -11.4681
|  | ST2 - RY | 20.7969 | 2.3751 | -14.6349 | 11.3522 | 3747.4952 | -1.8739
|  | ST2 - RZ | 12.4311 | -32.598 | -26.8202 | -4.958 | 15.8083 | 2340.9977
Sensors
|  |  | ST1 - H1 | ST1 - H2 | ST1 - H3 | ST1 - V1 | ST1 - V2 | ST1 - V3
| Actuators | ST1 - X | 1548.0186 | -730.4174 | -736.0559 | 27.5599 | -0.65736 | -7.3226
|  | ST1 - Y | -25.1542 | 1298.1793 | -1271.8197 | 12.2489 | 18.8722 | -0.48838
|  | ST1 - Z | -47.6979 | -6.8933 | -22.6215 | 648.5937 | 657.4534 | 656.9731
|  | ST1 - RX | -17.0849 | 127.4245 | -145.6122 | -2367.1164 | 2015.164 | 369.4109
|  | ST1 - RY | -180.7292 | 78.3375 | 58.0421 | -936.4675 | -1580.2124 | 2541.0495
|  | ST1 - RZ | 2786.7548 | 2757.6193 | 2763.5365 | 4.6077 | -3.6263 | -5.272
Sensors
|  |  | ST2 - H1 | ST2 - H2 | ST2 - H3 | ST2 - V1 | ST2 - V2 | ST2 - V3
| Actuators | ST2 - X | 616.7627 | -1210.9608 | 630.1197 | -35.6041 | -0.75188 | 23.2471
|  | ST2 - Y | 1055.6075 | 0.96914 | -1082.5165 | -15.3823 | 8.0785 | 13.4015
|  | ST2 - Z | 6.3389 | 7.5378 | 4.1295 | 948.4691 | 979.0128 | 971.9798
|  | ST2 - RX | -246.5687 | -2.7122 | 237.742 | -2197.7529 | 2226.9676 | -38.2856
|  | ST2 - RY | 144.7739 | -268.1452 | 151.7261 | -1340.6468 | -1210.6724 | 2546.2989
|  | ST2 - RZ | 1612.0863 | 1612.8644 | 1661.4545 | -18.2668 | 17.8138 | 21.4816
Range of motion | LHO_ISI_BSC8_Range_Of_Motion_20111206
Sensor readout (counts) | Negative drive | no drive | Positive drive | Amplitude count | mil
ST1 - H1 | -15865.252 | 357 | 16290.242 | 32155.494 | 38.2803
ST1 - H2 | -16123.448 | 21 | 15848.596 | 31972.044 | 38.062
ST1 - H3 | -15595.782 | 562 | 15303.954 | 30899.736 | 36.7854
ST1 - V1 | -12524.706 | -30 | 12482.594 | 25007.3 | 29.7706
ST1 - V2 | -12792.768 | -354 | 12069.422 | 24862.19 | 29.5978
ST1 - V3 | -12064.468 | 321 | 12667.594 | 24732.062 | 29.4429
ST2 - H1 | -8793.5532 | 539 | 9846.2876 | 18639.8408 | 5.5476
ST2 - H2 | -8511.8832 | 814 | 10166.036 | 18677.9192 | 5.5589
ST2 - H3 | -8430.9066 | 800 | 10025.4934 | 18456.4 | 5.493
ST2 - V1 | -10134.7426 | 983 | 12128.332 | 22263.0746 | 6.6259
ST2 - V2 | -10512.31 | 512 | 11491.758 | 22004.068 | 6.5488
ST2 - V3 | -11251.53 | -95 | 11035.326 | 22286.856 | 6.633
Pressure Sensor | LHO_ISI_BSC8_Pressure_Sensors_Check_20111206
Pressure (counts)
Sensors | Corner 1 | Corner 2 | Corner 3
ST1-L4C-D | 37 | 124 | 231
ST1-L4C-P | 24642 | 24647 | 24589
ST1-GS13-D | 625 | 1165 | 4065
ST1-GS13-P | 24425 | 24324 | 24266
ST1-T240-P | 14073 | 13642 | 13829
State | GPS Time
ISI with the Damping Loops | 1007316891
ISI Unlocked | 1007321196
ISI Locked | 1007334800
| Time | Tilt | Offset drive
ISI Tilted 1 | 1007327576 | +RX | 40000
ISI Tilted 2 | 1007327824 | -RX | -40000
ISI Tilted 3 | 1007328175 | +RY | 32000
ISI Tilted 4 | 1007328449 | -RY | -32000
ISI Tilted 5 | 1007328907 |  | 25K RX-25KRY
ISI Tilted 6 | 1007329120 |  | 25K RX+25KRY
Measurement on November 6, 2011
Stage 2
Mass type | Quantity | Total Weight (lb) | Total Weight (Kg) | Location
D1003136 | 12 | 600 | 272.154 | On Keel
D071200-06 | 1 | 27.2 | 12.3376 | On Keel
D071200: 6,5,5;4,4,3,2,2,2;1,1,1,0 | 13 | 89.2 | 40.4602 | Corner1 SideWall
D071200:6,3,3,1,0;5,1;1 | 8 | 55.7 | 25.265 | Corner2 SideWall
D071200:-;4,4;- | 2 | 9 | 4.0823 | Corner3 SideWall
| Total | 781.1 | 354.2991
Stage 2
Mass Type | Quantity | Weight (Kg) | Total weight (Kg) | Total weight (lb) | Location
D1003163 | 6 | 10 | 60 | 132.278 | On Optical Table
D1003164 | 6 | 5 | 30 | 66.139 | On Optical Table
D1003123 | 6 | 3 | 18 | 39.6834 | On Optical Table
D1003166 | 6 | 1 | 6 | 13.2278 | On Optical Table
D1003167 | 6 | 0.5 | 3 | 6.6139 | On Optical Table
D1003168 | 12 | 0.25 | 3 | 6.6139 | On Optical Table
|  | Total | 120 | 264.5561
Stage 1
D0902612 | 12 | lb
D0902614 | 10 | lb
D0902615 | 10 | lb
Stage 1
Mass Type | Location | Weight (lb) | Weight (Kg)
D0902612 | C1-1 | 12 | 5.496
D0902616-1,2,2,3 | C1-2 | 21.1 | 9.6638
D0902612 | C2-1 | 12 | 5.496
D0902616--1,1,2 | C2-2 | 11.1 | 5.0838
D0902612 | C3-1 | 12 | 5.496
D0902616--5 | C3-2 | 20.1 | 9.2058
| Total | 88.3 | 40.4414
Linearity TEST
| 0.54127 | 388.8773
| Inf | 1.4961
| 0.53906 | 594.9519
| 0.41658 | 86.2748
| 0.41424 | -270.1989
| 0.41311 | 382.7748
| 0.31152 | 589.3753
| 0.31224 | 833.2204
| 0.30772 | 803.2192
| 0.37255 | 1077.3245
| 0.36836 | 509.6524
| 0.37252 | -99.9903
|  | Slope | Offset | Average slope | Variation from average(%)
Stage 1 | ST1 - H1 | 0.54127 | 388.8773 | 0.54017 | 0.20324
ST1 - H2
| ST1 - H3 | 0.53906 | 594.9519 |  | -0.20324
| ST1 - V1 | 0.41658 | 86.2748 | 0.41464 | 0.46653
| ST1 - V2 | 0.41424 | -270.1989 |  | -0.097573
| ST1 - V3 | 0.41311 | 382.7748 |  | -0.36896
Stage 2 | ST2 - H1 | 0.31152 | 589.3753 | 0.31049 | 0.33055
| ST2 - H2 | 0.31224 | 833.2204 |  | 0.56214
| ST2 - H3 | 0.30772 | 803.2192 |  | -0.89268
| ST2 - V1 | 0.37255 | 1077.3245 | 0.37114 | 0.37906
| ST2 - V2 | 0.36836 | 509.6524 |  | -0.75053
| ST2 - V3 | 0.37252 | -99.9903 |  | 0.37148