LIGO Laboratory / LIGO Scientific Collaboration

LIGO- E1100300 	                 LIGO                   May 16th, 2014

aLIGO BSC-ISI, Pre-integration Testing report,

Unit 7 - Phase I (post-assembly)

E1100300 – V2

Jim Warner

Distribution of this document:

Advanced LIGO Project

This is an internal working note

of the LIGO Laboratory

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 1970 Mail Stop S9-02 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

Table of contents:

I. Pre-Assembly Testing	4

II. Tests to be performed during assembly	6

III. Tests to perform after assembly	14

IV. BSC-ISI testing Summary	36

Introduction

The BSC-ISI testing is performed in three phases:

1) BSC-ISI, Pre-integration Testing, Phase I (post-assembly)

2) BSC-ISI, Pre-integration Testing, Phase II: Tests done after Transport (and possible storage), during mating phase with Suspensions, before insertion.

3) BSC-ISI, Integration Phase Testing: Procedure and results related to the commissioning in the chamber.

This document presents the series of tests (Phase I) performed on the fourth BSC-ISI assembled at LHO.

The testing procedure document E1000486-v5 was used.

All results are posted on the SVN at:

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/X1/Unit_7/

The following type of document can be found in the SVN:

- Excel spreadsheet (.xls)

- Data location

- Figures location

- Masses distribution scheme (ppt)

# Pre-Assembly Testing

## Step 1 - CPS Test and calibration – E1100369

CPS sensors are tested (calibration and noise test) at MIT before being cleaned and baked at LHO. The list of installed sensors used for testing (phase I) are reported in step II.3.

All data related to the CPS testing can be found in the SVN at

/svn/seismic/Common/Data/aLIGO_BSC_ISI_CPS/

Test result:	Passed:   X   	Failed:         .	Waived :.     .

## Step 2 - GS13 – Inspection/Assembly – E1000058 – E1100740

GS13 are tested and podded at LLO.

We had several GS-13s fail on this unit. Initially, 1h and 1v, then another vertical. A third vertical failed after a transfer function with too high drive.

The list of installed sensors used for testing (phase I) are reported in step II.3.

All the data related to GS-13 post podding testing can be found in the SVN at : /svn/seismic/Common/Data/aLIGO_GS13_TestData/PostMod_TestResults_PDFs

E1000058 and E1100740 spreadsheets provide the status of each individual GS-13 at LLO site for HAM-ISI and BSC-ISI and the installation location of the geophones.

Test result:	Passed:   X   	Failed:         .	Waived :.     .

## Step 3 - L4C – Inspection/Assembly – E1000136 – E1100740

L4C are tested and podded at LLO. The list of installed sensors used for testing (phase I) are reported in step II.

One vertical L4C failed.

All the data related to L4C post podding testing can be found in the SVN at : /svn/seismic/Common/Data/aLIGO_L4C_TestData/TestResults_PDFs/

Test result:	Passed:   X   	Failed:         .	Waived :.     .

## Step 4 - T240 – Inspection/Assembly - E1100326 – E1100740

T240 are tested and podded at LLO. We haven’t had to replace the T240s on this Unit, and these are the ones with the new Voltage Regulator, it seems that they are working fine and keep the pressure sensor from dying. The list of installed sensors used for testing (phase I) are reported in step II.3.

All the data related to T240 post podding testing can be found in the SVN at :

seismic/Common/Data/aLIGO_T240_TestData/AsReceived_TestResults_PDFs.

E1100326 and E1100740 spreadsheets provide the status of each individual T240 at LLO site for BSC-ISI and the installation location of the geophones.

Test result:	Passed:   X   	Failed:         .	Waived :.     .

## Step 5 - Actuators - T0900564 - T1100234 – E1100741

The list of installed sensors used for testing (phase I) are reported in step II.2

Large actuators data can be found at: T0900564. Actuator inventory is made at Section II – Step 2.

Small actuators data can be found at: T1100234. Actuator inventory is made at Section II – Step 2.

Test result:	Passed:   X   	Failed:         .	Waived :.     .

# Tests to be performed during assembly

## Step 1 - Test stand level

The Test Stand was transformed and re-leveled to dock a BSC-ISI.

Test result:	Passed:   X   	Failed:         .	Waived :.     .

## Step 2 - Actuators Inventory

The actuators S/N are reported in the table below. Further information can be found in T0900564 and T1100234.

| Stage 1 | Stage 2 |
| --- | --- |
| Actuator | Actuator S/N | Actuator | Actuator S/N |
| ST1 - H1 | 192 | ST2 - H1 | 86 |
| ST1 - H2 | 180 | ST2 - H2 | 27 |
| ST1 - H3 | 169 | ST2 - H3 | 89 |
| ST1 - V1 | 113 | ST2 - V1 | 81 |
| ST1 - V2 | 111 | ST2 - V2 | 4 |
| ST1 - V3 | 168 | ST2 - V3 | 88 |

Table 1 - Actuators' inventory

Test result:	Passed:   X   	Failed:         .	Waived :.     .

## Step 3 - Sensors Inventory

The sensors S/N are reported in the table below.

| CPS Stage 1 | CPS S/N | ADE board serial # |
| --- | --- | --- |
| H1 | 12957 |  |
| H2 | 12893 |  |
| H3 | 13189 |  |
| V1 | 13413 |  |
| V2 | 13455 |  |
| V3 | 13429 |  |

Table 2 - Capacitive position sensors' inventory – Stage 1

| CPS Stage 2 | CPS S/N | ADE board serial # |
| --- | --- | --- |
| H1 | 13424 |  |
| H2 | 12862 |  |
| H3 | 13434 |  |
| V1 | 13200 |  |
| V2 | 12907 |  |
| V3 | 13442 |  |

Table 3 - Capacitive position sensors' inventory – Stage 2

| Geophones GS13 | Serial Number | POD |
| --- | --- | --- |
| H1 |  | 30 |
| H2 |  | 502 |
| H3 |  | 19 |
| V1 |  | 35 |
| V2 |  | 501 |
| V3 |  | 6 |

Table 4 - GS13 inventory

| Geophones L4C | Serial Number | POD |
| --- | --- | --- |
| H1 |  | 23 |
| H2 |  | 24 |
| H3 |  | 56 |
| V1 |  | 47 |
| V2 |  | 72 |
| V3 |  | 45 |

Table 5 - L4C inventory

| Geophones T240 | Serial Number | POD |
| --- | --- | --- |
| 1 |  | 300 |
| 2 |  | 33 |
| 3 |  | 26 |

Table 6 - T240 inventory

Test result:	Passed:   X   	Failed:         .	Waived :.     .

## Step 4 - Electronics Inventory

Write down in the table below all serial numbers all the electronic equipment:

| Hardware | Ligo reference | S/N |
| --- | --- | --- |
| Interface Chassis - Corner 1 | D1002432 | S110223 |
| Interface Chassis - Corner 2 |  | S1102224 |
| Interface Chassis - Corner 3 |  | S1102218 |
| Anti-Alliasing Chassis - Corner 1 | D1002693 | S1102693 |
| Anti-Alliasing Chassis - Corner 2 |  | S1102694 |
| Anti-Alliasing Chassis - Corner 3 |  | S1102679 |
| Anti-image Chassis | D070081 | S1000250 |
| Binary Input Chassis | D1001726 | S1101309 |
| Binary Input Chassis |  | S11031308 |
| Binary Output Chassis | D1001728 | S1101347 |
| T240 Interface - Corner 1 | D1002694 | S1101040 |
| T240 Interface - Corner 2 |  | S1101838 |
| T240 Interface - Corner 3 |  | S1101839 |
| I/O Chassis | n/a |  |
| Coil driver Pod 1 | D0902744 | S1000266 |
| Coil driver Pod 2 |  | S1000269 |
| Coil driver Pod 3 |  | S110692 |

Table 7 - Electronic equipment

Note:

Test result:	Passed:   X   	Failed:         .	Waived :.     .

## Step 5 - Check level of Stage 0 after top-bottom plate assembly

Note: This test has not been performed. The test stand was leveled before assembly began.

Test result:	Passed:      	Failed:         .	Waived :.   X   .

## Step 6 - Check gaps under the blade posts

Test result:	Passed:   X   	Failed:         .	Waived :.     .

## Step 7 - Blade post shim thickness

This table shows the shims thickness installed under the lockers.

| Stage 0-1 | Stage 1-2 |
| --- | --- |
| Lockers | Shim thickness (mil) | Lockers | Shim thickness (mil) |
| Corner 1 | 119 | Corner 1 | 117 |
| Corner 2 | 125 | Corner 2 | 118 |
| Corner 3 | 130 | Corner 3 | 120 |

Table 8 - Shims thickness

Acceptance criteria: Both D0901805 Stage 0-1 Locker Shims & D0902551 Stage 1-2 Locker Shims goes from .110” up to .130” with an increment of .001”.

Test result:	Passed:   X   	Failed:         .	Waived :.     .

## Step 8 - Blade 0-1 post launch angle

Test result:	Passed:      	Failed:         .	Waived :.   X  .

## Step 9 - Gap checks on actuators

Test result:	Passed:   X   	Failed:         .	Waived :.     .

## Step 10 - Mass budget

Note: The second version of the blade spacers was used. Consequently, the additional payload is expected to be close from design.

Six vibration absorbers were installed on stage 1. Masses on stage 2 are resting on Viton pads.

Stage 1:
The stage 1 payload is reported in the table below:

| Corner 1 | 2x vib abs |  |  |
| --- | --- | --- | --- |
| Corner 2 | 2x vib abs |  |  |
| Corner 3 | 2x vib abs |  |  |
| Stage 1 |
| Location | Weight (lb) |
| Corner 1 | 1xt1,2xt4,3xt5,1x2613 |
| Corner 2 | 1xt1,1xt2,1xt4,1x2613 |
| Corner 3 | 5xt5,1xt3 |
| Total |  |

Table 9 - Payload Stage 1

Nominal payload on stage 1: 109Kg – 240lb
Nominal mass of stage 1=916Kg - 2019lb

Stage 2:

The stage 2 payload is reported in the table below:

| Mass Budget | Quantity | Weight | Unit | Weight (lb) |
| --- | --- | --- | --- | --- |
|  | 3 | 610 | lb | 1830 |
|  | 2 | 233 | lb | 466 |
| type 0 | 3 |  | lb |  |
| type 1 | 7 |  | lb |  |
| type 2 | 1 |  | lb |  |
| type 3 |  |  | lb |  |
| type 4 | 1 |  | lb |  |
| type 5 | 3 |  | lb |  |
| type 6 | 5 |  | lb |  |

Table 10 - Payload Stage 2

Nominal payload: 1183.4Kg – 2609lb
Total nominal mass of Stage 2: 2913.9Kg – 6424lb
Additional stage 2 payload is 30lb lighter than the design.

Error mass on stage 0-1 blades: -(30+107)/(6424+2019) = -1.6%

Test result:	Passed:   X   	Failed:         .	Waived :.     .

## Step 12 – Cables inventory – E1100822

The final Class A cables have been used for the testing of this Unit.

| DCC Number | Description | Serial Number | Location | Inventory date | Tested |
| --- | --- | --- | --- | --- | --- |
| D1100154 | 25-pin M-to-two 9-pin F straight |  | L4C corner2 |  | YES |
| D1100155 | 25-pin M-to-two 9-pin F straight |  | GS-13 corner2 |  | YES |
| D1100148 | 2-wire, 14awg 3-pin M to 3-pin F |  | Act St1 V3 |  | YES |
| D1100148 | 2-wire, 14awg 3-pin M to 3-pin F |  | Act St2 V2 |  | YES |
| D1100148 | 2-wire, 14awg 3-pin M to 3-pin F |  | Act St2 H3 |  | YES |
| D1100148 | 2-wire, 14awg 3-pin M to 3-pin F |  | Act St2 H2 |  | YES |
| D1100148 | 2-wire, 14awg 3-pin M to 3-pin F |  | Act St1 H2 |  | YES |
| D1100150 | 2-wire, 14awg 2 pins to 3-pin F |  | Act St1 H3 |  | YES |
| D1100150 | 2-wire, 14awg 2 pins to 3-pin F |  | Act St1 V3 |  | YES |
| D1100150 | 2-wire, 14awg 2 pins to 3-pin F |  | Act St1 V2 |  | YES |
| D1100150 | 2-wire, 14awg 2 pins to 3-pin F |  | Act St1 V1 |  | YES |
| D1100150 | 2-wire, 14awg 2 pins to 3-pin F |  | Act St1 H2 |  | YES |
| D1100150 | 2-wire, 14awg 2 pins to 3-pin F |  | Act St1 H1 |  | YES |
| D1100151 | 2-wire, 14awg 2 pins to 3-pin F |  | Act St2 H2 |  | YES |
| D1100151 | 2-wire, 14awg 2 pins to 3-pin F |  | Act St2 V2 |  | YES |
| D1100151 | 2-wire, 14awg 2 pins to 3-pin F |  | Act St2 H3 |  | YES |
| D1100151 | 2-wire, 14awg 2 pins to 3-pin F |  | Act St2 V3 |  | YES |
| D1100151 | 2-wire, 14awg 2 pins to 3-pin F |  | Act St2 V1 |  | YES |
| D1100151 | 2-wire, 14awg 2 pins to 3-pin F |  | Act St H1 |  | YES |
| D1100152 | 25-pin F-to-25-pin F |  | T240 corner3 |  | YES |
| D1100152 | 25-pin F-to-25-pin F |  | T240 corner1 |  | YES |
| D1100152 | 25-pin F-to-25-pin F |  | T240 corner2 |  | YES |
| D1100153 | 25-pin F-to-25-pin F |  | L4C corner2 |  | YES |
| D1100153 | 25-pin F-to-25-pin F |  | GS-13 corner3 |  | YES |
| D1100153 | 25-pin F-to-25-pin F |  | L4C corner3 |  | YES |
| D1100153 | 25-pin F-to-25-pin F |  | GS-13 corner2 |  | YES |
| D1100153 | 25-pin F-to-25-pin F |  | L4C corner1 |  | YES |
| D1100153 | 25-pin F-to-25-pin F |  | GS-13 corner1 |  | YES |
| D1100148 | 2-wire, 14awg 3-pin M to 3-pin F |  | Act St2 V3 ext |  | YES |
| D1100148 | 2-wire, 14awg 3-pin M to 3-pin F |  | Act St2 H1 ext |  | YES |
| D1100148 | 2-wire, 14awg 3-pin M to 3-pin F |  | Act St1 V1 |  | YES |
| D1100148 | 2-wire, 14awg 3-pin M to 3-pin F |  | Act St1 H3 ext |  | YES |
| D1100148 | 2-wire, 14awg 3-pin M to 3-pin F |  | Act St1 H1 ext |  | YES |
| D1100148 | 2-wire, 14awg 3-pin M to 3-pin F |  | Act St1 V2 ext |  | YES |
| D1100148 | 2-wire, 14awg 3-pin M to 3-pin F |  | Act St2 V1 ext |  | YES |
| D1100154 | 25-pin M-to-two 9-pin F straight |  | L4C corner1 ext |  | YES |
| D1100155 | 25-pin M-to-two 9-pin F straight |  | GS-13 corner1 |  | YES |
| D1100155 | 25-pin M-to-two 9-pin F straight |  | GS-13 corner3 |  | YES |
| D1100154 | 25-pin M-to-two 9-pin F straight |  | L4C corner3 |  | YES |

Test result:	Passed:   X   	Failed:         .	Waived:.     .

## Step 13 - Cable routing

The final Class A cables have been used for the testing of this Unit.

The cabling has been done following E1101027 aLIGO BSC-ISI Cable Routing Manual.

Test result:	Passed:   X   	Failed:         .	Waived :.     .

# Tests to perform after assembly

## Step 1- Geophones pressure readout

Raw pressure measured by the geophones is reported in the table below:

|  | Pressure (kPa) |
| --- | --- |
| Sensors | Corner 1 | Corner 2 | Corner 3 |
| ST1-L4C-P |  |  |  |
| ST1-L4C-D |  |  |  |
| ST1-GS13-P |  |  |  |
| ST1-GS13-D |  |  |  |
| ST1-T240-P |  |  |  |

Table 11 - Raw Pressure

A screenshot of the MEDM pressure screen is saved in the Misc directory for Unit 6

Note: The T240’s chassis has the old gain for the pressure sensors.

Test result:	Passed:   X   	Failed:         .	Waived :.     .

## Step 2- Set up sensors gap – Locked vs unlocked position

During this step, sensors gap are adjusted. This step considers that the lockers have been finely setup during assembly.

|  | Table locked | Table unlocked | Difference locked - unlocked |
| --- | --- | --- | --- |
| Sensors | Offset (Mean) | Std deviation | Offset (Mean) | Std deviation | Offset (Mean) | mil |
| ST1 - H1 | -101 | 14 | 54 | 38 | -156 | -0.19 |
| ST1 - H2 | 70 | 7 | 509 | 18 | -440 | -0.52 |
| ST1 - H3 | -234 | 15 | -138 | 23 | -96 | -0.11 |
| ST1 - V1 | -459 | 18 | -649 | 26 | 190 | 0.23 |
| ST1 - V2 | -109 | 13 | -56 | 33 | 53 | -0.06 |
| ST1 - V3 | -550 | 9 | -250 | 32 | -300 | -0.36 |
| ST2 - H1 | 1133 | 55 | 1125 | 25 | 8 | 0 |
| ST2 - H2 | 1669 | 43 | 1638 | 34 | 31 | 0.01 |
| ST2 - H3 | 1160 | 43 | 1319 | 22 | -159 | -0.05 |
| ST2 - V1 | -134 | 43 | -188 | 81 | 54 | 0.02 |
| ST2 - V2 | -88 | 59 | -159 | 90 | 71 | 0.02 |
| ST2 - V3 | 454 | 40 | 536 | 84 | 82 | -0.02 |

Table 12 - Capacitive position sensors readout after gap set-up

Acceptance criteria:

- In the locked position, all mean values must be lower than 400 counts for stage 1 CPS and 1600 counts for stage 2 CPS on Dataviewer (a bit less than .0005”).

- In the locked position, all standard deviations below 25 counts for stage 1, 100 counts for stage 2

- Absolute values of the difference between the unlocked and the locked table must be below:

Stage 1

  - 1600 cts for horizontal sensors (~0.002”)

  - 1600 cts for vertical sensors (~0.002”)

Stage 2

  - 6500 cts for horizontal sensors (~0.002”)

  - 6500 cts for vertical sensors (~0.002”)

- Considering the acceptance criteria of step 2, all mean values must be lower than

Stage 1

  - 2000 cts for horizontal sensors (~0.0025”)

  - 2000 cts for vertical sensors (~0.0025”)

Stage 2

  - 8000 cts for horizontal sensors (~0.0025”)

  - 8000 cts for vertical sensors (~0.0025”)

Note: The CPS will be readjusted once SUS will installed

Test result:	Passed:   X   	Failed:         .	Waived :.      .

## Step 3 - Measure the Sensor gap

Test Failure mitigation:

This test was not performed. The sensor gaps have not been measured. These sensors have already been tested at LASTI. Moreover, risks of scratching the target are so high that we preferred not performing this test. In the future, this test will be removed from the testing procedure.

Test result:	Passed:      	Failed:         .	Waived :.   X   .

## Step 4- Performance of the limiters

## Step 4.1 - Test Nº1 - Push “in the general coordinates Z/RZ”

Test result:	Passed:      	Failed:         .	Waived :.  X   .

## Step 4.2 - Test Nº2 – Push “locally”

| Sensors | Push in negative direction | Push in positive direction | Mil | Mil | Railing | Actuator Gap Check |
| --- | --- | --- | --- | --- | --- | --- |
| ST1 - H1 | -18500 | 16000 | -24 | 21 |  | X |
| ST1 - H2 | -18000 | 18000 | -21 | 20 |  | X |
| ST1 - H3 | -19000 | 18500 | -20 | 19 |  | X |
| ST1 - V1 | -24500 | 24000 | -29 | 29 |  | X |
| ST1 - V2 | 27000 | 25000 | -32 | 30 |  | X |
| ST1 - V3 | -26500 | 25000 | -32 | 30 |  | X |
| ST2 - H1 |  |  |  |  | X | X |
| ST2 - H2 |  |  |  |  | X | X |
| ST2 - H3 |  |  |  |  | X | X |
| ST2 - V1 |  |  |  |  | X | X |
| ST2 - V2 |  |  |  |  | X | X |
| ST2 - V3 |  |  |  |  | X | X |

Table 13 - Stages range of motion – “Push locally”

Acceptance criteria:

- The vertical sensor readout must be positive when the optical table is pushed in the +Z direction

- The horizontal sensor readout on Stage 2 must be positive when the optic table is pushed in the +RZ direction

- Step 4.2

  - Absolutes value of all estimated motions must be higher than 15000counts for stage 1 (~0.018”)

  - Absolutes value of all estimated motions must be higher than 32000counts for stage 2 (~0.010”)

Test result:	Passed:   X   	Failed:         .	Waived :.     .

## Step 5 - Sensors spectra

The geophones spectra have been measured and can be found in the SVN:

seismic/BSC-ISI/X1/Unit_7/Data/Spectra/Undamped/

- X1_ISI_ITMX_ASD_m_LOC_CPS_T240_L4C_GS13_2014_05_15_125407.mat (Unlocked)

- X1_ISI_ITMX_ASD_m_LOC_CPS_T240_L4C_GS13_2014_05_15_160721.mat (Locked)

seismic/BSC-ISI/X1/Unit_7/Data/Figures/Spectra/Undamped/

- X1_ISI_ITMX_ASD_m_LOC_CPS_T240_L4C_GS13_2014_05_15_125407.fig (Unlocked)

- X1_ISI_ITMX_ASD_m_LOC_CPS_T240_L4C_GS13_2014_05_15_160721.fig (Locked)

Stage locked – unlocked

The spectra are measured in two different configurations:

- Stage 1 locked – Stage 2 locked

- Stage 1 unlocked – Stage 2 unlocked

Figure 1: Calibrated Spectra Stage 1 Locked and Stage 2 Locked

Figure 2: Calibrated Spectra Stage 1 Unlocked and Stage 2 Unlocked

Stage Tilted

The Spectra are measured when the ISI is unlocked a mass is placed on stage 2 to tilt Stage 1 and Stage 2.

The six configurations are the following in six different configurations:

- Mass placed in the actuator pocket at corner 1

- Mass placed in the pocket under the blade 0-1 at corner 1

- Mass placed in the actuator pocket at corner 2

- Mass placed in the pocket under the blade 0-1 at corner 2

- Mass placed in the actuator pocket at corner 3

- Mass placed in the pocket under the blade 0-1 at corner 3

/seismic/BSC-ISI/X1/Unit_7/Data/Spectra/Undamped/

- X1_ISI_ITMX_ASD_m_L4C_GS13_Stage_Tilted_2014_05_14.mat

seismic/BSC-ISI/X1/Unit_7/Data/Figures/Spectra/Undamped/

- X1_ISI_ITMX_Tilted_ASD_m_LOC_ST1_L4C_2014_05_14.fig

- X1_ISI_ITMX_Tilted_ASD_m_LOC_ST2_GS13_2014_05_14.fig

Figure 3 - ST1 L4C – Tilted

Figure 4 - ST1 GS13 – Tilted

Test result:	Passed:   X   	Failed:         .	Waived :.     .

## Step 6 - Coil Driver, cabling and resistance check

Resistances of the couple (actuators + in vacuum cables) were measured using the voltage and current the coil drivers read back. Resistances of the couple actuator + in-vacuum cables are reported in the table below:

| Actuator | Resistance (Ω) |
| --- | --- |
| ST1 H1 | 6.5 |
| ST1 H2 | 6.5 |
| ST1 H3 | 6.3 |
| ST1 V1 | 6.3 |
| ST1 V2 | 6.1 |
| ST1 V3 | 6.4 |
| ST2 H1 | 10.5 |
| ST2 H2 | 10.1 |
| ST2 H3 | 9.8 |
| ST2 V1 | 10.6 |
| ST2 V2 | 10.2 |
| ST2 V3 | 10.4 |

Table 14 - Actuator Resistance

Acceptance criteria:

- For the actuators of stage 1, the measured resistance between the middle pin and one side pin must be 6.3 +/-0.5 ohms

- For the actuators of stage 2, the measured resistance between the middle pin and one side pin must be 10.3 +/-0.5 ohms

- Actuator neutral pins must be connected on pin #1 (left side pin of the plug)

- Actuator drive pins must be connected on pin #2 (middle pin of the plug)

- Actuator ground shield pins must be connected on pin #3 (right pin of the plug)

- All LEDs on the coil driver front panel must be green the binary input bit must be in the upper state.

Test result:	Passed:   X   	Failed:         .	Waived :.     .

## Step 7- Actuators Sign and range of motion (Local drive)

## Step 7.1 - Actuators sign

Test result:	Passed:   X   	Failed:         .	Waived :.     .

## Step 7.2 - Range of motion - Local drive

In this step, range of motion of the two stages is checked when applying a local drive (30000 counts) on actuators.

seismic/BSC-ISI/X1/Unit_7/Data/Static_Tests/ :

X1_ISI_ITMX_Range_Of_Motion_20131210.mat

| Sensor readout (counts) | Negative drive | no drive | Positive drive | Amplitude count | mil |
| --- | --- | --- | --- | --- | --- |
| ST1 - H1 | -16779 | 67 | 16459 | 33239 | 40 |
| ST1 - H2 | -16533 | 527 | 15773 | 32307 | 38 |
| ST1 - H3 | -15789 | -127 | 16718 | 32505 | 39 |
| ST1 - V1 | -15426 | -662 | 14167 | 29594 | 35 |
| ST1 - V2 | -14853 | -66 | 14740 | 29593 | 35 |
| ST1 - V3 | -14741 | -284 | 14198 | 28939 | 34 |
| ST2 - H1 | -9087 | 1115 | 11308 | 20396 | 6.1 |
| ST2 - H2 | -8414 | 1638 | 11701 | 20116 | 6 |
| ST2 - H3 | -8453 | 1403 | 11220 | 19674 | 5.9 |
| ST2 - V1 | -12624 | -227 | 12147 | 24772 | 7.4 |
| ST2 - V2 | -12432 | -197 | 12019 | 24452 | 7.3 |
| ST2 - V3 | 11646 | 454 | 12525 | 24171 | 7.2 |

Table 15 - Range of motion - Local drive

Acceptance criteria:

- Amplitude p-p must be at least 30000 counts (0.035”) for Stage 1 CPS H

- Amplitude p-p must be at least 25000 counts (0.03”) for Stage 1 CPS V

- Amplitude p-p must be at least 18000 counts (0.005”) for Stage 2 CPS H

- Amplitude p-p must be at least 22000 counts (0.006”) for Stage 2 CPS V

- Signs of actuators drive and sensors read out have to be the same

Note: The motion of the platform can be computed. For a 30000 counts drive in the +Z direction, the platform should move by 12.6 mil on Stage 1 and 3.6mil on Stage 2.

In the Cartesian basis, the platform should move (calculation) by:

| Stage 1 - Platform move for 32K counts drive: | 12.63 | mil |
| --- | --- | --- |
| Stage 2 - Platform move for 32K counts drive: | 3.59 | mil |

Note: The range of motion in the case of a “local drive” is in agreements with the measurements done on the previous units.

Test result:	Passed:   X   	Failed:         .	Waived :.     .

## Step 8 - Vertical Sensor Calibration

Not done.

Test result:	Passed:      	Failed:         .	Waived :.   X   .

## Step 9 - Vertical Spring Constant

This test is realized by loading the ISI when one stage is locked and using the capacitive position sensors as reference.

The stiffness measurements of the spring are reported in the tables below. The nominal blade stiffness are:

- Stage 1: 1241lb/in

- Stage 2: 1465lb/in

Blade Stage 0-1

Stage 2 Locked & Stage 1 Unlocked. Stage 2 is loaded with 3 x 10kg masses and the measurements are repeated three times (by rotating the masses).

|  | Mean No load | Mean Load | Diff |  |
| --- | --- | --- | --- | --- |
| V1 | 88 | -14728 | 14816 |  |
| V2 | -3704 | -18105 | 14401 |  |
| V3 | 855 | -15103 | 15868 |  |
|  |  |  | 15028 | count |
|  |  |  | 18.1 | mil |
|  |  |  | 1219 | lb/in |
|  |  |  | -1.7 | % |

Blade Stage 1-2

Stage 1 Locked & Stage 2 Unlocked. Stage 2 is loaded with 3 x 5Kg masses and the measurements are repeated three times (by rotating the masses).

|  | Mean no load | Mean load | Diff 1 |  |
| --- | --- | --- | --- | --- |
| V1 | 493 | -25864 | 26357 |  |
| V2 | 272 | -26023 | 26296 |  |
| V3 | 679 | -25304 | 25984 |  |
|  |  |  | 26212 | count |
|  |  |  | 7.9 | mil |
|  |  |  | 1386 | lb/in |
|  |  |  | -5.4 | % |

Test mitigation: Blades are softer than design.

Test result:	Passed:   X   	Failed:       .	Waived :.     .

## Step 10 - Static Testing (Tests in the local basis)

The static tests results are reported in the SVN at :

seismic/BSC-ISI/X1/Unit_7/Data/Static_Tests/

- X1_ISI_ITMX_Offset_Local_Drive_20131210.mat

The table below shows the main and the cross-coupling when the actuators are driven in the local basis:

|  |  | Sensors |
| --- | --- | --- |
|  |  | ST1 - H1 | ST1 - H2 | ST1 - H3 | ST1 - V1 | ST1 - V2 | ST1 - V3 |
| Actuators | ST1 - H1 | 4204 | 1713 | 1681 | 16 | -9 | -39 |
|  | ST1 - H2 | 1608 | 4176 | 1689 | -7 | 9 | -6 |
|  | ST1 - H3 | 1701 | 1660 | 4123 | -19 | -11 | -2 |
|  | ST1 - V1 | -51 | -192 | 39 | 3451 | -649 | -648 |
|  | ST1 - V2 | 86 | 76 | -152 | -626 | 3450 | -634 |
|  | ST1 - V3 | -173 | 111 | 41 | -639 | -600 | -3363 |

Table 16 - Static test - Local to local - Stage 1

|  |  | Sensors |
| --- | --- | --- |
|  |  | ST2 - H1 | ST2 - H2 | ST2 - H3 | ST2 - V1 | ST2 - V2 | ST2 - V3 |
| Actuators | ST2 - H1 | 2401 | 359 | 354 | -5 | -37 | 0.9 |
|  | ST2 - H2 | 369 | 2371 | 354 | -35 | -14 | -16 |
|  | ST2 - H3 | 354 | 347 | 2315 | -47 | -26 | -5 |
|  | ST2 - V1 | 77 | 130 | -229 | -2880 | 288 | -47 |
|  | ST2 - V2 | -200 | 83 | 130 | -57 | 2844 | 293 |
|  | ST2 -V3 | 125 | -205 | 83 | 271 | -79 | 2815 |

Table 17 - Static test - Local to local - Stage 2

Acceptance criteria:

- Main couplings readout must be positive

- Comparison with the reference tables:

  - Main coupling differences mustn’t exceed 200 counts

  - Cross coupling differences mustn’t exceed 50 counts

Test result:	Passed:   X   	Failed:         .	Waived :.     .

## Step 11- Static Testing - In the general coordinate basis (Static test - CPS)

Not performed.

Test result:	Passed:       	Failed:         .	Waived :.   X .

## Step 11.2 – Base change matrices from Cartesian to Cartesian

Not performed.

Test result:	Passed:       	Failed:         .	Waived :.     X.

## Step 12 - Linearity test

The “Linearity test” was performed twice (rearranging the cables). The second time, all corners seemed to respond similarly.

The linearity test data can be found in the SVN at:

/seismic/BSC-ISI/X1/Unit_7/Data/Linearity_Test/

X1_ISI_ITMX_Linearity_test_20140515.mat

The linearity test figures can be found in the SVN at:

/seismic/BSC-ISI/X1/Unit_7/Data/Figures/Linearity_Test/

X1_ISI_ITMX_Linearity_test_20140515.fig

Figure 5 - Linearity Test

Slope – Offset:

|  |  | Slope | Offset | Average slope | Variation from average(%) |
| --- | --- | --- | --- | --- | --- |
| Stage 1 | ST1 - H1 | 0.61 | 55 | 0.60 | 1.72 |
|  | ST1 - H2 | 0.59 | 451 |  | -0.99 |
|  | ST1 - H3 | 0.59 | -88 |  | -0.73 |
|  | ST1 - V1 | 0.49 | -799 | 0.49 | 0.71 |
|  | ST1 - V2 | 0.49 | -56 |  | 0.77 |
|  | ST1 - V3 | 0.48 | -292 |  | -1.48 |
| Stage 2 | ST2 - H1 | 0.34 | 1060 | 0.34 | 1.63 |
|  | ST2 - H2 | 0.34 | 155 |  | 0.44 |
|  | ST2 - H3 | 0.33 | 1527 |  | -2.06 |
|  | ST2 - V1 | 0.42 | -642 | 0.41 | 1.41 |
|  | ST2 - V2 | 0.41 | -312 |  | -0.39 |
|  | ST2 - V3 | 0.41 | 101 |  | -1.03 |

Table 18 - Slopes and offset of the triplet Actuators - BSC-ISI – Sensors

Acceptance criteria:

- Horizontal and vertical slopes of the triplet actuators x BSC-ISI x sensors:  Average slope +/- 2.5%

Test result:	Passed:   X   	Failed:         .	Waived :.     .

## Step 13 – Transfer functions – Local to Local

Note: two vibration absorbers were installed in corner 1 and 2 vibration absorbers were installed in corner 3. No TMDs were installed on the stage 0-1 blades.

Data files measurement of local to local transfer functions in SVN at:

seismic/BSC-ISI/X1/Unit_7/Data/Transfer_Functions/Measurements/Undamped/

- X1_ISI_ITMX_TF_L2L_Raw_from_ST1_ACT_to_ST1_CPS_2014_05_14.mat

- X1_ISI_ITMX_TF_L2L_Raw_from_ST1_ACT_to_ST1_T240_2014_05_14.mat

- X1_ISI_ITMX_TF_L2L_Raw_from_ST1_ACT_to_ST2_CPS_2014_05_14.mat

- X1_ISI_ITMX_TF_L2L_Raw_from_ST1_ACT_to_ST2_GS13_2014_05_14.mat

- X1_ISI_ITMX_TF_L2L_Raw_from_ST2_ACT_to_ST1_L4C_2014_05_14.mat

- X1_ISI_ITMX_TF_L2L_Raw_from_ST2_ACT_to_ST1_T240_2014_05_14.mat

- X1_ISI_ITMX_TF_L2L_Raw_from_ST2_ACT_to_ST2_CPS_2014_05_14.mat

- X1_ISI_ITMX_TF_L2L_Raw_from_ST2_ACT_to_ST2_GS13_2014_05_14.mat

Script file for processing and plotting local to local transfer functions in SVN at:

/seisvn/seismic/BSC-ISI/X1/Unit_7/Scripts/Control_Scripts

- Step_1_TF_L2L_Raw_X1_ISI_TST.m

Figures of local to local transfer functions (Main couplings) in SVN at:

/seismic/BSC-ISI/X1/Unit_7/Data/Figures/Transfer_Functions/Measurements/Undamped/

- X1_ISI_ITMX_TF_L2L_Raw_from_ST1_ACT_to_ST1_CPS_2014_05_14.fig

- X1_ISI_ITMX_TF_L2L_Raw_from_ST1_ACT_to_ST1_L4C_2014_05_14.fig

- X1_ISI_ITMX_TF_L2L_Raw_from_ST1_ACT_to_ST1_T240_2014_05_14.fig

- X1_ISI_ITMX_TF_L2L_Raw_from_ST1_ACT_to_ST2_CPS_2014_05_14.fig

- X1_ISI_ITMX_TF_L2L_Raw_from_ST1_ACT_to_ST2_GS13_2014_05_14.fig

- X1_ISI_ITMX_TF_L2L_Raw_from_ST2_ACT_to_ST1_L4C_2014_05_14.fig

- X1_ISI_ITMX_TF_L2L_Raw_from_ST2_ACT_to_ST1_T240_2014_05_14.fig

- X1_ISI_ITMX_TF_L2L_Raw_from_ST2_ACT_to_ST2_CPS_2014_05_14.fig

- X1_ISI_ITMX_TF_L2L_Raw_from_ST2_ACT_to_ST2_GS13_2014_05_14.fig

Measured of local to local transfer functions in the SVN at:

/svncommon/seisvn/seismic/BSC-ISI/X1/Unit_7/Data/Transfer_Functions/Simulations/Undamped

- X1_ISI_ITMX_TF_L2L_Raw_2014_05_14.mat

Note 1:  The transfer functions are measured from the Output filter bank (excitation variable) to the input (IN1) of the input filter bank. The transfer functions presented below are raw transfer functions without any electronic compensation of the sensor electronic. The actuator and the coil driver electronic compensation are introduced in these transfer functions.

Figure 6: TF L2L Raw - ST1 Act to ST1 CPS

Figure 7: TF L2L Raw - ST1 Act to ST1 L4C

Figure 8 - TF L2L Raw - ST1 Act to ST1 T240

Figure 9: TF L2L Raw - ST2 Act to ST2 CPS

Figure 10: TF L2L Raw - ST2 Act to ST2 GS13

Test result:	Passed:   X   	Failed:         .	Waived :.

# BSC-ISI testing Summary

This is the third “aLigo BSC-ISI” tested at LHO. The testing procedure document E1000483-v5 was used. Tests were done during in March 2013.

The LHO ISI-BSC Unit 5 is validated per the tests presented in this report. All results are posted on the SVN at:

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/X1/Unit_7/Data

FAILED AND WAIVED TESTS

- List of tests that failed/waived and won’t be redone

- Step III.9 – Spring constant – The blades are slightly softer than the design. However, the blade softness’s are in good agreements with what was measured on the other units.

- List of tests that failed/waived, that need to be re-done during phase 2

- List of tests skipped that won't be performed because not feasible during phase II (i.e. stage 0 leveling)

- Step II.5 – Check level of Stage 0 after top-bottom plate assembly

- Step II.8 – Blade 0-1 Post Launch Angle – No need for this test, the budget mass looks good and we already reposition the Blades after noticing a gap between the Blade and its Spacer on Stage 0-1 (see comment on Step 9 – Vertical Spring Constant).

- Step III.3 – Measure the Sensor gap - This test was not performed. The sensor gaps have not been measured. These sensors have already been checked at LASTI. Moreover, risks of scratching the target are so high that we preferred not performing this test. In the future, this test will be removed from the testing procedure.

- Step III.8 – Vertical sensor calibration - The test is not realized in a proper way to evaluate accurately the calibration of the vertical CPS.

- Lists of tests skipped that needs to be done during phase II.

- Step III.14 – Symmetrization – Calibration

- Step III.17 – Lower Zero Moment Plan

- Step III.18.1 – Damping Loops – Stage 2

- Step III.18.2 – Damping Loops – Stage 1

- Step III.20 – Isolation loops