## Sheet1
COMMON MODE SERVO Table 1
COMPUTER SCREEN |  | EtherCAT -COMMON MODE SERVO TESTER
Step | NAME | FRONT PANEL LED | D No: | REMARKS |  | CM A P/F | CM B P/F
1 | GAIN IN 1 Slider | "-32" | D5 | Verify LED lighting | Step 1,2 and 23:
|  | "+16" | D4 | when the associated
|  | "+8" | D3 | input is selected | Starting with "Gain IN 1" slider click on computer screen "Gain Slider 1" and monitor the CM tester "Gain Slider 1" LED's. Each LED shall be "ON" according to slider setting. Repeat the same procedure for "Gain IN 2" and " Gain Fast" slider.
|  | "+4" | D2 | on computer screen
|  | "+2" | D1
|  | "+1" | D0
2 | GAIN IN 2 Slider | "-32" | D11
|  | "+16" | D10
|  | "+8" | D9
|  | "+4" | D8
|  | "+2" | D7
|  | "+1" | D6
3 | Input 1 Enable | "Enable" | D12 |  | Step 3 to 5, 8 to 22 and 24 to 29:
4 | Input 2 Enable | "Enable" | D13
5 | Output Switch | Enable IN 2/IN1 | D14 |  | Test each binary output ,one at a time. Select each listed tab on the computer screen monitoring tester LED's. Each LED shall be "ON" when the associated tab is selected.
6 | Boost Stage | "Boost +1" | D15
7 | Boost Stage | "Boost +2" | D16
8 | Common Comp. | "Comp." | D17
9 | Com. Exc. Enable | "Exc." | D18
10 | Com. Option | "Option" | D19
11 | Slow Polarity | "Sign + -" | D11 |  | Step 6,7: Select "Boost Stage" gain slider on computer screen and monitor the Boost Off" LED's on CM Tester front panel. LED's" shall be "ON" according to slider setting.
12 | Com. Filter | "Filter" | D21
13 | Fast Enable | "Enable" | D22
14 | Fast Polarity | "Sign + -" | D23
15 | Slow Option | "Option" | D24
16 | Slow Bypass | "Bypass" | D25
17 | Slow 5V Offset | "Fixed" | D26
18 | Slow Ofs. Enable | "Enable" | D27
19 | Slow Comp. | "Comp." | D28
20 | Slow Boost | "Boost" | D29
21 | Slow Filter | "Filter" | D30
COMMON MODE SERVO |  |  |  |  | Table 1
COMPUTER SCREEN |  | EtherCAT -COMMON MODE SERVO TESTER
Step | NAME | FRONT PANEL LED | D No: | REMARKS |  | CM A P/F | CM B P/F
22 | Fast Limiter | "Limiter" | D31
23 | GAIN FAST | "-32" | D46
|  | "+16" | D45
|  | "+8" | D44
|  | "+4" | D43
|  | "+2" | D42
|  | "+1" | D41
24 | Split Exc. Enable | "Excitation" | D47
25 | Fast Option | "Option" | D48
26 | Slow Exc. | "Fast-Slow" | D49
27 | Polarity IN 1 | "Sign + -" | D50
28 | Polarity IN 2 | "Sign + -" | D51
29 | Latch enable | "LE" | LE
Step | DISPLAYED | FRONT PANEL SWITCH | D No:
30 | Fast Limit | "Limiter" | 35 |  | Step 30,31:
31 | OK | "OK" | OK |  | Test each binary input one at a time. Turn the listed tester switches on and off, monitoring the computer screen. "Fast Limit" and "OK" LED's on computer screen shall switch on and off according to each switch setting.
Step | DISPLAYED | FRONT PANEL BNC | D No: |  | Step 32 to 34: Test each anlog input, one
at a time. Connect a DVM to the BNC
32 | - | Common Offset | 36 |  | connectors.
33 | - | Slow Offset | 37 |  | Type 9.00V then press "Enter". The DVM
34 | - | Slow Output Offset | 38 |  | shall read 9.00V DC +/- 1%.
Type -9.00V then press "Enter".
The DVM shall read -9.00V DC +/- 1%.
Type 0.00V and press "Enter".
The DVM shall read 0.00V.
COMMON MODE SERVO |  |  |  |  | Table 1
COMPUTER SCREEN |  | EtherCAT -COMMON MODE SERVO TESTER
Step | DISPLAYED | FRONT PANEL LED | D No: | REMARKS |  | CM A P/F | CM B P/F
35 | (-9.00V), 9.00V | Input Mon | 32 |  | Step 35 to 39: Test each analog output
36 | (-9.00V), 9.00V | Split Mon | 33 |  | one at a time.
37 | (-9.00V) 9.00V | Fast Mon | 34 |  | Apply -9.00V DC to each BNC connector.
38 | (-9.00V) 9.00V | Slow FB Monitor | 39 |  | The voltage applied shall be displayed
39 | (-9.00V) 9.00V | Slow Mon | 40 |  | on the computer screen.
Apply 9.00 V DC to each BNC connector.
The voltage applied shall be displayed
on the computer screen.
Remove the input voltage. On computer
screen shall be displayed 0.00V.
The voltage tolerance shall be +/- 1/%.
VOLTAGE CONTROLLED OSCILLATOR |  |  |  |  | Table 1
COMPUTER SCREEN |  | EtherCAT LOW NOISE VCO TESTER
Step | Displayed | FRONT PANEL BNC | D No: | REMARKS
40 | 9.00V, (-9.00V) | RF Pwr Mon. M1 | - |  | Step 40 to 46: Apply 9.00V DC to each
41 | 9.00V, (-9.00V) | RF Pwr Mon. M2 | - |  | BNC connector. 9.00 V DC shall be
42 | 9.00V, (-9.00V) | RF Pwr Mon. M3 | - |  | displayed on the computer screen.
43 | 9.00V, (-9.00V) | Temp. Mon. M1 | - |  | Apply -9.00V DC. The applied
44 | 9.00V, (-9.00V) | Temp. Mon. M2 | - |  | voltage shall be displayed on the
45 | 9.00V, (-9.00V) | Temp. Mon. M3 | - |  | computer screen.
46 | 9.00V, (-9.00V) | VCO Tune Mon. | - |  | The tolerance shall be +/- 1%.
47 | - | Man.Freq.Tuning | - |  | Step 47:
Connect a DVM to the "Man. Freq. Tuning" BNC connector. Type 9.00V and press "ENTER". DVM shall read 9.00V +/- 1%. Type - 9.00V and press "ENTER". DVM shall read -9.00V +/- 1%. Type 0.00V and press "Enter" DVM shall read 0.00 V.
Step | COMPUTER SCREEN | FRONT PANEL SWITCH | D No: |  | REMARKS
48 | EXCSW | " Exc Readback" | - |  | Step 48,49: Turn the listed switches on
|  |  | - |  | and off monitoring computer screen.
49 | OK | "OK" | - |  | "EXCSW" and "OK" on computer screen
shall switch between on and off
according to each switch setting.
Step | COMPUTER SCREEN | FRONT PANEL LED | D No: |  | REMARKS
50 | Exc. Enable | "Exc Enable" | - |  | Select "Exc Enable" on the computer
screen monitoring "Exc Enable" LED on
the tester front panel. LED shall be ON
when "Exc Enable" is selected.
DELAY LINE PHASE SHIFTER |  |  |  |  | Table 1
COMPUTER SCREEN |  | EtherCAT -DELAY LINE PHASE SHIFTER TESTER
Step | NAME | FRONT PANEL LED | D No: | REMARKS
51 | Delay A | "1/16" | - |  | Click on the "Delay Line A" gain slider on computer screen and monitor tester front panel LED's. Each LED shall be "ON" according to gain slider setting.
|  | "1/8" | -
|  | "1/4" | -
|  | "1/2" | -
|  | "1" | -
|  | "2" | -
|  | "4" | - |  | Select "LE A" tab on computer screen.
|  | "8" | - |  | "LE" LED shall be ON when tab is
|  | "16" | - |  | selected.
52 | LE A | "LE" | -
53 | Delay B | "1/16" | - |  | Click on the "Delay Line B" gain slider
|  | "1/8" | - |  | on computer screen and monitor tester
|  | "1/4" | - |  | front panel LED's. Each LED shall be "ON"
|  | "1/2" | - |  | according to gain slider setting.
|  | "1" | -
|  | "2" | - |  | Select "LE B" tab on computer screen.
|  | "4" | - |  | "LE" LED shall be ON when tab is
|  | "8" | - |  | selected.
|  | "16" | -
54 | LE B | "LE" | -