LIGO Laboratory / LIGO Scientific Collaboration

LIGO-T1100323-v1	LIGO	15-June-2011

aLIGO Modification of SpectraCom IRIG-B Breakout Adapter

E.Merilh

Distribution of this document:

LIGO Scientific Collaboration

This is an internal working note

of the LIGO Laboratory.

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 159 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

http://www.ligo.caltech.edu/

# Introduction

The purpose of this procedure is to modify the SpectraCom IRIG-B breakout cable by removing the DB9 connector that carries the DCLS signals and wire in a BNC cable end that will carry the IRIG DCLS+ signal.

# Tools Needed

- Razor knife

- Phillips screw driver

- Flush/Diagonal Cutters

- Shrink wrap; ¼”PVC

- RG-58 BNC Connector kit

- 24AWG Wire stripper

- .213dia Hexagonal crimp tool

- Heat gun

- IDEXPERT labeler w/xsl-32-427 black on white cartridge

# Procedure

1. Remove the DB9 shell at the end of the Digital I/O cable and carefully remove the heat shrink label with a razor knife and remove strain relief clamp and copper shielding.

2 .Cut all wires back to jacket. Strip jacket back 1” and trim back all wires except for the red wire and the ground wire.

3. Solder female RG-58 center conductor pin onto end of IRIG DCLS Input + (red) wire and slide 1” of 1/4”PVC shrink wrap and RG58 crimp sleeve over jacket end.

4. Press center conductor pin into connector housing.

5. Wrap ground wire around throat of connector. Solder may be used to ensure strong electrical connection.

6. Wrap copper shielding tape around connection on throat of connector then slide crimp sleeve over and crimp using .213 dia hexagonal crimp tool.

7. Slide ¼” PVC shrink wrap over crimp sleeve and apply heat.

8. Create and apply self-laminating label to jacket using the IDEXPERT labeler w/xsl-32-427 black on white cartridge, 9 pt. font and center justified text.