| AUTHOR: | DATE | CHECKED: |
| --- | --- | --- |
| Lisa C. Austin | 8-2-11 | Mike Smith |

This specification is for uncoated, laser-quality, plano-plano windows. They are to be used as viewports on UHV chambers.

Applicable Documents

LIGO- D1101548	Uncoated high quality, non-wedged, 6" Viewport, Optic

Requirements

Physical Configuration

| Diameter | 6.00( +/- 0.01( |
| --- | --- |
| Thickness | 0.75( +/- 0.01( |
| Wedge | < 5 arc minutes |
| Bevel | 45 deg. bevel on both edges, width approx. 0.03( or as appropriate to prevent chipping |
| Clear Aperture | > 5.20( |
| Markings | Part and serial number, as defined below, shall be marked on the barrel |

Substrate Material

Fused silica, Corning 7980 or equivalent, grade 1B or better

Part and Serial Number

The Part/Serial number shall be per D1101548 and of the format: 
D1101548   Where

YY 		is incremental for each optic type, starting at 01

Registration Mark

Registration mark shall be etched, ground or sandblasted

Surface Quality: Surfaces 1 and 2, within the clear aperture

Laser grade polish: 10-5 scratch-dig surface quality

Transmitted wavefront error, within the clear aperture

(/10 or less, at 633 nm

Testing and Documentation

| Specification | Test Method | Frequency of Inspection | Data Delivered |
| --- | --- | --- | --- |
| Physical Dimensions | Visual Inspection | 100% | Certification |
| Surface Quality | Visual Inspection | 100% | Certification |
| Transmitted Wavefront Error | Interferometry | 100% | Certification |