LIGO Laboratory / LIGO Scientific Collaboration

LIGO- E1100846 	                 LIGO                   Mar 18th, 2014

aLIGO BSC-ISI, Pre-integration Testing report,

Phase II (before and after cartridge install)

E1100846 – V3

Jim Warner

Distribution of this document:

Advanced LIGO Project

This is an internal working note

of the LIGO Laboratory

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 1970 Mail Stop S9-02 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

Table of contents:

1.   Phase II-a	4

1.   Hardware changes	4

1.   CPS – E1100369	4

2.   GS13 – E1100740	4

3.   L4C – E1100740	4

4.   T240 – E1100740	4

5.   Cables – E1100822	4

6.   Misc	4

2.   Electronic Inventory	4

3.   Models Modifications	4

4.   Mass distribution	5

5.   Basic functionalities just after installing the BSC-ISI on the teststand	6

1.    Pressure sensors	6

i.   Spectra	6

ii.   Actuators-cables resistance	7

iii.   Offsets CPS Unlocked vs locked 	8

iv.   Offset local drive	8

v.   Offset Cartesian drive	8

vi.   Range of motion	9

vii.   Linearity test	9

6.   Transfer functions and Comparison with measurements done in the staging building.	9

1.   At the end station	9

7.   Conclusion Phase II-a	13

3.   Conclusion Phase II-b	22

Introduction

The BSC-ISI testing is performed in three phases:

1) BSC-ISI, Pre-integration Testing, Phase I (post-assembly, in the staging building)

2) BSC-ISI, Pre-integration Testing, Phase II: Final tests done before insertion in the chamber
3) BSC-ISI, Integration Phase Testing: Procedure and results related to the commissioning in the chamber.

The ISI-BSC10 was moved from the Staging building to the VEA test stand July 2013.

This document presents results of tests (Phase II) performed on the ISI-BSC10 (ETMY) before installation in the chamber.

All results are posted on the SVN at:

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/H1/ETMY/

The following type of document can be found in the SVN:

- Excell spreadsheet (.xls)

- Data location

- Figures location

- Masses distribution scheme (ppt)

# Phase II-a

## Hardware changes

### CPS – E1100369

CPS have not been replaced since phase I testing in the staging building.

### GS13 – E1100740

GS13 have not been replaced since phase I testing in the staging building.

### L4C – E1100740

L4C have not been replaced since phase I testing in the staging building.

### T240 – E1100740

T240 have not been replaced since phase I testing in the staging building.

### Cables – E1100822

The cables installed on ISI-BSC3 are reported in the table below. Further information can be found in E1100822.

### Misc

No hardware changes since phase I testing in the staging building.

## Electronic Inventory

This table reports the electronic equipment used in the LVEA.

## Models Modifications

The model was updated and recompiled since MEDM screens were modified

## Mass distribution

Here is the payload distribution of the ISI. SUS-QUAD and TMS are installed.

| Stage 1 |  |  |  |  |
| --- | --- | --- | --- | --- |
| Location | Mass Type | Weight (lb) | Weight (Kg) |
| Corner 1 | D0902616 - 2,3,4,4 | 13 | 6.0 |
| Corner 2 | D0902616 - 2,2,2,3,3,1 | 10 | 4.6 |
| Corner 3 | D0902616 - None | 0 | 0.0 |
|  | 6 vibration absorbers (60Kg) | 90 | 41.2 |
|  |  | 113 | 51.8 |

Table 1 - Payload Stage 1

| Stage 2 |  |  |  |  |
| --- | --- | --- | --- | --- |
| Location | Mass Type | Weight (lb) | Weight (Kg) |
| Corner 1 | D071200 - 1,1,1,5,5 | 33 | 15 |
| Corner 2 | D071200 - 5,5,6 | 58 | 27 |
| Corner 3 | D071200 - 1,1,6 | 29 | 13 |
|  | Keel Masses (D1003136) | 600 | 275 |
|  | Ballast Masses (6x46lb) | 276 | 126 |
|  | Total | 997 | 457 |

Table 2 - Payload Stage 2

This is the Stage 2 Seismic payload. Add ~900 lbs for a quad and ~450 for the TMS and the total payload is 2347lb.

Test result:                                              Passed:   X             Failed:         .   W Waived:  __

## Basic functionalities just after installing the BSC-ISI on the teststand

### Pressure sensors

All pressure sensors are working.

Table 3 - Geophones Pressure sensors

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

### Spectra

Spectra of the instrument can be found in the SVN at:

seismic/BSC-ISI/H1/ETMY/Data/Spectra/Undamped/

- H1_ISI_ETMY_ASD_m_LOC_CPS_T240_L4C_GS13_2020_02_20_4 7:2:.mat

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/H1/ETMY/Data/Figures/Spectra/Undamped/

-    H1_ISI_ETMY_ASD_m_LOC_CPS_T240_L4C_GS13_2020_02_20_4 7:2:.fig

Figure 1 - Spectra inboard instruments - ISI Unlocked

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

### Actuators-cables resistance

Not done. Can be completed outside the chamber if needed.

Test result:                                              Passed:                 Failed:         .   W Waived:   X__

### Offsets CPS Unlocked vs locked

|  | Table locked | Table unlocked | Difference locked - unlocked |
| --- | --- | --- | --- |
| Sensors | Offset (Mean) | Std deviation | Offset (Mean) | Std deviation | Offset (Mean) | mil |
| ST1 - H1 | -49 | 9 | -193 | 110 | -144 | .17 |
| ST1 - H2 | 2290 | 5 | 1824 | 138 | -466 | .56 |
| ST1 - H3 | 270 | 5 | -2776 | 123 | -3047 | 3.63 |
| ST1 - V1 | 372 | 31 | 1552 | 80 | 1180 | 1.4 |
| ST1 - V2 | 1570 | 11 | 576 | 95 | -994 | 1.18 |
| ST1 - V3 | 342 | 17 | 975 | 115 | -1317 | 1.57 |
| ST2 - H1 | 695 | 31 | -307 | 199 | -1003 | .30 |
| ST2 - H2 | 1854 | 33 | 1273 | 224 | -580 | .17 |
| ST2 - H3 | 2253 | 20 | 2172 | 117 | -81 | .02 |
| ST2 - V1 | 161 | 20 | 200 | 207 | 39 | .01 |
| ST2 - V2 | -290 | 22 | -123 | 227 | 167 | .05 |
| ST2 - V3 | 3923 | 21 | 3758 | 275 | -165 | .05 |

Table 4 - Locked vs Unlocked Position

Test result:                                              Passed:  X               Failed:         .   W Waived:   __

### Offset local drive

Skipped due to time.

Test result:                                              Passed:   X              Failed:         .   W Waived:   X__

### Offset Cartesian drive

Skipped due to time.

### Range of motion

The range of motion of the table is measured by pushing on the table in a direction collinear to the CPS. The Static tests results can be found on the SVN at:

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/H1/ETMY/Data/Static_Tests/

| Sensor readout (counts) | Negative drive | no drive | Positive drive | Amplitude count | mil |
| --- | --- | --- | --- | --- | --- |
| ST1 - H1 | -15612 | -193 | 16542 | 32169 | 38 |
| ST1 - H2 | -15628 | 1824 | 16948 | 32150 | 38 |
| ST1 - H3 | -14882 | -2776 | 14759 | 28784 | 34 |
| ST1 - V1 | -11694 | 1552 | 13890 | 26027 | 31 |
| ST1 - V2 | -12381 | 576 | 13739 | 26276 | 31 |
| ST1 - V3 | -14945 | -974 | 11519 | 25983 | 31 |
| ST2 - H1 | -10347 | -307 | 9150 | 19280 | 5.7 |
| ST2 - H2 | -8289 | 1273 | 11022 | 19381 | 5.8 |
| ST2 - H3 | -7700 | 2172 | 11844 | 19315 | 5.7 |
| ST2 - V1 | -12616 | 199 | 10781 | 22504 | 6.7 |
| ST2 - V2 | -11426 | -122 | 10823 | 23188 | 6.9 |
| ST2 - V3 | -9812 | 3758 | 13120 | 22479 | 6.7 |

Table 5 - Range of motion - Actuator drive in the LVEA

Test result:                                              Passed:   X              Failed:         .   W Waived:

### Linearity test

Skipped due to time.

Note:

Test result:                                              Passed:    X              Failed:         .   W Waived:

## Transfer functions and Comparison with measurements done in the staging building.

### At the end station

Measurements data can be found in the SVN at:

SeiSVN/seismic/BSC-ISI/H1/ETMY/Data/Transfer_Functions/Measurements/Undamped:

- H1_ISI_ETMY_Data_L2L_500Hz_1000Hz_ST1_ST2_20140219-055819.mat

H1_ISI_ETMY_Data_L2L_100Hz_500Hz_ST1_ST2_20140219-032326.mat

H1_ISI_ETMY_Data_L2L_100mHz_700mHz_ST1_ST2_20140219-003031.mat

H1_ISI_ETMY_Data_L2L_10Hz_100Hz_ST1_ST2_20140219-183606.mat

H1_ISI_ETMY_Data_L2L_700mHz_10Hz_ST1_ST2_20140219-153718.mat

Once the data are processed, they can be found in the SVN at:

/seismic/BSC-ISI/H1/ETMY/Data/Transfer_Functions/Simulations/Undamped/
- H1_ISI_TST_TF_L2L_Raw_2014_02_21.mat

The transfer functions can be found in the SVN at:

seismic/BSC-ISI/H1/ETMY/Data/Figures/Transfer_Functions/Measurements/Undamped/

- H1_ISI_TST_TF_L2L_Raw_from_ST1_ACT_to_ST1_CPS_2014_02_21.fig

- H1_ISI_TST_TF_L2L_Raw_from_ST1_ACT_to_ST1_L4C_2014_02_21.fig

- H1_ISI_TST_TF_L2L_Raw_from_ST1_ACT_to_ST1_T240_2014_02_21.fig

- H1_ISI_TST_TF_L2L_Raw_from_ST1_ACT_to_ST2_CPS_2014_02_21.fig

- H1_ISI_TST_TF_L2L_Raw_from_ST1_ACT_to_ST2_GS13_2014_02_21.fig

- H1_ISI_TST_TF_L2L_Raw_from_ST2_ACT_to_ST1_L4C_2014_02_21.fig

- H1_ISI_TST_TF_L2L_Raw_from_ST2_ACT_to_ST1_T240_2014_02_21.fig

- H1_ISI_TST_TF_L2L_Raw_from_ST2_ACT_to_ST2_CPS_2014_02_21.fig

- H1_ISI_TST_TF_L2L_Raw_from_ST2_ACT_to_ST2_GS13_2014_02_21.fig

Figure 2 - TF ST1 ACT to ST1 CPS

Figure 3 - TF ST1 ACT to ST1 T240

Figure 4 - TF ST1 ACT to ST1 L4C

Figure 5 - TF ST2 ACT to ST2 CPS

Figure 6 - TF ST2 ACT to ST2 GS13

The tests show some noise and extra peaks between 1 & 100 hz. These were probably due to an improper set up of the TMS, and were resolved in-chamber.

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

## Conclusion Phase II-a

Tests performed during Phase II-a don’t show any major anomalies on ISI-BSC10.

Test result:                                              Passed:   X              Failed:         .   W Waived:   _

Phase IIb Tests after Cartridge Install

Basic Functionality after Cartridge Installation

    - Pressure Sensors

All pressure sensors are working.

Test result:                                    Passed:   X              Failed:         .   W Waived:

    - Spectra

Spectra of the seismometers can be found in the svn at:

seismic/BSC-ISI/H1/ETMY/Data/Spectra/Undamped/

H1_ISI_ETMY_ASD_m_LOC_CPS_T240_L4C_GS13_2020_03_17_4 7:7:.mat

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/H1/ETMY/Data/Figures/Spectra/Undamped/
H1_ISI_ETMY_ASD_m_LOC_CPS_T240_L4C_GS13_2020_03_17_4 7:7:.fig

Figure 7 - St1 CPS Power Spectra

Figure 8 - St1 T240 Power Spectra

Figure 9 - St1 L4C Power Spectra

Figure 10 - St2 CPS Power Spectra

Figure 11 - St2 GS13 Power Spectra

All the seismometers are functioning normally.

Test result:                                    Passed:   X              Failed:         .   W Waived:

    - Lock-Unlock Shifts

|  | Table locked | Table unlocked | Difference locked - unlocked |
| --- | --- | --- | --- |
| Sensors | Offset (Mean) | Std deviation | Offset (Mean) | Std deviation | Offset (Mean) | mil |
| ST1 - H1 | 278 | 3.6 | 425 | 21 | -146 | .17 |
| ST1 - H2 | 776 | 11 | -127 | 21 | 905 | 1.08 |
| ST1 - H3 | 388 | 6 | -231 | 21 | 619 | .74 |
| ST1 - V1 | -970 | 5 | -1243 | 21 | 273 | .33 |
| ST1 - V2 | 592 | 5 | 188 | 25 | 405 | .48 |
| ST1 - V3 | -830 | 4 | -603 | 36 | -226 | .27 |
| ST2 - H1 | -524 | 86 | -520 | 36 | -5 | 0! |
| ST2 - H2 | -194 | 33 | 188 | 37 | -383 | .11 |
| ST2 - H3 | -265 | 64 | 231 | 36 | -497 | .15 |
| ST2 - V1 | 90 | 148 | 64 | 63 | 27 | .01 |
| ST2 - V2 | 619 | 37 | 1181 | 62 | -562 | .17 |
| ST2 - V3 | 1403 | 98 | 1622 | 74 | -219 | .07 |

Table 6 - Lock/Unlock Shifts.

Locker shifts are within spec.

Test result:                                    Passed:   X              Failed:         .   W Waived:

    - Range of motion

| Sensor readout (counts) | Negative drive | no drive | Positive drive | Amplitude count | mil |
| --- | --- | --- | --- | --- | --- |
| ST1 - H1 | -16065 | -210 | 16038 | 23169 | 38 |
| ST1 - H2 | -15086 | -106 | 17221 | 32150 | 38 |
| ST1 - H3 | -14274 | -191 | 17491 | 28783 | 34 |
| ST1 - V1 | -13524 | 545 | 12711 | 26027 | 30 |
| ST1 - V2 | -13758 | 684 | 12284 | 26275 | 31 |
| ST1 - V3 | -14712 | 394 | 11741 | 25982 | 31 |
| ST2 - H1 | -9997 | 407 | 9762 | 19280 | 5.7 |
| ST2 - H2 | -9655 | 827 | 9762 | 19381 | 5.7 |
| ST2 - H3 | -10027 | 982 | 9552 | 19315 | 5.7 |
| ST2 - V1 | -10934 | -37 | 12746 | 22504 | 6.7 |
| ST2 - V2 | -11336 | -323 | 11032 | 23187 | 6.9 |
| ST2 - V3 | -11006 | 1289 | 11947 | 22478 | 6.7 |

Table 7 - Range of motion under actuator drive

Range of motion is within spec.

Test result:                                    Passed:   X              Failed:         .   W Waived:

    - CPS Linearity Test

Test result:                                    Passed:                 Failed:         .   W Waived:   X

  - Transfer Functions after Cartridge Install

Reference transfer functions L2L with ETMY, TMSY, ACB and all ISI viton damping elements.

Transfer Functions are located in the SVN at:

/SeiSVN/seismic/BSC-ISI/H1/ETMY/Data/Transfer_Functions//Undamped/

-H1_ISI_ETMY_Data_L2L_700mHz_10Hz_ST1_ST2_20140313-191146.mat

-H1_ISI_ETMY_Data_L2L_10Hz_100Hz_ST1_ST2_20140313-221042.mat

-H1_ISI_ETMY_Data_L2L_100mHz_700mHz_ST1_ST2_20140314-011735.mat

-H1_ISI_ETMY_Data_L2L_100Hz_500Hz_ST1_ST2_20140314-041028.mat

-H1_ISI_ETMY_Data_L2L_500Hz_1000Hz_ST1_ST2_20140314-061519.mat

L2L concatenated

/SeiSVN/seismic/BSC-ISI/H1/ETMY/Data/Transfer_Functions/Simulations/Undamped/

-H1_ISI_ETMY_TF_L2L_Raw_2014_03_14.mat

Figure 12 St1 Act to St1 CPS

Figure 13 St1 Act to T240

Figure 14 St1 Act to L4C

Figure 15 St2 Act to GS13

Figure 16 St2 Act to St2 CPS

Transfer functions look healthy.

Test result:                                    Passed:   X              Failed:         .   W Waived:

## Conclusion Phase II-b

Tests performed during Phase II-b don’t show any anomalies on ISI-BSC10.

Test result:                                              Passed:   X              Failed:         .   W Waived:   _