LIGO Laboratory / LIGO Scientific Collaboration

LIGO- E1100855-V3 	                 LIGO                   October 18th, 2012

L1 BSC 2 BSC-ISI, Pre-integration Testing report,

Phase II

E1100855-V3

Adrien Le Roux, Michael Puma, Céline Ramet

Distribution of this document:

Advanced LIGO Project

This is an internal working note

of the LIGO Laboratory

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 1970 Mail Stop S9-02 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

Table of contents:

Introduction	4

1.	Phase II-a	5

1.	Hardware changes	5

1.	CPS – E1100369	5

2.	GS13 – E1100740	5

3.	L4C – E1100740	5

4.	T240 – E1100740	5

5.	Cables – E1100822	5

6.	Misc	5

2.	Electronic Inventory	6

3.	Models Modifications	6

4.	Mass distribution	7

1.	Seismic	7

2.	Suspension	8

3.	Miscellaneous	8

4.	Total	8

5.	Basic functionalities just after installing the BSC-ISI on the test stand	8

1.	Pressure sensors	8

2.	Spectra	9

3.	Actuators-cables resistance	10

4.	Offsets CPS Unlocked vs locked	10

5.	Offset local drive	11

6.	Offset Cartesian drive	11

7.	Range of motion	12

8.	Linearity test	12

6.	Transfer functions and Comparison with measurements done in the staging building.	14

1.	In the LVEA	14

2.	Comparisons with measurements in the staging building	17

7.	Conclusion Phase II-a	21

2.	Phase II-b- After cartridge install	22

1.	Mass distribution	22

1.	Seismic	22

2.	Suspension	23

3.	Miscellaneous	23

4.	Total	23

2.	Basic functionalities just after installing the BSC-ISI on the test stand	23

1.	Pressure sensors	23

2.	Spectra	24

3.	Actuators-cables resistance	24

4.	Offsets CPS Unlocked vs locked	25

5.	Offset local drive	25

6.	Offset Cartesian drive	26

7.	Range of motion	27

8.	Linearity test	27

3.	Transfer functions and Comparison with measurements done in the staging building.	29

1.	In the LVEA	29

2.	Comparisons with measurements before cartridge install	33

4.	Conclusion Phase II-b	37

Conclusion	38

Introduction

The BSC-ISI testing is performed in three phases:

1) BSC-ISI, Pre-integration Testing, Phase I (post-assembly, in the staging building)

2) BSC-ISI, Pre-integration Testing, Phase II: Final tests done before insertion in the chamber
3) BSC-ISI, Integration Phase Testing: Procedure and results related to the commissioning in the chamber.

The ISI-BSC2 was moved from the Staging building to the LVEA test stand on May 17th, 2012.

This document presents results from the series of tests (Phase II) performed on the ISI-BSC2 (BS) in the corner station. The tests were done with the Beam Splitter suspension (BS) installed.

First tests started on August 6th, 2012. The first testing phase (II-a validation before cartridge installation) was completed on August 17th 2012.

All results are posted on the SVN at:

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/BS/

The following types of documents can be found in the SVN:

- Data location

- Figures location

# Phase II-a

## Hardware changes

### CPS – E1100369

13579 CPS was replaced by 13523, because 13579 was constantly saturating regardless how close the sensor and the target was.

### GS13 – E1100740

GS13 have not been replaced since phase I testing in the staging building.

### L4C – E1100740

L4C have not been replaced since phase I testing in the staging building.

### T240 – E1100740

The 3 T240s were replaced because their pressure sensors were not working at the time of Phase 1 testing. We know have:

- S/N 001 in corner 1

- S/N 028 in corner 2

- S/N 030 in corner 3

### Cables – E1100822

Only 4 cables were changed in order to equilibrate the use of the different cables between BSC-ISIs.

- GS-13 Corner 1 D1100154 S11074260 was replaced by D1100155 S1107397 
- Actuator 0-1 V Corner 2 D1100150 S1107076 was replaced by D1100151 S1107199 
- Actuator 0-1 H Corner 2 D1100150 S1107074 was replaced by D1100151 S1107190 
- Trillium Corner 1 D1100152 S1107223 was replaced by D1100153 S1107275

### Misc

No hardware changes since phase I testing in the staging building.

## Electronic Inventory

This table reports the electronic equipment used in the LVEA.

| Hardware | Ligo reference | S/N |
| --- | --- | --- |
| Interface Chassis Pod 1 | D1002432 | S1106357 |
| Interface Chassis Pod 2 | D1002432 | S1200666 |
| Interface Chassis Pod 3 | D1002432 | S1200665 |
| Anti-aliasing Chassis Pod 1 | D1002693 | S1200947 |
| Anti-aliasing Chassis Pod 2 | D1002693 | S1200946 |
| Anti-aliasing Chassis Pod 3 | D1002693 | S1200948 |
| Anti-image Chassis | D070081 | S1200995 |
| Binary Input Chassis 1 | D1001726 | S1101291 |
| Binary Input Chassis 2 | D1001726 | S1101292 |
| Binary Output Chassis | D1001728 | S1101325 |
| T240 Interface Pod 1 | D1002694 | S1104427 |
| T240 Interface Pod 2 | D1002694 | S1104425 |
| T240 Interface Pod 3 | D1002694 | S1104421 |
| Anti-image Chassis | D1000305 | N/A |
| Coil driver Pod 1 | D0902744 | S1103305 |
| Coil driver Pod 2 | D0902744 | S1103337 |
| Coil driver Pod 3 | D0902744 | S1103360 |
| Expansion chassis | L1seib62 | S1101148 |

Table 2 - Electronic inventory

## Models Modifications

No model modifications were done between the beginning and the end of Phase 2a testing.

## Mass distribution

This final mass distribution will be presented once all elements will be installed on the ISI (during phase II-b). These elements are the vibration absorbers on stage 1 and the QUAD structure.

### Seismic

Stage 1

| 8/16/2012 | D0902612 | D0902613 | D0902614 | D0902615 | D0902616 | D0901075 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | 01 | 02 | 03 | 04 | 05 | 1 | 2 |  |  |
|  | 12 | 15 | 10 | 10 | 3.5 | 1 | 2 | 5 | 0.5 | 11 | 22 | lbs | kgs |
| C1-1 | 1 |  |  |  |  |  |  |  |  |  |  | 12 | 5.44 |
| C1-2 |  | 0 |  |  |  |  |  |  |  |  |  | 0 | 0.00 |
| C1-3 |  |  |  |  |  |  | 1 | 1 |  |  |  | 7 | 3.18 |
| C2-1 | 1 |  |  |  |  |  |  |  |  |  |  | 12 | 5.44 |
| C2-2 |  | 1 |  |  |  |  |  |  |  |  |  | 15 | 6.80 |
| C2-3 |  |  |  |  |  | 1 | 1 |  |  |  |  | 3 | 1.36 |
| C3-1 | 1 |  |  |  |  |  |  |  |  |  |  | 12 | 5.44 |
| C3-2 |  | 1 |  |  |  |  |  |  |  |  |  | 15 | 6.80 |
| C3-3 |  |  |  |  |  | 1 |  | 1 |  |  |  | 6 | 2.72 |
| Stage 1 | 3 | 2 | 0 | 0 | 0 | 2 | 2 | 2 | 0 | 0 | 0 | 82 | 37.19 |

Stage 2

| 8/16/2012 | D1003136 | D1003161 | D071200 |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  |  | 0 | 1 | 2 | 3 | 4 | 5 | 6 |  |  |
|  | 50 | 19.6 | 20.6 | 21.6 | 22.6 | 23.6 | 0.6 | 1.1 | 2.2 | 4.5 | 7.9 | 15.6 | 27.2 | lbs | kgs |
| Keel | 12 |  |  |  |  |  |  |  |  |  |  |  |  | 600 | 272.40 |
| Optical |  | 1 | 1 | 4 | 1 | 1 |  |  |  |  |  |  |  | 172.8 | 78.45 |
| E-1 |  |  |  |  |  |  |  |  | 1 | 1 |  | 3 |  | 53.5 | 24.29 |
| E-2 |  |  |  |  |  |  | 4 | 2 |  | 3 |  |  |  | 18.1 | 8.22 |
| E-3 |  |  |  |  |  |  | 4 |  |  |  |  | 1 | 1 | 45.2 | 20.52 |
| F1 |  |  |  |  |  |  | 4 | 5 |  |  |  |  | 1 | 35.1 | 15.94 |
| F2 |  |  |  |  |  |  |  |  |  |  |  |  | 2 | 54.4 | 24.70 |
| F3 |  |  |  |  |  |  | 1 |  |  |  |  |  | 2 | 55 | 24.97 |
| Stage 2 | 12 | 1 | 12 | 1 | 1 | 1 | 13 | 7 | 1 | 4 | 0 | 4 | 6 | 1034.1 | 469.48 |

| 8/16/2012 | D972213 | D972214 | D0901075 | D071200 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | 5 kg | 10 kg | 01 | 02 | 03 | 04 | 05 | 06 |  |  |
|  | 610 | 376 | 11 | 22 | 1.1 | 2.2 | 4.5 | 7.9 | 15.6 | 27.2 | lbs | kgs |
| A | 1 |  |  |  |  |  |  |  | 1 | 1 | 652.8 | 296.10 |
| B |  | 1 |  |  |  |  |  |  |  | 4 | 338.8 | 153.68 |
| C | 1 |  |  |  |  |  |  |  |  |  | 610 | 276.69 |
| D |  |  |  |  |  |  |  |  |  |  | 0 | 0.00 |
| E-1 |  |  |  |  |  |  |  |  |  |  | 0 | 0.00 |
| E-2 |  |  |  |  |  |  |  |  |  |  | 0 | 0.00 |
| E-3 |  |  |  |  |  |  |  |  |  |  | 0 | 0.00 |
| F1 |  |  |  |  |  | 1 |  |  | 1 |  | 17.8 | 8.07 |
| F2 |  |  |  |  |  | 1 |  |  |  | 4 | 111 | 50.35 |
| F3 |  |  |  |  | 1 | 1 |  |  |  | 2 | 57.7 | 26.17 |
| Stage 2 | 2 | 1 | 0 | 0 | 1 | 3 | 0 | 0 | 2 | 11 | 1788.1 | 811.07 |

### Suspension

The BS weight was estimated to be 204.70 kgs. (E1000337)

### Miscellaneous

The added cables were estimated to weigh 5.86 kgs. (E1000337)

### Total

Nominal mass hanging on stage 0-1 blades (without stage 2): 912Kg – 2010lb

Nominal mass hanging on stage 1-2 blades: 2830Kg – 6239lb

Nominal payload on stage 1: 109Kg – 240lb

Nominal payload on stage 2: 1185Kg – 2612lb

|  | Plan | 3/9/2012 | LVEA |
| --- | --- | --- | --- |
| Stage 1 (kg) | 108.86 | 36.29 | 37.19 |
| Stage 2 (kg) | 1183.42 | 1096.83 | 811.07 |
| Suspension (kg) |  |  | 204.7 |
| Miscellaneous (kg) |  |  | 5.86 |
| Total (kg) | 1292.28 | 1133.12 | 1058.82 |

Test result:                                              Passed:                 Failed:         .   W Waived:   X

## Basic functionalities just after installing the BSC-ISI on the test stand

### Pressure sensors

All pressure sensors are working.

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/BS/Data/Static_Tests/
LLO_ISI_BS_Pressure_Sensors_Check_Calibrated_2012_08_06_163846.mat

|  | Pressure (KPa) |
| --- | --- |
| Sensors | Corner 1 | Corner 2 | Corner 3 |
| ST1-L4C-P | 66.93 | 99.11 | 98.97 |
| ST1-L4C-D | 1.68 | 1.21 | 1.48 |
| ST1-GS13-P | 99.92 | 100.14 | 99.92 |
| ST1-GS13-D | -0.93 | -1.12 | -0.91 |
| ST1-T240-P | 153.48 | 153.00 | 153.57 |

Table 3 - Geophones Pressure sensors

Note/comment about this test: The direct pressure of the corner 1 L4-C was identified to be a problem with the BSC-ISI interface and not with the L4-C itself.

Test result:                                              Passed:    X             Failed:         .   W Waived:   _

### Spectra

Spectra of the instrument can be found in the SVN at:

seismic/BSC-ISI/L1/BS/Data/Spectra/Undamped/

LLO_ISI_BS_ASD_m_LOC_CPS_T240_L4C_GS13_2012_08_06_200624.mat

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/BS/Data/Figures/Spectra/Undamped/
LLO_ISI_BS_ASD_m_LOC_CPS_T240_L4C_GS13_2012_08_06_200624.fig

Figure 1 - Spectra inboard instruments - ISI Unlocked

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

### Actuators-cables resistance

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/BS/Data/Static_Tests/

- L1_ISI_BSC2_Actuators_Resistance_20120817.mat

|  | Stage 1 | Stage 2 |
| --- | --- | --- |
|  | H1 | H2 | H3 | V1 | V2 | V3 | H1 | H2 | H3 | V1 | V2 | V3 |
| Script | 7.39 | 7.38 | 7.49 | 7.30 | 7.41 | 7.23 | 11.31 | 11.75 | 11.59 | 11.32 | 11.78 | 11.19 |
| Hand | 7.3 | 7.4 | 7.5 | 7.3 | 7.4 | 7.5 | 11.5 | 11.5 | 11.5 | 11.3 | 11.6 | 11.3 |

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

### Offsets CPS Unlocked vs locked

The table is not perfectly balanced but it is considered sufficiently good to perform the series of test before the cartridge installation. Working conditions at the end station are difficult. A fine balancing will be done during phase II-b.

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/BS/Data/Static_Tests/

- LLO_ISI_BS_CPS_Read_Back_ISI_Locked_2012_08_06_155728.mat

- LLO_ISI_BS_CPS_Read_Back_ISI_Unlocked_2012_08_06_163846.mat

|  | Table locked | Table unlocked | Difference locked - unlocked |
| --- | --- | --- | --- |
| Sensors | Offset (Mean) | Std deviation | Offset (Mean) | Std deviation | Offset (Mean) | mil |
| ST1 - H1 | -13.85 | 3.94 | 60.30 | 33.00 | -74.14 | -0.09 |
| ST1 - H2 | 159.97 | 3.65 | -1631.67 | 24.41 | 1791.64 | 2.13 |
| ST1 - H3 | 31.39 | 4.03 | -1860.89 | 33.13 | 1892.28 | 2.25 |
| ST1 - V1 | -149.63 | 6.23 | 492.39 | 42.35 | -642.02 | -0.76 |
| ST1 - V2 | -43.47 | 3.88 | 4010.50 | 47.66 | -4053.97 | -4.83 |
| ST1 - V3 | -128.72 | 4.93 | 1616.65 | 58.94 | -1745.37 | -2.08 |
| ST2 - H1 | 56.46 | 9.44 | -833.36 | 31.50 | 889.82 | 0.26 |
| ST2 - H2 | -214.09 | 11.24 | -2306.60 | 52.65 | 2092.51 | 0.62 |
| ST2 - H3 | 627.68 | 18.58 | 1860.72 | 43.70 | -1233.05 | -0.37 |
| ST2 - V1 | 9.90 | 14.88 | 7394.22 | 120.64 | -7384.31 | -2.20 |
| ST2 - V2 | 78.49 | 13.75 | 4126.33 | 129.11 | -4047.84 | -1.20 |
| ST2 - V3 | 280.27 | 21.68 | 3503.18 | 146.70 | -3222.91 | -0.96 |

Table 4 - Locked vs Unlocked Position

Additionally note that during most of the testing, the balance was actually not as good as presented here. Indeed the testing was completed after alignment of the beam splitter and the balancing of this one was adjusted. They also added additional hardware onto the table for alignment purposes.

Test result:                                              Passed:   _              Failed:         .   W Waived:   X__

### Offset local drive

Note: Due to longer cable s, offsets measured by CPS for a 7000 count drive are slightly lower than offsets measured in the staging building.

Results of this test can be found in the SVN at:

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/BS/Data/Static_Tests/

- LLO_ISI_BS_Offset_Local_Drive_20120814.mat

|  | Sensors |
| --- | --- |
|  |  | ST1 - H1 | ST1 - H2 | ST1 - H3 | ST1 - V1 | ST1 - V2 | ST1 - V3 |
| Actuators | ST1 - H1 | 3743.11 | 1507.15 | 1507.14 | 8.52 | -9.00 | -7.16 |
|  | ST1 - H2 | 1495.88 | 3715.49 | 1494.25 | -4.39 | 2.70 | 0.85 |
|  | ST1 - H3 | 1492.10 | 1485.77 | 3688.95 | -16.73 | -3.91 | 3.93 |
|  | ST1 - V1 | 34.40 | -135.01 | 74.02 | 2853.91 | -524.92 | -496.35 |
|  | ST1 - V2 | 85.76 | 47.43 | -120.90 | -505.97 | 2946.74 | -514.12 |
|  | ST1 - V3 | -144.21 | 77.87 | 17.48 | -531.26 | -510.58 | 2916.33 |

Table 5 - Static Tests – Local to Local - Stage 1

|  |  | Sensors |
| --- | --- | --- |
|  |  | ST2 - H1 | ST2 - H2 | ST2 - H3 | ST2 - V1 | ST2 - V2 | ST2 - V3 |
| Actuators | ST2 - H1 | 2153.24 | 336.54 | 307.33 | -18.60 | -31.00 | 23.59 |
|  | ST2 - H2 | 347.74 | 2165.22 | 320.63 | 16.31 | -39.44 | 13.02 |
|  | ST2 - H3 | 332.41 | 330.77 | 2155.23 | 18.07 | -83.47 | 11.85 |
|  | ST2 - V1 | 66.54 | 97.52 | -183.37 | 2563.58 | 308.05 | -22.06 |
|  | ST2 - V2 | -194.94 | 66.84 | 114.64 | -9.30 | 2639.97 | 291.64 |
|  | ST2 - V3 | 114.60 | -182.12 | 65.12 | 339.85 | -4.82 | 2617.36 |

Table 6 - Static Tests – Local to Local - Stage 2

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

### Offset Cartesian drive

The test was not performed because it only tests the matrices, which can be changed at any time.

Test result:                                              Passed:   _              Failed:         .   W Waived:   X_

### Range of motion

The range of motion of the table is measured by pushing on the table in a direction collinear to the CPS. The Static tests results can be found on the SVN at:

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/BS/Data/Static_Tests/
LLO_ISI_BS_Range_Of_Motion_20120814.mat

| Sensor readout (counts) | Negative drive | no drive | Positive drive | Amplitude count | mil |
| --- | --- | --- | --- | --- | --- |
| ST1 - H1 | -15723.20 | 237.00 | 15708.62 | 31431.81 | 37.42 |
| ST1 - H2 | -15526.72 | -1592.00 | 14300.75 | 29827.47 | 35.51 |
| ST1 - H3 | -16008.36 | -2017.00 | 13769.04 | 29777.40 | 35.45 |
| ST1 - V1 | -11517.92 | 676.00 | 12886.89 | 24404.81 | 29.05 |
| ST1 - V2 | -7663.12 | 4922.00 | 17545.11 | 25208.23 | 30.01 |
| ST1 - V3 | -12405.74 | 115.00 | 12649.79 | 25055.53 | 29.83 |
| ST2 - H1 | -10453.53 | -1296.00 | 7816.67 | 18270.20 | 5.44 |
| ST2 - H2 | -10705.45 | -1452.00 | 7747.15 | 18452.60 | 5.49 |
| ST2 - H3 | -9038.56 | 228.00 | 9461.42 | 18499.98 | 5.51 |
| ST2 - V1 | -2806.55 | 8095.00 | 18990.87 | 21797.41 | 6.49 |
| ST2 - V2 | -5277.62 | 6051.00 | 17332.00 | 22609.62 | 6.73 |
| ST2 - V3 | -10869.26 | 335.00 | 11546.60 | 22415.86 | 6.67 |

Table 7 - Range of motion - Actuator drive in the LVEA

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

### Linearity test

The data of the linearity test can be found on the SVN at:

seismic/BSC-ISI/L1/BS/Data/Linearity_Test/

- LLO_ISI_BSC2_Linearity_test_20120814.mat

The figures of the linearity test can be found on the SVN at:

seismic\BSC-ISI\H2\BS\Data\Figures\Linearity_Test\

- LLO_ISI_BSC2_Linearity_test_20120814.fig

Figure 2 - Linearity test –L1 - BS – In LVEA

|  |  | Slope | Offset | Average slope | Variation from average(%) |
| --- | --- | --- | --- | --- | --- |
| Stage 1 | ST1 - H1 | 0.533697 | 226.1222 | 0.529864 | 0.723451 |
|  | ST1 - H2 | 0.530224 | -1597.27 |  | 0.067976 |
|  | ST1 - H3 | 0.52567 | -2017.53 |  | -0.79143 |
|  | ST1 - V1 | 0.407374 | 689.2569 | 0.414936 | -1.82241 |
|  | ST1 - V2 | 0.419246 | 4942.302 |  | 1.038839 |
|  | ST1 - V3 | 0.418187 | 152.1449 |  | 0.783571 |
| Stage 2 | ST2 - H1 | 0.305848 | -1274.94 | 0.307575 | -0.56131 |
|  | ST2 - H2 | 0.308059 | -1456.79 |  | 0.157416 |
|  | ST2 - H3 | 0.308817 | 269.3765 |  | 0.403897 |
|  | ST2 - V1 | 0.362087 | 8128.792 | 0.371402 | -2.50818 |
|  | ST2 - V2 | 0.377928 | 6102.843 |  | 1.757234 |
|  | ST2 - V3 | 0.374191 | 330.0906 |  | 0.750941 |

Table 8 - Slope – Offset Linearity test

Note: A few slopes are over 2% but nothing appears really bad.

Test result:                                              Passed:   _              Failed:         .   W Waived:   X_

## Transfer functions and Comparison with measurements done in the staging building.

### In the LVEA

The parameters for the measurements in the LVEA are slightly different from those in the staging building. We chose to have weaker excitation but longer averages in an effort to reduce risk of the attached suspension.

At this point, only the tuned mass dampers on the spring are installed (No vibration absorbers, no Viton under the keel masses, No vibration absorbers on the BS)

Measurements data can be found in the SVN at:

SeiSVN/seismic/BSC-ISI/L1/BS/Data/Transfer_Functions/Measurements/Undamped:

- LLO_ISI_BSC2_Data_L2L_10mHz_100mHz_ST1_ST2_20120816-171919.mat

- LLO_ISI_BSC2_Data_L2L_100mHz_700mHz_ST1_ST2_20120815-025159.mat

- LLO_ISI_BSC2_Data_L2L_700mHz_10Hz_ST1_ST2_20120816-194216.mat

- LLO_ISI_BSC2_Data_L2L_10Hz_100Hz_ST1_ST2_20120814-214508.mat

- LLO_ISI_BSC2_Data_L2L_100Hz_500Hz_ST1_ST2_20120814-191017.mat

- LLO_ISI_BSC2_Data_L2L_500Hz_1000Hz_ST1_ST2_20120814-170550.mat

Data after processing can be found in the SVN at:

SeiSVN/seismic/BSC-ISI/L1/BS/Data/Transfer_Functions/Simulations/Undamped

- LLO_ISI_BSC2_TF_L2L_Raw_10mHz_1000Hz_2012_08_17.mat

The transfer functions can be found in the SVN at:

seismic/BSC-ISI/L1/BS/Data/Figures/Transfer_Functions/Measurements/Undamped/

- LLO_ISI_BSC2_TF_L2L_Raw_from_ST1_ACT_to_ST1_CPS_2012_08_17.fig

- LLO_ISI_BSC2_TF_L2L_Raw_from_ST1_ACT_to_ST1_L4C_2012_08_17.fig

- LLO_ISI_BSC2_TF_L2L_Raw_from_ST1_ACT_to_ST1_T240_2012_08_17.fig

- LLO_ISI_BSC2_TF_L2L_Raw_from_ST2_ACT_to_ST2_CPS_2012_08_17.fig

- LLO_ISI_BSC2_TF_L2L_Raw_from_ST2_ACT_to_ST2_GS13_2012_08_17.fig

Note 1:  The transfer functions are measured from the Output filters bank excitation point to the input (IN1) of the input filters bank. The transfer functions presented below are raw transfer functions without any electronic compensation.

Note 2: The L4Cs are out of phase (should be -90 before 1Hz).  A minus sign is added in the calibration filters that convert count to nm/s.

Note 3: On the ST1-ACT-H to ST1-CPS-H transfer functions, we can see the first resonances of the LVEA test stand at 20 Hz, 21.5Hz and 28.8Hz (matches within 1 Hz with LHO’s results).

Note 4: Even without vibration absorbers on the BS structure, the first resonance visible on Stage 2 is at 154 Hz.

Because we haven’t done the comparison plot, we leave the H2 BSC 6 transfer functions for comparison.

Figure 3 - TF ST1 ACT to ST1 CPS

Figure 4 - TF ST1 ACT to ST1 L4C

Figure 5 - TF ST1 ACT to ST1 T240

Figure 6 - TF ST2 ACT to ST2 CPS

Figure 7 - TF ST2 ACT to ST2 GS13

### Comparisons with measurements in the staging building

The script used to compare transfer function can be found in the SVN at:

SeiSVN/seismic/BSC-ISI/Common/Plot_Functions_BSC_ISI/Version_0

- Plot_Comparison_TF_L2L_BSC_ISI.m

The figure that shows the comparison between the transfer functions of the staging building and the LVEA are located in the SVN at:

SeiSVN/seismic/BSC-ISI/L1/BS/Data/Figures/Transfer_Functions/Measurements/Comparison/L2L/

- COMP_LVEA_L1_ISI_BS_ST1_CPS_2012_10_11_.fig

- COMP_LVEA_L1_ISI_BS_ST2_CPS_2012_10_11_.fig

- COMP_LVEA_L1_ISI_BS_ST1_L4C_2012_10_11_.fig

- COMP_LVEA_L1_ISI_BS_ST2_GS13_2012_10_11_.fig

Main differences are:
- The DC gains (cables resistance)
- Resonances of the rigid body modes [1; 10]Hz (different payload)
- ST1-CPS resonances different in the staging building and the LVEA (Test stand short legs vs long legs). Resonance at 53Hz on ST1 CPS V1.
- Resonance at 78.5Hz on stage 2 sensors (Payload)
- Similar at high frequencies

Figure 9 - Transfer functions comparison - ST1 ACT to ST1 CPS

Figure 11 - Transfer functions comparison - ST1 L4C V to ST1 L4C V

Figure 12 - Transfer functions comparison - ST2 ACT to ST2 CPS

Figure 12 – Zoom in on transfer functions comparison - ST2 ACT H to ST2 CPS H

Figure 13 - Zoom in on transfer functions comparison - ST2 ACT V to ST2 CPS V

Figure 15 - Transfer functions comparison – ST2 ACT V to ST2 GS13 V

Figure 15 – Zoom in transfer functions comparison – ST2 ACT H to ST2 GS13 H

Figure 15 – Zoom in transfer functions comparison – ST2 ACT V to ST2 GS13 V

## Conclusion Phase II-a

All results appear satisfying, apart from a few test waived (will be redone during phase II-b):

- Offset unlocked vs locked (Stage 1 is not perfectly balanced but it is sufficient for phase II-a testing)

- Static tests in the Cartesian basis (redundant with the static test in the local basis)

- Pressure sensor issue on L4C 1 (interface issue)

- Linearity test: some corners have slightly different slopes but nothing strikingly alarming. We have seen in the past (HAM-ISI phase 2 testing) stronger slopes different, blamed on different cable lengths inside the LVEA.

- Balancing on the test stand revealed a 75 kgs difference between the expected and the actual payload. This may be due to mistakes in the mass budget reporting, or poor weight estimation of the payload !?!

Test result:                                              Passed:   X              Failed:         .   W Waived:   _

# Phase II-b- After cartridge install

The cartridge install was done on September 7th 201. The mass distribution was finalized, while vibration absorbers were installed. Also cable brackets were installed and the cabling was improved. Phase II-b testing was completed on October 10th and the dome was close on October 11th.

## Mass distribution

This is the final mass distribution after cartridge install. We followed D1101852-v7 for the optical table and ballast masses and then balanced the ISI using the side masses. Please note that between Phase II a and Phase II b the vibration absorbers were installed on Stage 1 of the ISI.

### Seismic

Stage 1

| 10/10/2012 | D0902612 | D0902613 | D0902616 | D1001760 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | 1 | 2 | 3 | 4 | 5 |  |  |  |
|  | 12 | 15 | 3.5 | 1 | 2 | 5 | 0.5 | 15.86 | lbs | kgs |
| C1-1 |  |  |  |  |  |  |  |  | 0 | 0 |
| C1-2 |  |  |  |  |  |  |  |  | 0 | 0 |
| C1-3 |  |  |  | 2 |  | 1 |  | 1 | 22.86 | 10.37844 |
| C2-1 |  |  |  |  |  |  |  |  | 0 | 0 |
| C2-2 |  |  |  |  |  |  |  |  | 0 | 0 |
| C2-3 |  |  |  |  |  |  |  | 2 | 31.72 | 14.40088 |
| C3-1 |  |  |  |  |  |  |  |  | 0 | 0 |
| C3-2 |  | 1 |  |  |  |  |  |  | 15 | 6.81 |
| C3-3 |  |  |  | 1 |  |  |  | 1 | 16.86 | 7.65444 |
| Stage 1 | 0 | 1 | 0 | 3 | 0 | 1 | 0 | 4 | 82 | 39.24376 |

Stage 2

| 10/9/2012 | D1003136 | D1003161 | D071200 |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  |  | 0 | 1 | 2 | 3 | 4 | 5 | 6 |  |  |
|  | 50 | 43.2 | 45.4 | 47.6 | 49.8 | 52 | 0.6 | 1.1 | 2.2 | 4.5 | 7.9 | 15.6 | 27.2 | lbs | kgs |
| Keel | 12 |  |  |  |  |  |  |  |  |  |  |  |  | 600 | 272.40 |
| Optical |  | 1 | 1 | 19 | 1 | 1 |  |  |  |  |  |  |  | 1095.3 | 497.25 |
| E-1 |  |  |  |  |  |  |  |  | 1 | 1 |  | 3 |  | 53.5 | 24.29 |
| E-2 |  |  |  |  |  |  | 4 | 2 |  | 3 |  |  |  | 18.1 | 8.22 |
| E-3 |  |  |  |  |  |  | 4 |  |  |  |  | 1 | 1 | 45.2 | 20.52 |
| F1 |  |  |  |  |  |  | 4 | 5 |  |  |  |  | 1 | 35.1 | 15.94 |
| F2 |  |  |  |  |  |  |  |  |  |  |  |  | 2 | 54.4 | 24.70 |
| F3 |  |  |  |  |  |  | 1 |  |  |  |  |  | 2 | 55 | 24.97 |
| Stage 2 | 12 | 1 | 12 | 1 | 1 | 1 | 13 | 7 | 1 | 4 | 0 | 4 | 6 | 1956.6 | 888.28 |

### Suspension

The BS weight was estimated to be 204.70 kgs. (E1000337)

### Miscellaneous

The added cables were estimated to weigh 5.86 kgs. (E1000337)

### Total

|  | Plan | 3/9/2012 | LVEA (08/16/12) | Chamber (10/12/12) |
| --- | --- | --- | --- | --- |
| Stage 1 (kg) | 108.86 | 36.29 | 37.19 | 39.24 |
| Stage 2 (kg) | 1183.42 | 1096.83 | 811.1 | 888.28 |
| Suspension (kg) |  |  | 204.7 | 204.70 |
| Miscellaneous (kg) |  |  | 5.86 | 5.86 |
| Total (kg) | 1292.28 | 1133.12 | 1058.8 | 1138.08 |

This confirms that the BS weight was well estimated and that the difference in weight found in side chamber testing was due to a problem estimating the ISI masses.

Test result:                                              Passed:   X              Failed:         .   W Waived:

## Basic functionalities just after installing the BSC-ISI on the test stand

### Pressure sensors

All pressure sensors are working.

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/BS/Data/Static_Tests/
L1_ISI_BS_Pressure_Sensors_Check_Calibrated_2012_10_14_163321.mat

|  | Pressure (KPa) |
| --- | --- |
| Sensors | Corner 1 | Corner 2 | Corner 3 |
| ST1-L4C-P | 66.9776 | 98.96156 | 98.95647 |
| ST1-L4C-D | 1.223052 | 1.955529 | 1.801736 |
| ST1-GS13-P | 100.1601 | 100.3851 | 100.1741 |
| ST1-GS13-D | -0.91096 | -1.08985 | -0.925 |
| ST1-T240-P | 153.6562 | 153.2118 | 153.8986 |

Table 3 - Geophones Pressure sensors

Note/comment about this test: The direct pressure of the corner 1 L4-C was identified to be a problem with the BSC-ISI interface and not with the L4-C itself.

Test result:                                              Passed:    X             Failed:         .   W Waived:   _

### Spectra

Spectra of the instrument can be found in the SVN at:

seismic/BSC-ISI/L1/BS/Data/Spectra/Undamped/

LLO_ISI_BS_ASD_m_LOC_CPS_T240_L4C_GS13_2012_10_01_202526.mat

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/BS/Data/Figures/Spectra/Undamped/

- LLO_ISI_BS_ASD_m_LOC_CPS_T240_L4C_GS13_2012_10_01_202526.fig

Figure 1 - Spectra inboard instruments - ISI Unlocked

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

### Actuators-cables resistance

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/BS/Data/Static_Tests/

- L1_ISI_BS_Actuators_Resistance_20121015.mat

|  | Stage 1 | Stage 2 |
| --- | --- | --- |
|  | H1 | H2 | H3 | V1 | V2 | V3 | H1 | H2 | H3 | V1 | V2 | V3 |
| Script | 7.02 | 6.95 | 6.90 | 6.90 | 7.02 | 6.94 | 11.01 | 11.40 | 10.78 | 10.96 | 11.41 | 11.17 |

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

### Offsets CPS Unlocked vs locked

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/BS/Data/Static_Tests/

- LLO_ISI_BS_CPS_Read_Back_ISI_Locked_2012_10_08_160124.mat

- LLO_ISI_BS_CPS_Read_Back_ISI_Unlocked_2012_10_08_163641.mat

|  | Table locked | Table unlocked | Difference locked - unlocked |
| --- | --- | --- | --- |
| Sensors | Offset (Mean) | Std deviation | Offset (Mean) | Std deviation | Offset (Mean) | mil |
| ST1 - H1 | -23.009 | 14.13914 | -1559.94 | 90.73972 | -1536.93 | -1.82968 |
| ST1 - H2 | -189.17 | 8.526574 | -539.702 | 27.94949 | -350.532 | -0.4173 |
| ST1 - H3 | -217.21 | 9.926204 | -798.299 | 66.35433 | -581.089 | -0.69177 |
| ST1 - V1 | -229.904 | 17.321 | -1077.55 | 56.48563 | -847.65 | -1.00911 |
| ST1 - V2 | 20.91925 | 9.871158 | -894.045 | 57.53425 | -914.964 | -1.08924 |
| ST1 - V3 | -306.683 | 10.96414 | -434.712 | 66.00704 | -128.029 | -0.15242 |
| ST2 - H1 | -167.819 | 136.5499 | 499.088 | 111.0348 | 666.9067 | 0.198484 |
| ST2 - H2 | -697.277 | 59.06676 | 1802.584 | 130.2406 | 2499.861 | 0.744006 |
| ST2 - H3 | 50.68177 | 85.01798 | 2062.966 | 44.03003 | 2012.284 | 0.598894 |
| ST2 - V1 | -1002.9 | 155.9407 | 68.46054 | 172.6333 | 1071.357 | 0.318856 |
| ST2 - V2 | 1091.066 | 237.3824 | 254.6616 | 172.3709 | -836.404 | -0.24893 |
| ST2 - V3 | -667.076 | 55.35058 | -1037.48 | 179.5833 | -370.404 | -0.11024 |

Table 4 - Locked vs Unlocked Position

Note: Stage 1 level is not as good as we usually like to get it. Our limitation here, is that in order to improve the vertical level, we would have needed to remove one of the vibration absorbers, which we did not want to do (as we already had installed only 4 of those on Stage 1).

Test result:                                              Passed:   X_              Failed:         .   W Waived:   __

### Offset local drive

Note: Due to longer cables, offsets measured by CPS for a 7000 count drive are slightly lower than offsets measured in the staging building.

Results of this test can be found in the SVN at:

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/BS/Data/Static_Tests/

- LLO_ISI_BS_Offset_Local_Drive_20121008.mat

|  | Sensors |
| --- | --- |
|  |  | ST1 - H1 | ST1 - H2 | ST1 - H3 | ST1 - V1 | ST1 - V2 | ST1 - V3 |
| Actuators | ST1 - H1 | 3945.138 | 1589.714 | 1587.759 | 8.40346 | 11.05842 | -4.06386 |
|  | ST1 - H2 | 1607.562 | 3968.233 | 1592.45 | 4.45278 | -0.79596 | -1.47768 |
|  | ST1 - H3 | 1594.809 | 1584.72 | 3920.97 | -12.0188 | 18.85602 | -1.92364 |
|  | ST1 - V1 | 24.4812 | -141.38 | 74.46578 | 2940.304 | -493.381 | -486.627 |
|  | ST1 - V2 | 99.1584 | 46.283 | -120.667 | -480.811 | 3031.996 | -508.308 |
|  | ST1 - V3 | -132.198 | 73.04544 | 16.65176 | -507.342 | -471.681 | 2977.197 |

Table 5 - Static Tests – Local to Local - Stage 1

|  |  | Sensors |
| --- | --- | --- |
|  |  | ST2 - H1 | ST2 - H2 | ST2 - H3 | ST2 - V1 | ST2 - V2 | ST2 - V3 |
| Actuators | ST2 - H1 | 2213.337 | 354.623 | 322.746 | -30.3301 | -17.0111 | -5.06274 |
|  | ST2 - H2 | 345.4151 | 2229.945 | 332.914 | -13.6421 | -56.5943 | 7.61952 |
|  | ST2 - H3 | 329.8847 | 339.9476 | 2211.603 | -34.1825 | -52.8874 | 5.0095 |
|  | ST2 - V1 | 75.46992 | 108.659 | -167.07 | 2569.149 | 313.6153 | 12.25218 |
|  | ST2 - V2 | -200.895 | 67.7532 | 104.3314 | -7.05884 | 2643.61 | 301.7954 |
|  | ST2 - V3 | 102.7038 | -144.555 | 65.5426 | 291.1055 | -26.7397 | 2572.482 |

Table 6 - Static Tests – Local to Local - Stage 2

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

### Offset Cartesian drive

The test was not performed because it only tests the matrices, which can be changed at any time.

Test result:                                              Passed:   _              Failed:         .   W Waived:   X_

### Range of motion

The range of motion of the table is measured by pushing on the table in a direction collinear to the CPS. The Static tests results can be found on the SVN at:

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/BS/Data/Static_Tests/
- L1_ISI_BS_Range_Of_Motion_20121015.mat

| Sensor readout (counts) | Negative drive | Positive drive | Amplitude count | mil |
| --- | --- | --- | --- | --- |
| ST1 - H1 | -16264.6 | 15353.09 | 31617.68 | 37.6401 |
| ST1 - H2 | -15814.7 | 16021.22 | 31835.91 | 37.89989 |
| ST1 - H3 | -16690.3 | 15373.55 | 32063.83 | 38.17123 |
| ST1 - V1 | -15071.2 | 10145.66 | 25216.83 | 30.02003 |
| ST1 - V2 | -14914.3 | 11028.69 | 25943 | 30.88452 |
| ST1 - V3 | -13000 | 12586.91 | 25586.88 | 30.46057 |
| ST2 - H1 | -8648.96 | 10188.87 | 18837.84 | 5.606499 |
| ST2 - H2 | -8217.55 | 10750.42 | 18967.97 | 5.645228 |
| ST2 - H3 | -7054.84 | 11710.8 | 18765.64 | 5.585011 |
| ST2 - V1 | -15541.9 | 6716.805 | 22258.66 | 6.624602 |
| ST2 - V2 | -13600.9 | 9285.09 | 22886.03 | 6.811319 |
| ST2 - V3 | -12871.7 | 9495.85 | 22367.52 | 6.656999 |

Table 7 - Range of motion - Actuator drive in the LVEA

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

### Linearity test

The data of the linearity test can be found on the SVN at:

seismic/BSC-ISI/L1/BS/Data/Linearity_Test/

- LLO_ISI_BSC2_Linearity_test_20121009.mat

The figures of the linearity test can be found on the SVN at:

seismic\BSC-ISI\H2\BS\Data\Figures\Linearity_Test\

- LLO_ISI_BSC2_Linearity_test_20121009.fig

Figure 2 - Linearity test –L1 - BS – In Chamber

|  |  | Slope | Offset | Average slope | Variation from average(%) |
| --- | --- | --- | --- | --- | --- |
| Stage 1 | ST1 - H1 | 0.562152 | -1601.38 | 0.562146 | 0.001176 |
|  | ST1 - H2 | 0.565076 | -584.167 |  | 0.521231 |
|  | ST1 - H3 | 0.559209 | -801.554 |  | -0.52241 |
|  | ST1 - V1 | 0.420188 | -1049.65 | 0.425896 | -1.34037 |
|  | ST1 - V2 | 0.432098 | -873.607 |  | 1.456235 |
|  | ST1 - V3 | 0.425403 | -343.901 |  | -0.11587 |
| Stage 2 | ST2 - H1 | 0.31708 | 469.5074 | 0.317112 | -0.01015 |
|  | ST2 - H2 | 0.318163 | 1728.48 |  | 0.331479 |
|  | ST2 - H3 | 0.316093 | 2019.604 |  | -0.32132 |
|  | ST2 - V1 | 0.371271 | 5.412806 | 0.375579 | -1.14717 |
|  | ST2 - V2 | 0.383337 | 275.1606 |  | 2.065527 |
|  | ST2 - V3 | 0.37213 | -951.599 |  | -0.91836 |

Table 8 - Slope – Offset Linearity test

Note: One slope is slightly over 2% but nothing dramatic.

Test result:                                              Passed:   _              Failed:         .   W Waived:   X_

## Transfer functions and comparison with measurements done before the cartridge install

### In the LVEA

We kept the excitations identical to what was used on the LVEA test stand before the cartridge install.

At this point, only the tuned mass dampers on the spring are installed (No vibration absorbers, no Viton under the keel masses, no vibration absorbers on the BS)

Measurements data can be found in the SVN at:

SeiSVN/seismic/BSC-ISI/L1/BS/Data/Transfer_Functions/Measurements/Undamped:

- LLO_ISI_BSC2_Data_L2L_10mHz_100mHz_ST1_ST2_20121010-224952.mat

- LLO_ISI_BSC2_Data_L2L_100mHz_700mHz_ST1_ST2_20121010-172700.mat

- LLO_ISI_BSC2_Data_L2L_700mHz_10Hz_ST1_ST2_20121011-180140.mat

- LLO_ISI_BSC2_Data_L2L_10Hz_100Hz_ST1_ST2_20121011-022506.mat

- LLO_ISI_BSC2_Data_L2L_100Hz_500Hz_ST1_ST2_20121011-011248.mat

- LLO_ISI_BSC2_Data_L2L_500Hz_1000Hz_ST1_ST2_20121011-015239.mat

Data after processing can be found in the SVN at:

SeiSVN/seismic/BSC-ISI/L1/BS/Data/Transfer_Functions/Simulations/Undamped

- LLO_ISI_BSC2_TF_L2L_Raw_10mHz_1000Hz_2012_10_11.mat

The transfer functions can be found in the SVN at:

seismic/BSC-ISI/L1/BS/Data/Figures/Transfer_Functions/Measurements/Undamped/

- LLO_ISI_BSC2_TF_L2L_Raw_from_ST1_ACT_to_ST1_CPS_2012_10_11.fig

- LLO_ISI_BSC2_TF_L2L_Raw_from_ST1_ACT_to_ST1_L4C_2012_10_11.fig

- LLO_ISI_BSC2_TF_L2L_Raw_from_ST1_ACT_to_ST1_T240_2012_08_17.fig

- LLO_ISI_BSC2_TF_L2L_Raw_from_ST2_ACT_to_ST2_CPS_2012_10_11.fig

- LLO_ISI_BSC2_TF_L2L_Raw_from_ST2_ACT_to_ST2_GS13_2012_10_11.fig

Note 1:  The transfer functions are measured from the Output filters bank excitation point to the input (IN1) of the input filters bank. The transfer functions presented below are raw transfer functions without any electronic compensation.

Note 2: The L4Cs are out of phase (should be -90 before 1Hz).  A minus sign is added in the calibration filters that convert count to nm/s.

Note 3: On the ST1-ACT-H to ST1-CPS-H transfer functions, we can see the first HEPI resonances at 11 Hz.

Note 4: With the installation of the vibration absorbers, the first resonance on Stage 1 is at 219 Hz (moved from 154 Hz).

Because we haven’t done the comparison plot, we leave the H2 BSC 6 transfer functions for comparison.

Figure 3 - TF ST1 ACT to ST1 CPS

Figure 4 - TF ST1 ACT to ST1 L4C

Figure 5 - TF ST1 ACT to ST1 T240

Figure 6 - TF ST2 ACT to ST2 CPS

Figure 7 - TF ST2 ACT to ST2 GS13

### Comparisons with measurements before cartridge install

The script used to compare transfer function can be found in the SVN at:

SeiSVN/seismic/BSC-ISI/Common/Plot_Functions_BSC_ISI/Version_0/

- Plot_Comparison_TF_L2L_BSC_ISI.m

The figure that shows the comparison between the transfer functions of the staging building and the LVEA are located in the SVN at:

SeiSVN/seismic/BSC-ISI/L1/BS/Data/Figures/Transfer_Functions/Measurements/Comparison/L2L/

- COMP_CARTRIDGE_L1_ISI_BS_ST1_CPS_2012_10_11_.fig

- COMP_CARTRIDGE_L1_ISI_BS_ST2_CPS_2012_10_11_.fig

- COMP_CARTRIDGE_L1_ISI_BS_ST1_L4C_2012_10_11_.fig

- COMP_CARTRIDGE_L1_ISI_BS_ST2_GS13_2012_10_11_.fig

Main differences are:

- Resonances of the rigid body modes [1; 10]Hz (different payload- with masses on optical table and on the keel)

- ST1-CPS resonances different between test stand and HEPI piers (visible between 10 and 70 Hz)

- Some noise around 3.2 Hz, especially visible on Stage 2 GS-13

- Transfer functions look similar at high frequencies with some improvement at 220 Hz on Stage 1 L4-Cs due to the installation of Stage 1 vibration absorbers.

Figure 9 - Transfer functions comparison - ST1 ACT to ST1 CPS

Figure 11 - Transfer functions comparison - ST1 L4C to ST1 L4C

Figure 12 – Zoom in on transfer functions comparison – ST1 ACT to ST1 L4C –effect of ST 1 vibration absorbers

Figure 12 - Transfer functions comparison - ST2 ACT to ST2 CPS

Figure 12 – Zoom in on transfer functions comparison - ST2 ACT to ST2 CPS

Figure 15 - Transfer functions comparison – ST2 ACT to ST2 GS13

Figure 15 – Zoom in transfer functions comparison – ST2 ACT to ST2 GS13

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

### 3 -Comparisons with measurements after closing the dome

The script used to compare transfer function can be found in the SVN at:

SeiSVN/seismic/BSC-ISI/Common/Plot_Functions_BSC_ISI/Version_0/

- Plot_Comparison_TF_L2L_BSC_ISI.m

The figure that shows the comparison between the transfer functions of the staging building and the LVEA are located in the SVN at:

SeiSVN/seismic/BSC-ISI/L1/BS/Data/Figures/Transfer_Functions/Measurements/Comparison/L2L/

- COMP_DOME_L1_ISI_BS_ST1_CPS_2012_10_11_.fig

- COMP_DOME_L1_ISI_BS_ST2_CPS_2012_10_11_.fig

- COMP_DOME_L1_ISI_BS_ST1_L4C_2012_10_11_.fig

- COMP_DOME_L1_ISI_BS_ST2_GS13_2012_10_11_.fig

Comments:
- There is logically great agreement between  those measurements as the only difference was the closing of the dome on October 11th. The main difference is that the noise previously visible around 3.2 Hz disappeared on those measurements.

Figure 9 - Transfer functions comparison - ST1 ACT to ST1 CPS

Figure 11 - Transfer functions comparison - ST1 L4C to ST1 L4C

Figure 12 - Transfer functions comparison - ST2 ACT to ST2 CPS

Figure 15 - Transfer functions comparison – ST2 ACT to ST2 GS13

Figure 15 – Zoom in transfer functions comparison – ST2 ACT to ST2 GS13

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

## Conclusion Phase II-b

All results appear satisfying, apart from a few tests waived:

- Static tests in the Cartesian basis (redundant with the static test in the local basis)

- Pressure sensor issue on L4C 1 (shown to be an interface issue)

- Linearity test: some corners have slightly different slopes but nothing strikingly alarming. We have seen in the past (HAM-ISI phase 2 testing) stronger slopes different, blamed on different cable lengths inside the LVEA.

The only worry we had after install was about the noise around 3.2 Hz and this noise disappeared right after the dome was closed.

Test result:                                              Passed:   X              Failed:         .   W Waived:   _

Conclusion

The ISI-BSC2 was moved from the Staging building to the LVEA test stand in May 2012

This document presents series of tests (Phase II) performed on the ISI-BSC2 (BS).

First tests started in early August 2012, after BS installation and re-balancing of the ISI. Testing took 2 weeks (with breaks for Initial Alignment on the BS, the ISI was approved for Install on August 17th.

The cartridge install was done on September 7th 2012, after which payload was finalized and vibration absorbtion elements installed. Finally, Phase II-b testing was completed on October 10th and the dome was close on October 11th.