## Slide 1
HAM-ISI Medm Screens and Troubleshooting
(introduction)
Hugo Paris for the SEI team

## Slide 2
MEDM Screens

## Slide 3: Main
Sensors
L4C
CPS
GS13
Basis change
Local => Cartesian
Processing
Sensor blend
Damping
Isolation
Basis change
Cartesian => Local
Actuators
ON/OFF
Safety
Notes: MEDM
Motif Editor and Display Manager (GUI for control screens)
GUI
Graphical User Interface
MEDM screens will change a bit. FFW line must be changed
Presentation from left to right

## Slide 4: Sensors
Calm environment:
L4C 0 +/- 400cts
CPS 0 +/- 400cts
GS13 +/- 12000

## Slide 5: Input Filters (INF)
L4C
Calibration
Gain
CPS
Calibration
GS13
Calibration
Symmetrization
Gain
Dewhitening

## Slide 6: Input Filters (INF)

## Slide 7: Coordinate Transform Matrices
Open Matlab®
Load Paths: HAM-ISI SEI
Set Variables:
  IFO = 'X1';
  Chamber = 'HAMX';
  Direction = 'X';
  Set_Gains = 1;
Run:
  Populate_MEDM_Screen_HAM_ISI(IFO,Chamber,Direction,Set_Gains)

## Slide 8: Damping

## Slide 9: Blend
CPS
Seismometer

## Slide 10: Blend
¥
Notes: MEDM
Motif Editor and Display Manager (GUI for control screens)
GUI
Graphical User Interface

## Slide 11: ISO Loops
We don’t know the perturbation:???
Measure the motion of the platform: Blend
Counteract with actuators: ISO Loops
???
Platform motion
ISO Loops
+ High performances
- Risks of Instability

## Slide 12: Watchdogs
Notes: GS13 tripped
  Look @ excitation
  Look @ Damping/Isolation. Stable???
L4C limit is over safe value
Make sure no “special measurement” is running and correct

## Slide 13: CPU check

## Slide 14: Conclusion
We want:
Input filters - ON
Coordinate Transform Matrices - Loaded
Damping loops - ON
Isolation filters - ON
WD - Safe
CPU - OK state
Guardian will help
Notes: MEDM
Motif Editor and Display Manager (GUI for control screens)
GUI
Graphical User Interface

## Slide 15
Troubleshooting

## Slide 16: Power Spectra
Broken Flexure…
DTT or Matlab®
Notes: MEDM
Motif Editor and Display Manager (GUI for control screens)
GUI
Graphical User Interface

## Slide 17: Transfer Functions
Measured/processed via Matlab®
More details in E1000309
    Run/Save excitation: Run_Exc_Batch_X1_HAMX.m.
    Retrieve time series + Process/Save TF data: Run_Get_Batch.m
    Concatenate/display/save TFS: Step_1_Plot_TF_Loc_to_Loc_X1_ISI_HAMX.m
We want
Symmetry
Low noise
Results comparable to reference
Notes: MEDM
Motif Editor and Display Manager (GUI for control screens)
GUI
Graphical User Interface

## Slide 18: Example #1
IN-Field cables misconnected
Bad connection on GS13 cables => Gain of 1/2
GS13 connections mismatched => reading cross-coupling
Notes: MEDM
Motif Editor and Display Manager (GUI for control screens)
GUI
Graphical User Interface

## Slide 19: Example #2
Mechanical issue
GS13 not torqued against the ISI
Notes: MEDM
Motif Editor and Display Manager (GUI for control screens)
GUI
Graphical User Interface

## Slide 20
Thank you