## Slide 1: Global longitudinal quad damping vs. local damping
Brett Shapiro
Stanford University
LIGO

## Slide 2: Summary
Background: local vs. global damping
Part I: global common length damping
  Simulations
  Measurements at 40 m lab
Part II: global differential arm length damping without OSEMs
  Simulations
  Measurements at LIGO Hanford
Conclusions
LIGO

## Slide 3: Usual Local Damping
ETMX
ETMY
Cavity control
uy,2
uy,3
uy,4
ux,2
ux,3
ux,4
0.5
0.5
-1
-1
Local damping
Local damping
The nominal way of damping
OSEM sensor noise coupling to the cavity is non-negligible for these loops.
The cavity control influences the top mass response.
Damping suppresses all Qs
ux,1
uy,1

## Slide 4: Common Arm Length Damping
ETMX
ETMY
Common length DOF independent from cavity control
The global common length damping injects the same sensor noise into both pendulums
Both pendulums are the same, so noise stays in common mode, i.e. no damping noise to cavity!
-0.5
+
Common length damping
uy,2
uy,3
uy,4
ux,2
ux,3
ux,4
-0.5
Cavity control
0.5
0.5
ux,1
uy,1

## Slide 5: Differential Arm Length Trans. Func.
ETMX
ETMY
*
*
The differential top mass longitudinal DOF behaves just like a cavity-controlled quad.
Assumes identical quads (ours are pretty darn close).
See `Supporting Math’ slides.
longitudinal
-
* Differential top to differential top transfer function
uy,2
uy,3
uy,4
ux,2
ux,3
ux,4
Cavity control
0.5
0.5
ux,1
uy,1
-0.5
-0.5

## Slide 6: Simulated Common Length Damping
Realistic quads - errors on the simulated as-built parameters are:
Masses: ± 20 grams
d’s (dn, d1, d3, d4): ± 1 mm
Rotational inertia: ± 3%
Wire lengths: ± 0.25 mm
Vertical stiffness: ± 3%
ETMX
ETMY
-0.5
+
Common length damping
uy,2
uy,3
uy,4
ux,2
ux,3
ux,4
-0.5
Cavity control
0.5
0.5
ux,1
uy,1

## Slide 7: Simulated Common Length Damping

## Slide 8: Simulated Damping Noise to Cavity
Red curve achieved by scaling top mass actuators so that TFs to cavity are identical at 10 Hz.

## Slide 9: Simulated Damping Ringdown

## Slide 10: 40 m Lab Noise Measurements
Seismic noise
Laser frequency noise
OSEM sensor noise

## Slide 11: 40 m Lab Noise Measurements
Ideally zero. Magnitude depends on quality of actuator matching.
Plant
Damp control
Cavity control
+
OSEM noise
cavity signal
Global common damping
Local ITMY damping
Ratio of local/global

## Slide 12: 40 m Lab Damping Measurements

## Slide 13: Differential Arm Length Damping
*
ETMX
ETMY
Control Law
0.5
0.5
*
If we understand how the cavity control produces this mode, can we design a controller that also damps it?
If so, then we can turn off local damping altogether.
longitudinal
0.5
0.5
-
* Differential top to differential top transfer function
uy,2
uy,3
uy,4
ux,2
ux,3
ux,4
?

## Slide 14: Differential Arm Length Damping
Pendulum 1
f2
x4
The new top mass modes come from the zeros of the TF between the highest stage with large cavity UGF and the test mass. See more detailed discussion in the ‘Supporting Math’ section.
This result can be generalized to the zeros in the cavity loop gain transfer functions (based on observations, no hard math yet).

## Slide 15: Differential Arm Length Damping

## Slide 16: Differential Arm Length Damping

## Slide 17: Differential Arm Length Damping

## Slide 18: Differential Arm Length Damping

## Slide 19: Differential Arm Length Damping
Test UGF: 300 Hz
PUM UGF: 50 Hz
UIM UGF: 10 Hz

## Slide 20: Differential Arm Length Damping
Test UGF: 300 Hz
PUM UGF: 50 Hz
UIM UGF: 5 Hz

## Slide 21: Differential Arm Length Damping
The top mass longitudinal differential mode resulting from the cavity loop gains on the previous slides. Damping is OFF!

## Slide 22: Differential Arm Length Damping
Top mass damping from cavity control. No OSEMs!

## Slide 23: LHO Damping Measurements Setup
M1
MC2 triple suspension
f2
x1
f3
-1
C2
C3
M2
M3
IMC Cavity signal
g2
Variable gain
f1
Test procedure
Vary g2 and observe the changes in the responses of x1 and the cavity signal to f1.
Terminology Key
IMC: input mode cleaner,
the cavity that makes the laser beam nice and round
M1: top mass
M2: middle mass
M3: bottom mass
MC2: Mode cleaner triple suspension #2
C2: M2 feedback filter
C3: M3 feedback filter

## Slide 24: LHO Damping Measurements
Terminology Key
M1: top mass
M2: middle mass
M3: bottom mass
MC2: triple suspension
UGF: unity gain frequency or bandwidth

## Slide 25: LHO Damping Measurements
Terminology Key
IMC: cavity signal, bottom mass sensor
M1: top mass
M2: middle mass
M3: bottom mass
MC2: triple suspension
UGF: unity gain frequency or bandwidth

## Slide 34: Conclusions
Very simple implementation. A matrix transformation and a little bit of actuator tuning.
Overall, global damping isolates OSEM sensor noise in two ways:
  1: common length damping -> damp global DOFs that couple weakly to the cavity
  2: differential length damping -> cavity control damps its own DOF
Can isolate nearly all longitudinal damping noise.
If all 4 quads are damped globally, the cavity control becomes independent of the damping design.
LIGO

## Slide 35: Conclusions cont.
Broadband noise reduction, both in band (>10 Hz) and out of band (<10 Hz).
Applies to other cavities, e.g. IMC.
Limitations:
  Only works for longitudinal, maybe pitch and yaw
  Adds an extra constraint to the cavity control, encouraging more drive on the lower stages
LIGO

## Slide 36: Acknowledgements
Caltech: 40 m crew, Rana Adhikari, Jenne Driggers, Jamie Rollins
LHO: commissioning crew
MIT: Kamal Youcef-Toumi, Jeff Kissel.
LIGO

## Slide 37: Backups
LIGO

## Slide 38: Differential Damping – all stages

## Slide 39: Supporting Math
Dynamics of common and differential modes
  Rotating the pendulum state space equations from local to global coordinates
  Noise coupling from common damping to DARM
  Double pendulum example
Change in top mass modes from cavity control – simple two mass system example.

## Slide 40: Dynamics of common and differential modes

## Slide 41: Rotating all ETMX and ETMY local long. DOFs into global diff. and comm. DOFs
Local to global transformations:
R = sensing matrix
n = sensor noise

## Slide 42: Rotating all ETMX and ETMY local long. DOFs into global diff. and comm. DOFs
Now, substitute in the feedback and transform to Laplace space:
Grouping like terms:
Determining the coupling of common mode damping to DARM

## Slide 43: Rotating all ETMX and ETMY local long. DOFs into global diff. and comm. DOFs
Solving c in terms of d and :
Plugging c in to d equation:
Defining intermediate variables to keep things tidy:
Then d can be written as a function of :

## Slide 44: Rotating all ETMX and ETMY local long. DOFs into global diff. and comm. DOFs
Then the transfer function from common mode sensor noise to DARM is:
As the plant differences go to zero, N goes to zero, and thus the coupling of common mode damping noise to DARM goes to zero.

## Slide 45: Simple Common to Diff. Coupling Ex.
To show what the matrices on the previous slides look like.
ETMX
ETMY
DARM Error
ux1
mx1
my1
mx2
my2
kx1
kx2
ky1
ky2
0.5
0.5
+
Common damping
uy1
c1
d2
x1
x2

## Slide 47: Simple Common to Diff. Coupling Ex

## Slide 48: Simple Common to Diff. Coupling Ex
Plugging in sus parameters for N:

## Slide 49: Change in top mass modes from cavity control – simple two mass system example.

## Slide 50: Change in top mass modes from cavity control – simple two mass ex.
m1
m2
k1
k2
x1
x2
f2
Question: What happens to x1 response when we control x2 with f2?
When f2 = 0,
The f1 to x1 TF has two modes
f1

## Slide 51: Change in top mass modes from cavity control – simple two mass ex.
This is equivalent to
m1
m2
k1
k2
x1
x2
C
As we get to C >> k2, then x1 approaches this system
m1
k1
k2
x1
The f1 to x1 TF has one mode. The frequency of this mode happens to be the zero in the TF from f2 to x2.

## Slide 52: Change in top mass modes from cavity control – simple two mass ex.
Discussion of why the single x1 mode frequency coincides with the f2 to x2 TF zero:
The f2 to x2 zero occurs at the frequency where the k2 spring force exactly balances f2. At this frequency any energy transferred from f2 to x2 gets sucked out by x1 until x2 comes to rest. Thus, there must be some x1 resonance to absorb this energy until x2 comes to rest. However, we do not see x1 ‘blow up’ from an f2 drive at this frequency because once x2 is not moving, it is no longer transferring energy. Once we physically lock, or control, x2 to decouple it from x1, this resonance becomes visible with an x1 drive.
The zero in the TF from f2 to x2. It coincides with the f1 to x1 TF mode when x2 is locked.
m2
fk2
k2
x2
f2
…

## Slide 53: Change in top mass modes from cavity control – Full Quad example.

## Slide 54: Cavity Control Influence on Damping
ETMX
ETMY
uy,2
uy,3
uy,4
ux,2
ux,3
ux,4
* Top to top mass transfer function
*
- Case 1: All cavity control on Pendulum 2
What you would expect – the quad is just hanging free.
Note: both pendulums are identical in this simulation.
longitudinal
Cavity control
1
0

## Slide 55: Cavity Control Influence on Damping
ETMX
ETMY
uy,2
uy,3
uy,4
ux,2
ux,3
ux,4
* Top to top mass transfer function
*
- Case 2: All cavity control on Pendulum 1
The top mass of pendulum 1 behaves like the UIM is clamped to gnd when its ugf is high.
Since the cavity control influences modes, you can use the same effect to apply damping (more on this later)
longitudinal
Cavity control
0
1

## Slide 56: Cavity Control Influence on Damping
ETMX
ETMY
uy,2
uy,3
uy,4
ux,2
ux,3
ux,4
* Top to top mass transfer function
*
- Case 3: Cavity control split evenly between both pendulums
The top mass response is now an average of the previous two cases -> 5 resonances to damp.
Control up to the PUM, rather than the UIM, would yield 6 resonances.
aLIGO will likely behave like this.
longitudinal
Cavity control
0.5
0.5

## Slide 57: Global Damping RCG Diagram

## Slide 59: Cavity loop gain measurements on a triple with 2 stages of actuation
P32
P33
+
C2
C3
+
+
c
D2
D3
e3
e2
If we measure the stage 3 loop gain by driving D3, we get
If we measure the stage 2 loop gain by driving D2, we get

## Slide 60: Scratch

## Slide 62: Scratch: Rotating all ETMX and ETMY local long. DOFs into global diff. and comm. DOFs
For DARM we measure the test masses with the global cavity readout, no local sensors are involved. The cavity readout must also have very low noise to measure GWs. So make the assumption that nx-ny=0 for cavity control and simplify to:
Now, substitute in the feedback and transform to Laplace space: