# LIGO Laboratory / LIGO Scientific Collaboration

LIGO- E1201012 - v8                     Advanced LIGO                      9-10-13

TransMon-SUS Assembly & Installation Procedure

Michael Smith, Ken Mailand, Craig Conley, Keita Kawabe, Bram Slagmolen, Lisa C. Austin

Distribution of this document:

DCC

# This is an internal working note

of the LIGO Project.

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 1970 Mail Stop S9-02 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

http://www.ligo.caltech.edu/

# Table of Contents

LIGO Laboratory / LIGO Scientific Collaboration	1

This is an internal working note	1

Table of Contents	2

1	Introduction	3

1.1	Purpose and Scope	3

1.2	References	3

2	Preparation	3

2.1	Receiving/inventory	3

2.1.1	Receiving/inventory of metal parts	3

2.2	Cleaning/Baking	3

2.3	HeliCoils	3

3	D0901880: aLIGO AOS Transmission Monitor System Assembly Overview	3

3.1	TMS Principal Assemblies	3

3.2	TMS Accessory Assemblies	3

3.3	Subassembly Part Number Assignment	3

3.4	Weighing Subassemblies	3

4	D1101064: aLIGO TMS FIRST ARTICLE & PRODUCTION INSTALLATION AND ASSEMBLY TOOLING	3

4.1	D1101361: TMS TOOLING TELESCOPE I-P ALIGNMENT MIRROR ASSEMBLY	3

4.2	D1101487: TMS TOOLING TELESCOPE O-P ALIGNMENT MIRROR ASSEMBLY	3

4.3	D1100933: aLIGO TMS TELESCOPE on BENCH SUPPORT ASSEMBLY (2 Req'd)	3

4.4	D1100531: aLIGO TMS OPTICAL TABLE SUPPORT BRIDGE	3

for alignment of (Telescope With Optical Table)	3

4.5	D1100649: aLIGO TMS TELESCOPE + OPTICAL TABLE,  BALANCE BRIDGE ASSEMBLY	3

4.6	D1100841: aLIGO TMS TELE W/OPT TABLE, INSTALL TOOL ASSEMBLY	3

4.7	D1100613: aLIGO TMS TELESCOPE ASSEMBLY ALIGNMENT STABILIZATION TOOL	3

4.8	D1101510: aLIGO TMS TELE ALIGNMENT 6in MIRROR MOUNT ASSY	3

4.9	D1100460: aLIGO TMS TELESCOPE ALIGNMENT,  SUPPORT BRIDGE ASSEMBLY	3

4.10	D1100886: MODIFIED GENIE LIFT FORKS	3

4.11	D1100807:  aLIGO TMS Setup Table Bosch-Frame Assembly	3

4.12	Other Tools	3

5	Assembly Procedure for TMS	3

5.1	Assembly Procedure for TOP MASS ASSEMBLY	3

5.1.1	D1200453 aLIGO TMS Upper Structure Weldment	3

5.1.2	D060324 SUS, Quad N-Ptype Top Stage, Blade Cartridge	3

5.1.2.1	Pre-loading of Top Mass Suspension Blades	3

5.1.3	D1101526 aLIGO TMS Mass Balanced Assembly. Note: this is the Production Assembly    (all but the WBSC6)	3

5.1.4	D060310 ADVANCED LIGO, SUS, QUAD N-PTYPE TABLECLOTH	3

5.1.4.1	D060311 Quad N-Ptype Tablecloth, Front and Rear Plate, Noise Tablecloth	3

5.1.4.2	D060313 Quad N-Ptype Tablecloth, Top Plate, Noise Tablecloth	3

5.1.4.3	D060314 Quad N-Ptype Tablecloth, Left Hand Plate, Noise Tablecloth	3

5.1.4.4	D060315 Quad N-Ptype Tablecloth, Tablecloth Pinch Plate, Upper Structure	3

5.1.4.5	D060316 Quad N-Ptype Tablecloth, OSEM and ECD Adjuster, Noise Tablecloth	3

5.1.4.6	D060320 Quad N-Ptype Tablecloth, Right Hand Plate, Noise Tablecloth	3

5.1.4.7	D060322 Quad N-Ptype Tablecloth, Right Hand Plate, Noise Tablecloth	3

5.1.4.8	D060218 BOSEM Assembly	3

5.1.5	D1101166 TMS Upper Suspension Wire Assembly	3

5.1.5.1	T1200388 TMS Wire Jig Assembly	3

5.2	D1100827: VERTICAL SAFETY WIRE ASSEMBLY	3

5.3	D1101130: aLIGO TMS TELESCOPE SAFETY SUPPORT BEAM ASSY	3

5.4	D1001891: aLIGO TMS VERTICAL SEISMIC SAFETY STOP ASSEMBLY	3

5.5	D1101163: aLIGO TMS Telescope SUS Wire Assembly	3

5.6	D1000484: ISC Transmon Assembly	3

5.7	D1102291: aLIGO TMS ISC Table Structural Assembly	3

5.8	D1102361: aLIGO TMS Telescope Frame Assembly	3

5.8.1	Parts List for TMS Telescope Frame Assembly	3

5.8.2	T1300639-v1 TMS Primary Mirror Installation Procedure	3

5.8.3	Assembly of D1000243-v3 aLIGO_ETM_TELE_TUBE_AND_INSERTS_ASSEMBLY	3

5.8.4	Assembly of D1200244: aLIGO TMS F1 Mirror Assembly	3

5.8.5	Assembly of D1102365: aLIGO TMS Telescope Secondary Mirror Assembly	3

5.8.5.1	Assembly of D1201077-aLigo Secondary Mirror Mount Tip-tilt Sub-assembly	3

5.8.5.2	Assembly of End Telescope Plates and Spacer Tubes	3

5.8.5.3	Assembly of FM2	3

5.8.5.4	Assembly of Primary Mirror	3

5.8.5.5	Assembly of FM2 and Secondary Mirror	3

5.8.6	Initial Alignment of TMS Telescope	3

5.8.7	T1100603: TMS Telescope Alignment Procedure	3

5.8.8	Mounting the ISC Table Assembly to the TMS Telescope	3

5.8.9	Balancing the TMS Telescope Frame Assembly /ISC Table Assembly	3

5.8.9.1	Assemble the TMS Tele-opt Table Balance Bridge Assembly, D1100649	3

5.8.9.2	Prepare the Pin-in-Socket Balance Point	3

5.8.9.3	Support the ISC Table / TMS Telescope Assembly with the Balance Bridge	3

5.8.9.4	Balance the ISC Table / TMS Telescope Assembly, See Table Assembly for Nominal Position of Trim Masses. First Article 80kg., Production 82kg.	3

5.8.10	Suspending the TMS Telescope Frame Assembly/ISC Table Assembly from the TMS Optical Table Support Bridge Assy D1100531	3

5.8.11	Mounting the AOS ETM Telescope Upper Suspension Frame and Top Mass to the Tele-opt Table (Bosch Frame)	3

5.8.12	TMS Test Mass Assembly, D1002097	3

5.8.12.1	Assembly of the Test Mass	3

5.8.12.2	Balancing the Test Mass	3

5.8.12.3	Installing the Test Mass: 80kg First Article (WBSC6) Production 82kg	3

5.8.12.4	Adjust the TMS Control System with Suspended Test Mass	3

5.8.12.5	Removing the Test Mass	3

5.8.13	Mounting the TMS Telescope Frame Assembly/ISC Table Assembly to the Top Mass	3

5.8.14	Adjustment of TMS Control System will Suspended TMS Telescope/ISC Table Assembly	3

6	TMS Installation Procedure	3

6.1	Transporting and Mounting the TMS System to the ISI Cartridge	3

6.1.1	Remove TMS Telescope/ISC Table Assembly from the Upper Suspension Frame and Top Mass	3

6.1.2	Prepare Upper Suspension Frame and Top Mass	3

6.1.3	Attach Upper Suspension Frame and Top Mass to Cartridge	3

6.1.4	Attach TMS Telescope/ISC Table Assembly to the Top Mass Suspension Wires	3

6.2	Final Installation inside BSC	3

6.2.1	Preparing the TMS System for Cartridge Flight into the BSC	3

6.2.1.1	TMS Transportation-Installation Restraint Procedure, E1100841	3

6.2.2	Cartridge Flight into the BSC	3

6.2.3	TMS Seismic Safety Stop Structure Attachment to TMS	3

6.3	Pull Back Procedure for Maintenance of ETM	3

6.3.1	Disconnect the Vertical Safety Wire Assembly	3

6.3.2	Disengage Earthquake Restraint from the TMS	3

6.3.2.1	Attach the Seismic Safety Stop Chain	3

6.3.2.2	Free the TMS Stop Rod	3

6.3.3	Attach Vertical Safety Chains and Pull Back	3

7	Appendix A: Useful Balancing and Alignment Data	3

The Telescope and Optical Table Assembly First Article Assembly: Total mass 80 kg	3

The Telescope and Optical Table Assembly Production Assembly: Total mass 82 kg	3

Appendix C: Related Documents	3

8	TMS Swing Stop Tool Used on Cartridge Installation w/ETM	3

9	TMS BEAM CENTER LINE DISTANCE FROM CARTRIDGE TABLE	3

10	TMS TOOLING USED DURING VARIOUS STAGES OF INSTALL ON BOSCH FRAME AND CARTRIDGE	3

# Introduction

## Purpose and Scope

This document describes the assembly and installation procedure for the TransMonSus (TMS). Open DCC for Latest Drawing Versions and Notes

## References

- LIGO-E1100841: TMS Transportation-Installation Restraint Procedure

- LIGO-E1000097: ETM TransMon Telescope SUS Installation Plan

- LIGO-E1100651: TMS Top Mass Assembly Procedure

- LIGO-D1100908: aLIGO TMS TELE-OPT TABLE TOOL ASSY-GENIE FORK CONNECTION

- LIGO-D1101307: TOOLING LAYOUT-aLIGO TMS Telescope Cartridge lift Safety Supports

- LIGO-D1100807: aLIGO TMS TELE-OPT TABLE INSTALL, FRAME-PLATE ASSY

- LIGO-D1101096: aLIGO TMS OPT TABLE-TELE- SUPPORT BRIDGE CONNECTION

- LIGO-D1101095: aLIGO TMS TELESCOPE-SUPPORT BRIDGE CONNECTION

- LIGO-D1101097: aLIGO TMS TELE-OPT TABLE-BALANCE BRIDGE interface

- LIGO-E1200049: aLIGO Transmission Monitor Suspension Installation Plan

- LIGO-E1200453: Fine Initial Alignment Procedure of the Transmission Monitor Telescope

- LIGO-T1100600: Requirement For Adjustment Procedure Of Advanced LIGO Transmission Monitor Telescope

- LIGO-T1100603: TMS Telescope Alignment Procedure

- LIGO-G1101254: TMSY Telescope focal tuning results.

- LIGO-T1100604: TMS Change Suggestions after First Build

- LIGO-D1101096: aLIGO TMS OPT TABLE-TELE- SUPPORT BRIDGE CONNECTION

- LIGO-E1200793: aLIGO TMS System Drawing Assembly Tree

- LIGO-D1002097: aLIGO_TMS_TEST_MASS_ASEM

# Preparation

Advanced LIGO has implemented a new Inventory Control System (ICS) which is designed to record all aLIGO hardware as it moves through receiving, inspection, clean, bake, storage, shipment, and assembly processes.  The ICS is meant to replace the shipping type paper traveler used in iLIGO.  While the ICS is still in final development as of this writing, the hope is that the engineering teams will be able to utilize ICS to record many aspects of the lifetime of a part from its initial receipt through the clean and bake processes previously documented in the iLIGO traveler.  The sites have dedicated staff to help with managing the data related to the processing of parts in ICS.  Engineering staff should become familiar with the ICS such that they can utilize it for their own record keeping and data management.  If the ICS fails to facilitate data that you need to record, process travelers (PT) can be placed on the DCC.  In either case, make sure to record all serial numbers and data in the ICS or the DCC during the following steps.

## Receiving/inventory

### Receiving/inventory of metal parts

Upon receipt of shipments of SUS parts, the following steps should be performed:

Basic inspection by the receiver prior to unpacking the shipment (crate damage, etc).  Packing slips should be sent (hardcopy or emailed) to Jennie Murdock at LHO.  Person performing this step should notify site subassembly lead of the shipment arrival.

Inventory Control and inspection performed by ICS person and site subsystem lead as parts are unpacked.  Drawing numbers, serial numbers, and quantities will be imported into the ICS database via spreadsheet templates (F0900052).  This is a good time for QA/QC and engineering inspections.  The following processes can now be recorded in ICS by grouping the parts into Loads.

Parts get separated into cleaning loads based on their level of cleanliness, and moved to the appropriate cleaning station.

Parts get separated into clean and bake loads based on their material – see E960022.  Sorting should be reflected in the Load records in ICS, where instructions to technicians can be added for any special handling or material considerations.

Parts will be processed as per E960022.

Parts will be stored in clean storage areas until assembly.

## Cleaning/Baking

Process all parts except for the Test Masses as Class A per E960022.  Test Mass D1002097 is to be processed as Class B, since it is used as tooling and will be removed.  All Parts should be processed as Class A or B prior to Helicoil installation.

## HeliCoils

Install all the HeliCoils in all the parts and make sure they are free running and not cross threaded, remove the tangs.  Perform HeliCoil inspections as per the instructions listed on the Advanced LIGO HeliCoil Wiki page for installation and inspection.

# D0901880: aLIGO AOS Transmission Monitor System Assembly Overview

## TMS Principal Assemblies

The Transmission Monitor Suspension (TMS) is an in-vacuum component that is located behind the ETM (End Test Mass) and is mounted to the internal seismic isolation (ISI) optics platform inside the BSC, as shown in Figure 1. The TMS Telescope reduces the size of the interferometer (IFO) beam transmitted through the ETM. The reduced TMS beam is used for monitoring the alignment of the IFO beam in the arms. The TMS is also used for injection of the green laser beam used in the Arm Length Stabilization (ALS) scheme. Finally, the TMS also accommodates probe beams for the Hartmann sensor that monitors the curvature of the ETM HR surface.

The TMS consists of three principal assemblies: 1) suspension frame and top mass with springs and controls, 2) ISC TMS Optical Assembly, and 3) TMS Telescope.

NOTE: SEE SECTION 9.0 FOR BEAM DISTANCE FROM CARTRIDGE TABLE

TMS Telescope

ISC TMS Optical Assembly

Suspension Frame and Top Mass with Springs and Controls

Figure 1: TMS shown behind the ETM quad

## TMS Accessory Assemblies

In addition to the principal assemblies, the following accessory assemblies are also part of the TMS.

D1100827: ALIGO VERTICAL SAFETY WIRE ASSEMBLY

D1101130: ALIGO TMS TELESCOPE SAFETY SUPPORT BEAM ASSY

D1001891: ALIGO TMS VERTICAL SEISMIC SAFETY STOP ASSEMBLY

D1101163: ALIGO TMS TELESCOPE SUS WIRE ASSEMBLY

Note: 	In Cartridge assembly D1300448 TMS Cartridge Swing Stop Tool is used to prevent TMS and ETM from Bumping.

D1000484: TMS ISC Table Structural Assembly

D1102361: aLIGO TMS Telescope Frame Assembly

D1101163: TMS Telescope SUS Wire Assembly

D1100827: aLIGO Vertical Safety Wire Assembly

D1101130: aLIGO TMS Telescope Safety Support Beam Assy

D1001891: aLIGO TMS Vertical Seismic Safety Stop Assembly

D1000484: ISC Transmon Assembly

D1000549: aLIGO Upper Intermediate Mass Top Assembly

Figure 2: TMS Principal and Accessory Assemblies

## Subassembly Part Number Assignment

Each subassembly will be assigned a unique serial number based on the parent number which can be used for referencing data taken on that subassembly.  For example, if 3 UPPER INTERMEDIATE MASS TOP Assemblies are assembled from drawing number D1000549, the units should be assigned serial numbers like:

D1000549-001

D1000549-002

D1000549-003, and so on.

As individual parts are added to the subassembly, record their serial numbers as part of that subassembly.  The overall subassembly number (i.e. D1000549-001) can now be used in the ICS to track further operations performed on that subassembly.  These subassemblies will eventually become associated with their parent D0901880: aLIGO AOS Transmission Monitor System Assembly, which will have its own serial number, such as D0901880-001 (aka TMS 001).  Label the bag with the newly designated subassembly and serial number after wrapping and bagging.

## Weighing Subassemblies

When weighing subassemblies, use the high precision scale dedicated for the SUS assemblies.

# D1101064: aLIGO TMS FIRST ARTICLE & PRODUCTION INSTALLATION AND ASSEMBLY TOOLING

Table 1: D1101064: aLIGO TMS FIRST ARTICLE & PRODUCTION INSTALLATION AND ASSEMBLY TOOLING

## D1101361: TMS TOOLING TELESCOPE I-P ALIGNMENT MIRROR ASSEMBLY

## D1101487: TMS TOOLING TELESCOPE O-P ALIGNMENT MIRROR ASSEMBLY

## D1100933: aLIGO TMS TELESCOPE on BENCH SUPPORT ASSEMBLY (2 Req'd)

## D1100531: aLIGO TMS OPTICAL TABLE SUPPORT BRIDGE

## for alignment of (Telescope With Optical Table)

## D1100649: aLIGO TMS TELESCOPE + OPTICAL TABLE,  BALANCE BRIDGE ASSEMBLY

## D1100841: aLIGO TMS TELE W/OPT TABLE, INSTALL TOOL ASSEMBLY

See: D1100908 For application w/ Genie Lift

Note: Use D1200355 optional Height Spacers + longer screws, under #2, for added Working Room between underside of Tele -Table and Plate #1

## D1100613: aLIGO TMS TELESCOPE ASSEMBLY ALIGNMENT STABILIZATION TOOL

## D1101510: aLIGO TMS TELE ALIGNMENT 6in MIRROR MOUNT ASSY

## D1100460: aLIGO TMS TELESCOPE ALIGNMENT,  SUPPORT BRIDGE ASSEMBLY

## D1100886: MODIFIED GENIE LIFT FORKS

See: D1100908 For application w/ Genie Lift

## D1100807:  aLIGO TMS Setup Table Bosch-Frame Assembly

## Other Tools

Ian’s list

| Test stand: | Mechanical Test Stand mounted with Solid Stack Assembly |
| --- | --- |
| Manual fork truck: | Similar to Caltech Genie |
| Bench: | May be an optics bench but this is not mandatory |
| Tools: | All the appropriate hand tools and measuring devices |
| Masses: | These will be necessary to load blades flat. |
| Lower structure assembly tooling: |  |
| Wire jig: |  |

Brett’s list

inch nut driver or wrench for axial OSEM positioning.

inch nut driver or wrench for lateral OSEM positioning.

, , , and  inch allen wrenches.

A flat head screw driver for turning the top mass pitch adjusters.

Torque wrench for the blade clamp bolts capable of 400 in-lb (33 ft-lbs, 45 Nm).

Slip or block gauges for measuring 5 mm, 9.6 mm, and 15 mm gaps.

Dentist Mirror.

Flashlight or small lamp.

Structure pushers for rotating the structure on the optical table (see Figure 14).

5 axis table for safety while rotating the structure.

Lower structure tooling for use with the 5 axis table.

Safety goggles for working around the wires.

An optical alignment tool with 10 Rad accuracy, such as an autocollimator.

A small, light, reliable level to place on suspended masses (optional).

# Assembly Procedure for TMS

## Assembly Procedure for TOP MASS ASSEMBLY

The D1000549 Upper Intermediate Mass Top Assembly consists of:

- D1200453 aLIGO TMS Upper Structure Weldment

- D060324 SUS, Quad N-Ptype Top Stage, Blade Cartridge

- D1101526 aLIGO TMS Mass Balanced Assembly

- D060310 Advanced LIGO, SUS, QUAD N-PType Tablecloth

- D1101166 TMS Upper Suspension Wire Assembly

In addition, it comprises the Control System Hardware for controlling the TMS Telescope Assembly.

Refer to E1100651

### D1200453 aLIGO TMS Upper Structure Weldment

The aLIGO TMS Upper Structure Weldment is shown in Figure 3.

Figure 3: TMS Upper Structure Weldment

### D060324 SUS, Quad N-Ptype Top Stage, Blade Cartridge

Figure 4: Quad N-Ptype Top Stage, Blade Cartridge Assembly

#### Pre-loading of Top Mass Suspension Blades

The pre-load (blade straightening) of the Top Mass Suspension Blade is applied by using the Straightening Tooling Assy - Quad Top Stage Blade Tooling D060370, as shown in Figure 5 and Figure 6. The top blade springs are each pre-loaded with approximately 140 lbs.

Figure 5: Insertion of Blade in Straightening Fixture

Figure 6: Straightening of Blade Spring

After straightening the blade, it is inserted into the cartridge holding assembly. The completed blade cartridge assembly is shown in Figure 7.

Figure 7: Completed Blade Cartridge Assembly

### D1101526 aLIGO TMS Mass Balanced Assembly. Note: this is the Production Assembly    (all but the WBSC6)

Top Mass Assembly Drawings in Order:

1. D1100525

2. D1200527

3. D1200526

The assembly procedure is described in D060370 Picture Book.

Figure 8: aLIGO TMS Mass Balanced Assembly

NOTE: THIS ASSEMBLY (D1000444) IS FIRST ARTICLE (WBSC6)

### D060310 ADVANCED LIGO, SUS, QUAD N-PTYPE TABLECLOTH

#### D060311 Quad N-Ptype Tablecloth, Front and Rear Plate, Noise Tablecloth

#### D060313 Quad N-Ptype Tablecloth, Top Plate, Noise Tablecloth

#### D060314 Quad N-Ptype Tablecloth, Left Hand Plate, Noise Tablecloth

#### D060315 Quad N-Ptype Tablecloth, Tablecloth Pinch Plate, Upper Structure

#### D060316 Quad N-Ptype Tablecloth, OSEM and ECD Adjuster, Noise Tablecloth

#### D060320 Quad N-Ptype Tablecloth, Right Hand Plate, Noise Tablecloth

#### D060322 Quad N-Ptype Tablecloth, Right Hand Plate, Noise Tablecloth

#### D060218 BOSEM Assembly

##### D060216 IRLED Assembly

##### D060217 PD Assembly

### D1101166 TMS Upper Suspension Wire Assembly

- Safety glasses must be worn during all wire work. Safety glasses are provided in the garbing areas and in all clean rooms.

- Glove liners should be worn under latex clean-room gloves, as a protective layer and an extra barrier. For information on glove liners, see the Contamination Control Plan, E0900047, page 13.

- Take the end of the wire and bend a small section, say 3” or so, for easier holding. The bent wire section can be hooked around your thumb, and held by your index finger.

- Un-spool the proper length of wire, including extra for handling, and control the area of the wire that needs to be cut.  Add a bend at the other end, if handling it that way is easier for you.

- Clean the end with the first bend. Change your gloves and grab the cleaned, bent end around your thumb.

- Prepare the cleaning wipes with methanol. Wipe wire clean with methanol changing wipes until the wire is completely clean. Clean the wire while it is coiled; do not stretch the wire until it is taut for cleaning. It can be laid down on a clean surface during this process. Clean a section at a time.

Transfer the wire to the wire jig. Use the wire jig clamps to hold down the wire. Cut the bent ends off and remove, after the wire is secured, according to the wire assembly procedure, T1000674-v2.

Figure 9: TMS Upper Suspension Wire Assembly

#### T1200388 TMS Wire Jig Assembly

The TMS Upper Suspension Wire will be stretched under load and set to the correct length by using a modified version of the D060517 Quad Wire Jig Assembly shown in Figure 10.

The modified version of the TMS wire jig assembly is shown in Figure 11, Figure 12, and Figure 13.

Figure 10: Quad Wire Jig Assembly

Figure 11: Modified TMS Wire Jig Assembly

Figure 12: Modified TMS Wire Jig Assembly, Clamping Detail

Figure 13: Modified TMS Wire Jig Assembly, with Loading Weight

## D1100827: VERTICAL SAFETY WIRE ASSEMBLY

The suspension wire segments are cut to desired length, cleaned and then assembled into clamp-wire-clamp assemblies before they are attached to the full TMS assembly, according to the wire assembly procedure, T1000674-v2.

## D1101130: aLIGO TMS TELESCOPE SAFETY SUPPORT BEAM ASSY

## D1001891: aLIGO TMS VERTICAL SEISMIC SAFETY STOP ASSEMBLY

This assembly is used only during alignment and servicing of the ETM.

## D1101163: aLIGO TMS Telescope SUS Wire Assembly

The assembly procedure is the same as described in D1101166 TMS Upper Suspension Wire Assembly—see section 5.1.5.

## D1000484: ISC Transmon Assembly

The ISC group is responsible for the ISC Transmon Assembly D1000484, which is placed on top of the TMS ISC Table Structural Assembly D1102291, as shown in .

Table 2: ISC Transmon Assembly

## D1102291: aLIGO TMS ISC Table Structural Assembly

Table 3: BOM TMS ISC Table Structural Assembly

| ITEM NO. | PART NUMBER | DESCRIPTION | MATERIAL | Default/REQ |
| --- | --- | --- | --- | --- |
| 1 | D1102293 | aLIGO ETM OPTICAL TABLE | 6061-T6 Al | 1 |
| 2 | D1001163 | aLIGO OPTICS TABLE GUSSET PLATE, LARGE | 6061-T6 Al | 2 |
| 3 | - | SCREW, SOCKET HEAD CAP, 1/4-20 UNC-2A X 0.625 LONG | 300 SSTL | 46 |
| 4 | - | SCREW, SOCKET HEAD CAP, 5/16-18 UNC-2A X 0.875 LONG | 300 SSTL | 8 |
| 5 | - | SCREW, SOCKET HEAD CAP, 1/4-20 UNC-2A X 0.75 LONG | 300 SSTL | 14 |
| 6 | - | SCREW, SOCKET HEAD CAP, 1/4-20 UNC-2A X 0.5 LONG | 300 SSTL | 12 |
| 7 | .250 VENT FL WSHR | .255 ID, .468 OD, .032 TH | 18-8 SSTL | 98 |
| 8 | .312 VENT FL WSHR | .328 ID, .562 OD, .032 TH | 18-8 SSTL | 12 |
| 9 | #10 VENT FL WSHR | .195 ID, .354 OD, .032 TH | 18-8 SSTL | 2 |
| 10 | D1200310 | aLIGO OPTICS TABLE SPEADER BRIDGE | 6061-T6 Al | 2 |
| 11 | D1001164 | aLIGO OPTICS TABLE GUSSET PLATE, SMALL | 6061-T6 Al | 2 |
| 12 | D1002696 | aLIGO ETM TELESCOPE BRIDGE SUPPORT TEE PLATE, FRONT | 6061-T6 Al | 1 |
| 13 | D1001109 | aLIGO ETM TELESCOPE BRIDGE SUPPORT TEE PLATE, BACK | 6061-T6 Al | 1 |
| 14 | D1002755 | aLIGO TMS TELESCOPE ROLL SLIDE ROD | 304 SSTL | 2 |
| 15 | - | SCREW, SOCKET HEAD CAP, 1/4-20 UNC-2A X 0.75 LONG, VENTED | Ag- PLATED 300 SSTL | 4 |
| 16 | D1002698 | aLIGO TMS TELESCOPE ROLL SLIDE MASS | 304 SSTL | 1 |
| 17 | D1002759 | aLIGO TMS TELESCOPE ROLL LOCK | 6061-T6 Al | 1 |
| 18 | - | SCREW, SOCKET HEAD CAP, 1/4-20 UNC-2A X 0.75 LONG | Ag- PLATED 300 SSTL | 6 |
| 19 | - | SCREW, SOCKET HEAD CAP, 1/4-20 UNC-2A X 1 LONG | Ag- PLATED 300 SSTL | 4 |
| 20 | 99934A760 (McM-Carr) or equiv | SKT SET SCRW, 1/4-20 X 3/8 L, SILVER TIP | 18-8 SSTL | 1 |
| 21 | D1002768 | aLIGO TMS TELESCOPE ROLL MASS ADJ. ROD | 316 SSTL | 1 |
| 22 | 8492A167 (McM-Carr) or equiv | DRILL BUSHING, PRESS FIT 5/16" ID, 1/2" OD, 3/8" LENGTH | Alloy Steel | 1 |
| 23 | 60705K46 (McM-Carr) or equiv | THREADED COLLAR, CLAMP ON 5/16"-18 UNC, 11/16" OD, 9/32" WIDTH | 300 SSTL | 2 |
| 24 | D1002769 | aLIGO TMS TELESCOPE ROLL SLIDE BKT. | 6061-T6 Al | 1 |
| 25 | BU-2406-NA (U-C Components) or equiv | BUTTON HEAD SOCKET SCREW 10-24 X 3/8 | Ag- PLATED 300 SSTL | 2 |
| 26 | D1002758 | aLIGO TMS TELESCOPE PITCH GUIDE | 304 SSTL | 2 |
| 27 | D1002699 | aLIGO PITCH SLIDE MASS | 304 SSTL | 2 |
| 28 | D1002757 | aLIGO TMS TELESCOPE PITCH SLIDE BKT. | 6061-T6 Al | 2 |
| 29 | D1003162 | aLIGO TMS TELESCOPE PITCH MASS ADJ. ROD | 316 SSTL | 2 |
| 30 | 9421T300 (McM-Carr) or equiv | 1/4" ID SHAFT COLLAR | 316 SSTL | 2 |
| 31 | D1001161 | aLIGO OPTICS TABLE TOP MODIFIED EYEBOLT | 304 SSTL | 2 |
| 32 | D1100989 | NICKEL-COPPER ALLOY 400 .312-18 HEX NUT, MODIFIED | NICKEL- COPPER ALLOY 400 | 2 |
| 33 | D1200470 | aLIGO TMS OPTICS TABLE CYLINDER MASS | 304 SSTL | 6 |
| 34 | D1100358 | TMS TELESCOPE MASS ATTACHMENT SCREW | TITANIUM | 7 |
| 35 | D1200471 | aLIGO TMS OPTICS TABLE SMALL CYLINDER MASS | 304 SSTL | 1 |
| 36 | D1100904 | aLigo_TMS_Periscope_Mount_Clamp | 6061-T6 | 6 |
| 37 | D1001435 | TMS, SISKIYOU MOUNT BRACKET, LOWER INBOARD | 6061-T6 Al | 1 |
| 38 | D1002504 | TMS, SISKIYOU MOUNT BRACKET, LOWER OUTBOARD | 6061-T6 Al | 1 |

NOTE: SEE SECTION 11. FOR ROLL AND PITCH ADJUSTER PROCEDURE

## D1102361: aLIGO TMS Telescope Frame Assembly

### Parts List for TMS Telescope Frame Assembly

### T1300639-v1 TMS Primary Mirror Installation Procedure

Table 4: Bill of Material, D1102361: aLIGO TMS Telescope Frame Assembly

| ASSEMBLY - 1 | QTY | ASSEMBLY - 2 | QTY | ASSEMBLY - 3 | QTY | PART NUMBER | QTY | DESCRIPTION |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| D1102361 | 1 | D1000243 | 3 |  | 1 | C-2006 | 6 | SCREW, SOCKET HEAD CAP, 1/4-20 UNC-2A x 0.375 LONG, VENTED |
| D1102361 | 1 | D1000243 | 3 |  | 1 | D1000072 | 1 | ADLIGO ETM TELE TUBE |
| D1102361 | 1 | D1000243 | 3 |  | 1 | D1000242 | 2 | aLIGO ETM TELE TUBE THD INSERT |
| D1102361 | 1 | D1102365 | 1 | D1200246 | 1 | BU-404 | 7 | BUTTON HEAD SOCKET SCREW 4-40 x .25, vented |
| D1102361 | 1 | D1102365 | 1 | D1200246 | 1 | D1200247 | 1 | aLIGO TMS FOCUS Carriage |
| D1102361 | 1 | D1102365 | 1 | D1200246 | 1 | D1200248 | 1 | aLIGO TMS FOCUS SLIDE |
| D1102361 | 1 | D1102365 | 1 | D1201077 | 1 | 92778A121 | 1 | 18-8 SS Oval Point Socket Set Screw 10-32 Thread, 1/2" Length |
| D1102361 | 1 | D1102365 | 1 | D1201077 | 1 | D0901565 | 1 | ADLIGO_ETM_TELE_SECONDARY |
| D1102361 | 1 | D1102365 | 1 | D1201077 | 1 | D1100114 | 3 | TMS Telescope Secondary Mirror Clamp |
| D1102361 | 1 | D1102365 | 1 | D1201077 | 1 | D1200826 | 1 | aLIGO secondary mount plate |
| D1102361 | 1 | D1102365 | 1 | D1201077 | 1 | D1201076 | 1 | aLigo secondary mirror mount tip-tilt sub assem. |
| D1102361 | 1 | D1102365 | 1 | D1201077 | 1 | D1201078 | 1 | ball for the secondary mount screw |
| D1102361 | 1 | D1102365 | 1 |  | 1 | 94125K577 | 5 | COMPRESSION SPRING, 7.1 MM OD, 0.8 MM WD, 15.5 MM L, 2.52 N/MM |
| D1102361 | 1 | D1102365 | 1 |  | 1 | 9716K36 | 2 | DISC SPRING .174" ID x .322" OD x .0050" THICK 1.50 Lb @ .034" DEFLECTION |
| D1102361 | 1 | D1102365 | 1 |  | 1 | 98017A625 | 4 | #8 FLAT WASHER, .17 ID x .38 OD x .032 THK, NAS 1149-CN832R |
| D1102361 | 1 | D1102365 | 1 |  | 1 | 99040A416 | 3 | SHIM WASHER .188 ID, .250 OD, .016 TH |
| D1102361 | 1 | D1102365 | 1 |  | 1 | BU-402-N | 3 | BUTTON HEAD SOCKET SCREW 4-40 x .13 |
| D1102361 | 1 | D1102365 | 1 |  | 1 | C-2006-N | 1 | SCREW, SOCKET HEAD CAP, #1/4-20 x .38 LG |
| D1102361 | 1 | D1102365 | 1 |  | 1 | C-2406-N | 2 | SCREW, SOCKET HEAD CAP, #10-24 x .38 LG |
| D1102361 | 1 | D1102365 | 1 |  | 1 | C-406-N | 2 | SCREW, SOCKET HEAD CAP, #4-40 x 0.38 LONG |
| D1102361 | 1 | D1102365 | 1 |  | 1 | C-408 | 3 | SCREW, SOCKET HEAD CAP, #4-40 x .50 LONG, VENTED |
| D1102361 | 1 | D1102365 | 1 |  | 1 | C-804-N | 2 | SCREW, SOCKET HEAD CAP, #8-32 x .25 LONG |
| D1102361 | 1 | D1102365 | 1 |  | 1 | D1001813 | 1 | aLIGO TMS THREAT MT. SEC. FOCUS |
| D1102361 | 1 | D1102365 | 1 |  | 1 | D1001834 | 1 | aLIGO TMS HOLDER SEC. FOCUS |
| D1102361 | 1 | D1102365 | 1 |  | 1 | D1001973 | 1 | aLIGO TMS FOCUS SCREW |
| D1102361 | 1 | D1102365 | 1 |  | 1 | D1200321 | 1 | aLIGO, AOS, TMS TELESCOPE, SECONDARY MIRROR, MT. BRACKET |
| D1102361 | 1 | D1102365 | 1 |  | 1 | D1200579 | 1 | aLIGO TMS FOCUS SLIDE LOCK |
| D1102361 | 1 | D1102365 | 3 |  | 1 | WFV-04 | 2 | WASHER, VENTED, NO. 4 |
| D1102361 | 1 | D1102365 | 1 |  | 1 | WFV-08 | 2 | #8 FLAT WASHER, .169 ID x .304 OD x .032 THK,VENTED |
| D1102361 | 1 | D1200243 | 1 |  | 1 | 70466S | 3 | COMPRESSION SPRING .210 O.D., .018 W.D., .63 L, 3.4#/IN |
| D1102361 | 1 | D1200243 | 1 |  | 1 | 94035A999 | 3 | PRECISION SKT SHLDR SCRW, 5/32 OD, 5/8 L, #6-32 x 3/16 THD |
| D1102361 | 1 | D1200243 | 1 |  | 1 | 94355A614 | 1 | SKT SET SCRW, FLAT PT, 1/4-28 x 1/2 |
| D1102361 | 1 | D1200243 | 1 |  | 1 | 9529K44 | 2 | 440C SS BALL 7/32" DIAMETER, GRADE 24 |
| D1102361 | 1 | D1200243 | 1 |  | 1 | 99934A330 | 1 | SKT SET SCRW, 6-32 x 1/8 L, SILVER TIP |
| D1102361 | 1 | D1200243 | 1 |  | 1 | C-2014-N | 6 | SCREW, SOCKET HEAD CAP, 1/4-20 UNC-2A x 0.875 LONG |
| D1102361 | 1 | D1200243 | 1 |  | 1 | C-2416-N | 3 | SCREW, SOCKET HEAD CAP, 10-24 x 1.0 LONG |
| D1102361 | 1 | D1200243 | 1 |  | 1 | D1002458 | 2 | aLIGO AOS F2 MIRROR SIDE CLAMP LOWER |
| D1102361 | 1 | D1200243 | 1 |  | 1 | D1002459 | 1 | aLIGO AOS F2 MIRROR SIDE CLAMP UPPER |
| D1102361 | 1 | D1200243 | 1 |  | 1 | D1002465 | 3 | aLIGO AOS TRANSMON F2 MIRROR CLAMP |
| D1102361 | 1 | D1200243 | 1 |  | 1 | D1102334 | 1 | aLIGO TMS TELESCOPE SECOND FOLD MIRROR |
| D1102361 | 1 | D1200243 | 1 |  | 1 | D1200249 | 1 | aLIGO TMS F2 MIRROR BASE |
| D1102361 | 1 | D1200243 | 1 |  | 1 | - | 3 | SISKIYOU 1" LEFT HAND MIRROR MOUNT |
| D1102361 | 1 | D1200243 | 1 |  | 1 | WFV-25 | 6 | VENTED FLAT WASHER .255 ID, .468 OD, .032 TH |
| D1102361 | 1 | D1200244 | 1 |  | 1 | 94035A999 | 3 | PRECISION SKT SHLDR SCRW, 5/32 OD, 5/8 L, #6-32 x 3/16 THD |
| D1102361 | 1 | D1200244 | 1 |  | 1 | 94355A614 | 1 | SKT SET SCREW, FLAT PT, 1/4-28 x 1/2 |
| D1102361 | 1 | D1200244 | 1 |  | 1 | 9435K420 | 3 | COMPRESSION SPRING .240 OD, .032 WD, .63 L, 23.32#/IN |
| D1102361 | 1 | D1200244 | 1 |  | 1 | 9529K44 | 2 | 440C SS BALL 7/32" DIAMETER, GRADE 24 |
| D1102361 | 1 | D1200244 | 1 |  | 1 | 99934A330 | 1 | SKT SET SCRW, 6-32 x 1/8 L, SILVER TIP |
| D1102361 | 1 | D1200244 | 1 |  | 1 | C-2014-N | 6 | SCREW, SOCKET HEAD CAP, 1/4-20 UNC-2A x 0.875 LONG |
| D1102361 | 1 | D1200244 | 1 |  | 1 | C-2416-N | 3 | SCREW, SOCKET HEAD CAP, 10-24 x 1.0 LONG |
| D1102361 | 1 | D1200244 | 1 |  | 1 | D1001136 | 3 | aLIGO F1 MIRROR CLAMP |
| D1102361 | 1 | D1200244 | 1 |  | 1 | D1002336 | 2 | aLIGO ETM F1 MIRROR SIDE CLAMP LOWER |
| D1102361 | 1 | D1200244 | 1 |  | 1 | D1002337 | 1 | aLIGO ETM F1 MIRROR SIDE CLAMP UPPER |
| D1102361 | 1 | D1200244 | 1 |  | 1 | D1102335 | 1 | aLIGO TMS TELESCOPE FIRST FOLD MIRROR |
| D1102361 | 1 | D1200244 | 1 |  | 1 | D1200245 | 1 | aLIGO TMS F1 MIRROR BASE |
| D1102361 | 1 | D1200244 | 1 |  | 1 | - | 3 | SISKIYOU 1" LEFT HAND MIRROR MOUNT |
| D1102361 | 1 | D1200244 | 1 |  | 1 | WFV-25 | 6 | VENTED FLAT WASHER .255 ID, .468 OD, .032 TH |
| D1102361 | 1 |  | 1 |  | 1 | 90585A628 | 6 | FLAT HEAD SOCKET SCREW, 3/8-16 x 1-1/2 |
| D1102361 | 1 |  | 1 |  | 1 | 92510A780 | 1 | SPACER ALUMINUM, .252 ID, 5/8 OD, 5/8 L |
| D1102361 | 1 |  | 1 |  | 1 | 9263K696 | 12 | O-RING, 2.5MM WIDTH, 7MM ID APPROVED VITON ELASTOMER, SHORE A 70-75 |
| D1102361 | 1 |  | 1 |  | 1 | 94355A614 | 1 | SKT SET SCRW, FLAT PT, 1/4-28 x 1/2 |
| D1102361 | 1 |  | 1 |  | 1 | 94518A418 | 2 | FLAT HEAD SOCKET SCREW, 10-24 x 3/4 |
| D1102361 | 1 |  | 1 |  | 1 | 9713K61 | 24 | DISC SPRING .255" ID, .500" OD, .018" THICK 47Lb @ .008" DEFLECTION |
| D1102361 | 1 |  | 1 |  | 1 | 99934A330 | 1 | SKT SET SCREW, 6-32 x 1/8 L, SILVER TIP |
| D1102361 | 1 |  | 1 |  | 1 | 99934A330 | 2 | SKT SET SCREW, 6-32 x 1/8 L, SILVER TIP |
| D1102361 | 1 |  | 1 |  | 1 | AS568A-106 | 10 | O-RING, 3/32" WIDTH, 3/16" ID, APPROVED VITON ELASTOMER, SHORE A 70-75 |
| D1102361 | 1 |  | 1 |  | 1 | C -816-N | 2 | SCREW, SOCKET HEAD CAP, #8-32 UNC-2A x 1 long |
| D1102361 | 1 |  | 1 |  | 1 | C -820-N | 2 | SCREW, SOCKET HEAD CAP, #8-32 UNC-2A x 1.25 long |
| D1102361 | 1 |  | 1 |  | 1 | C-2020-N | 2 | SCREW, SOCKET HEAD CAP, 1/4-20 UNC-2A x 1.25 long |
| D1102361 | 1 |  | 1 |  | 1 | C-2024-N | 3 | SCREW, SOCKET HEAD CAP, 1/4-20 UNC-2A x 1.5 long |
| D1102361 | 1 |  | 1 |  | 1 | C-2028-N | 1 | SCREW, SOCKET HEAD CAP, 1/4-20 UNC-2A x 1.75 long |
| D1102361 | 1 |  | 1 |  | 1 | C-2412-N | 2 | SCREW, SOCKET HEAD CAP, 10-24 x .75 LONG |
| D1102361 | 1 |  | 1 |  | 1 | D0902829 | 1 | aLIGO, AOS, TMS, ALIGNMENT TELESCOPE, MT. ADAPTER |
| D1102361 | 1 |  | 1 |  | 1 | D1000075 | 1 | aLIGO PRIMARY PARABOLIC MIRROR |
| D1102361 | 1 |  | 1 |  | 1 | D1001108 | 6 | aLIGO ETM TELESCOPE TUBE LARGE WASHER |
| D1102361 | 1 |  | 1 |  | 1 | D1001241 | 1 | aLIGO ETM TELESCOPE BEAM DUMP GLASS |
| D1102361 | 1 |  | 1 |  | 1 | D1002728 | 2 | aLIGO TMS PRIMARY MIRROR SIDE CLAMP LOWER |
| D1102361 | 1 |  | 1 |  | 1 | D1002731 | 1 | aLIGO TMS PRIMARY MIRROR SIDE CLAMP UPPER |
| D1102361 | 1 |  | 1 |  | 1 | D1002732 | 3 | aLIGO TMS PRIMARY MIRROR CLAMP |
| D1102361 | 1 |  | 1 |  | 1 | D1101688 | 1 | aLIGO TMS TELE ENTRY END PLATE |
| D1102361 | 1 |  | 1 |  | 1 | D1101727 | 1 | aLIGO TMS TELE PRIMARY 3 TUBE END PLATE |
| D1102361 | 1 |  | 1 |  | 1 | D1102297 | 1 | aLIGO TRANSMON TELESCOPE SSTL PLATE |
| D1102361 | 1 |  | 1 |  | 1 | D1200527 | 1 | aLIGO, AOS, TMS, ALIGNMENT TELESCOPE, CENTERING SLEEVE |
| D1102361 | 1 |  | 1 |  | 1 | D-271 | 1 | AUTOCOLLIMATOR |
| D1102361 | 1 |  | 1 |  | 1 | FA-2014-N | 13 | FLAT HEAD SOCKET SCREW, 1/4-20 x 7/8 |
| D1102361 | 1 |  | 1 |  | 1 | WF-08 | 16 | WASHER, FLAT, NO. 8 |
| D1102361 | 1 |  | 1 |  | 1 | WF-25 | 28 | .255 ID, .468 OD, .032 TH |

VIEW D1102361 SHEET 4 OF 7,  THE STARTING ASSEMBLY ORIENTATION OF THE TELESCOPE.

NOTE: END PLATES ARE INVERTED ON A FLAT SURFACE TO ALIGN WHAT WILL BECOME THE TOP SURFACES PARALLEL

### Assembly of D1000243-v3 aLIGO_ETM_TELE_TUBE_AND_INSERTS_ASSEMBLY

### Assembly of D1200244: aLIGO TMS F1 Mirror Assembly

### Assembly of D1102365: aLIGO TMS Telescope Secondary Mirror Assembly

#### Assembly of D1201077-aLigo Secondary Mirror Mount Tip-tilt Sub-assembly

#### Assembly of End Telescope Plates and Spacer Tubes

One person holds the bottoms of the End Plate assembly, with invar support tubes between loosely attached between them, against the table top, while a second person tightens the flat-head cap screws and fastens the tubes to the end plates with a torque value of 240 in-lbs using three iterations.( Alcohol on screws and heads to Torque )

Next, the TMS Telescope Bench Support Assembly D1100933 is attached to the bottoms of each End Plate.

Note: VIEW D1102361 SHEET 4 OF 7,  THE STARTING ASSEMBLY ORIENTATION OF THE TELESCOPE.

Note: END PLATES ARE INVERTED ON A FLAT SURFACE TO ALIGN AT START, WHAT WILL BECOME THE TOP SURFACES PARALLEL

#### Assembly of FM2

Bench Support Assembly

Alignment Support Bridge

Fold Mirror 2

Figure 14: FM2 assembly

#### Assembly of Primary Mirror

Primary Mirror

Figure 15: Primary Mirror Assembly

#### Assembly of FM2 and Secondary Mirror

Fold Mirror 1

2ndary Mirror

Figure 16: FM1 and Secondary Mirror Assembly

### Initial Alignment of TMS Telescope

The assembled TMS Telescope, which weighs approximately 58 lb, will be attached to the TMS Tele Alignment Support Bridge Assy D1100460, and then lifted 1/8 in above the table by turning the support screws on the Bridge Assy for initial alignment, as shown in Figure 17.

Remove the D1100933 stabilization bars during the optical alignment procedure.

TMS Telescope Alignment Support Bridge Assy

TMS Telescope Frame Assy

Figure 17: TMS Telescope Mounted to TMS Tele Alignment Support Bridge Assy D1100460 for Initial Alignment

VIEW D1102361 SHEET 4 OF 7,  THE STARTING ASSEMBLY ORIENTATION OF THE TELESCOPE.

NOTE: END PLATES ARE INVERTED ON A FLAT SURFACE TO ALIGN WHAT WILL BECOME THE TOP SURFACES PARALLEL

Figure 18: Completed TMS Telescope Assembly, Hanging from Alignment Support Bridge for Initial Alignment

### T1100603: TMS Telescope Alignment Procedure

The TMS Telescope alignment procedure is described in T1100603: TMS Telescope Alignment Procedure.

### Mounting the ISC Table Assembly to the TMS Telescope

After the optical alignment of the TMS Telescope on the Support Bridge, re-attach the D1100933 stabilization bars. Turn the support screws on the TMS Tele Alignment Support Bridge and lower the TMS Telescope to rest on the assembly table.

Remove the TMS Tele Alignment Support Bridge Assy to storage, as it will no longer be needed during the subsequent procedures.

Four people will lift the ISC Table Assembly and place it on top of the TMS Telescope. The ISC Table will then be attached to the TMS Telescope using the attachment hardware, as shown in

Figure 19: ISC Table Assembly being mounted to Top of the TMS Telescope

### Balancing the TMS Telescope Frame Assembly /ISC Table Assembly

#### Assemble the TMS Tele-opt Table Balance Bridge Assembly, D1100649

The assembled aLIGO TMS Tele-Opt Table Balance Bridge Assembly is shown in Figure 20

Note: Use this Balance Bridge with the assembled Telescope and ISC table to Balance it, prior to using the  support bridge D1100531 for Telescope to ISC table alignment.

Figure 20: TMS Tele-opt Table Balance Bridge Connection-Interface

#### Prepare the Pin-in-Socket Balance Point

Insert the TMS Tele-Opt Table Balance Bolt into the ISC Table Extension, as shown in Figure 21.

Figure 21: Balance Bolt Inserted into ISC Table Extension

Set the height of the pointed pivot screw to the nominal value, so that the pin and the socket will contact when the Balance Bridge is placed below the ISC Optics Table.

Figure 22: TMS Tele-Opt Table Balance Bolt

Figure 23: Detail of Balance Pin and Socket Positions

Figure 24: Detail of Pivot Pin Mating with Pivot Socket

#### Support the ISC Table / TMS Telescope Assembly with the Balance Bridge

Place the Balance Bridge on the assembly table underneath the ISC Table, by straddling the lower invar tube of the Telescope, as shown in Figure 25. Position the Balance Bridge so that the Pivot Pin on the Bridge and the Socket Bolt on the ISC Table are aligned.

Raise the leveling pad feet of the Balance Bridge iteratively, until the Telescope is lifted approximately ¼ in above the assembly table. Clamp the feet against the assembly table with dog clamps.

Remove the stabilizing bars on the bottoms of the Telescope End Plates.

#### Balance the ISC Table / TMS Telescope Assembly, See Table Assembly for Nominal Position of Trim Masses. First Article 80kg., Production 82kg.

Move the adjustable balance weights on the underside of the ISC Table until the ISC Table is balanced in a horizontal plane. Note the nominal location of the trim masses per D1102291

Figure 25: Balance Bridge supporting the ISC Table/ TMS Telescope Assembly

### Suspending the TMS Telescope Frame Assembly/ISC Table Assembly from the TMS Optical Table Support Bridge Assy D1100531

The Optical Alignment between the ISC Table and the TMS Telescope will be accomplished while the TMS Telescope Frame Assembly/ISC Table Assembly is suspended.  To do this, the TMS Telescope Frame Assembly/ISC Table Assembly will be suspended with straps from the TMS Optical Table Support Bridge Assy to simulate the suspension wires.

Place the TMS Optical Table Support Bridge Assy D1100531 on the work table to straddle the combined TMS Telescope Frame Assembly /ISC Table Assembly structure, as shown in the figures below. Table clamps, shown in Figure 26, are fastened to immobilize the suspended Telescope Assembly. Note: Telescope /Optical table is a complete as installed in Chamber

Table Clamps

Figure 26: TMS Optical Table Support Bridge Assy D1100531

Work Table

Optical Table Support Bridge

ISC Table Assembly

TMS Telescope Frame Assembly

Figure 27: TMS Optical Table Support Bridge Assy Straddling the Combined TMS Telescope/ISC Table Assembly Structure

Attach the suspension straps to the mounting holes of the ISC Table Assembly structure and adjust the leveling pad feet equally, in increments, to raise the Telescope assembly approximately 1/8 in above the assembly table, as shown below.

Optical Table Support Bridge

Suspension Straps

ISC Table Assembly

Figure 28: TMS Telescope/ISC Table Assembly Suspended by Straps to the ISC Table Assembly structure

### Mounting the AOS ETM Telescope Upper Suspension Frame and Top Mass to the Tele-opt Table (Bosch Frame)

Level the Tele-opt Table Frame Plate D1100807 on the Bosch Frame by adjusting the leveling pad feet, as shown in Figure 29.

Tele-opt Table Frame Plate

Leveling Pad Foot

Figure 29: TMS Tele-opt Table on Bosch Frame

A Genie lift, with special lifting forks attached (see Figure 30) and with the legs splayed appropriately to avoid tipping over of the load, will lift the Upper Suspension Frame and Top Mass (see Figure 31) up to the underside of the Tele-opt Table Frame Plate D1100807 on the Bosch Frame, as shown in Figure 32.

Figure 30: Modified Genie Lift Forks, D1100886

The combined Upper Suspension Frame and Top Mass weight is 273 lbs (124kg).

Figure 31: Upper Suspension Frame and Top Mass

Tele-opt Table Frame Plate on Bosch Frame

Genie Lift

Upper Suspension Frame and Top Mass

Figure 32: Genie Lift supporting Upper Suspension Frame and Top Mass

### TMS Test Mass Assembly, D1002097

A Test Mass will be used as a payload for the preliminary adjustment of the Top Mass control system.

#### Assembly of the Test Mass

The assembled TMS Test Mass is shown in Figure 33. Weighs: 80 kg First Article, 82kg Production. The various attached weights can be repositioned to change the location of COM. The suspension wires from the Upper Mass attach to the D1002240 aLIGO TMS Tele Wire Clamp Base, as shown in Detail A.

Figure 33: TMS Test Mass Assembly

#### Balancing the Test Mass

The Test Mass will be lifted by four people, while a fifth person places the

D1201330 aLIGO TMS Test Mass Balance Ball and spacer under the Test Mass Plate to mate with the locating hole, as shown in Figure 34.

The movable balance weights will be positioned so that the Test Mass is balanced in a level horizontal plane.

Figure 34: Placement of Balance Ball beneath the Test Mass

#### Installing the Test Mass: 80kg First Article (WBSC6) Production 82kg

The Genie lift adjustable forks will be positioned beneath the Test Mass, and the Test Mass will be transported under the Bosch Frame and lifted to a height that enables the hanging suspension wires from the mounted Top Mass to be connected to the Test Mass. The suspension wires will then be secured to the ISC Table mounting locations, as shown in the figures below.

Figure 35: Test Mass Suspended from Top Mass

#### Adjust the TMS Control System with Suspended Test Mass

#### Removing the Test Mass

After the Top Mass OSEM Controls have been adjusted, the suspension wires will be disconnected from the Test Mass, and it will be lowered using the Genie lift, as described in section 5.8.11.3, and transported to a storage area.

### Mounting the TMS Telescope Frame Assembly/ISC Table Assembly to the Top Mass

A modified Genie lift, with adjustable forks and with the D1100841 TMS Tele-opt Table Install Tool attached, is shown in Figure 36 and Figure 37. See: D1100908 For application w/ Genie Lift.

Note: Use D1200355 Height Spacers, under #2, to for additional Working Room between Underside of Table and Tool Plate #1

Figure 36: TMS Tele-opt Table Install Tool Assy

Figure 37: TMS Tele-opt Table Tool Assy-Genie Fork Connection

Lift the TMS Telescope/ISC Table Assembly from the assembly table using the Genie lift and transport it to the Bosch Frame (just as was done to transport and lift the Test Mass). Lift it to the level of the suspension wires hanging from the Top Mass, and connect the Top Mass suspension wires to the TMS Telescope/ISC Table--in the middle hole of the attachment bracket)  and suspend it, as shown in Figure 38.

Final adjustment of the TMS control system will be done while the TMS Assembly is mounted to the Bosch Frame.

NOTE: SEE SECTION 9.0 FOR BEAM DISTANCE FROM CARTRIDGE TABLE

Figure 38: TMS Assembly attached to Tele-opt Table Frame Plate on Bosch Frame

TMS Assembly

Genie Lift

### Adjustment of TMS Control System will Suspended TMS Telescope/ISC Table Assembly

TMS Installation Procedure

TMS Assembly

Genie Lift

## Transporting and Mounting the TMS System to the ISI Cartridge

The TMS System will be transported and installed on the ISI Cartridge in two steps: 1) First install the Top Mass, and 2) suspend the TMS Telescope/ISC Table Assembly from the Top Mass. This procedure requires reversing the procedure described in section 5.8.12 to detach the suspended TMS Telescope/ISC Table Assembly from the Upper Suspension Frame and Top Mass in the Bosch Frame.

Whenever the ISC Table Assembly is moved or transported, the Protective Covers (D1102411, D1102349, D1102329, D1102328) shown in Figure 39 should be in place. Ref D1102327 assem.

Figure 39: Transmon ISC Protective Cover

### Remove TMS Telescope/ISC Table Assembly from the Upper Suspension Frame and Top Mass

1. Release the tension on the suspension wired by raising the Genie lift against the Safety Support Assembly, shown in Figure 2, until the blade springs rest against their stops.

2. Unscrew the 1/4-20 SHCS which hold the wire clamp onto the short cross bar (two per

wire clamp). Although not difficult, they are awkward to reach, especially at the side of the secondary mirror!! Have one person underneath the Tele working on the screws, while a second person keeps the suspension wire under tension from above. Because the suspension wire is quite thick you can control the orientation of the wire clamp. This will greatly help the person underneath unscrewing the SHCS. Note: someone with small hands has an advantage.

2. Attach the TMS Telescope Bench Support Assembly D1100933 to the bottoms of each End Plate.

3. Once all wires are removed, the Genie will be rolled away, and the TMS Telescope/ISC Table Assembly will be placed temporarily on a work table.

### Prepare Upper Suspension Frame and Top Mass

1. Reposition the four electrical cables such that they do not cross underneath the Upper Suspension Frame and Top Mass.

2. Install the 5-axis SUS platform onto the Genie forks. The total nominal weight of the cube + cage + 5-axis SUS platform is approximately 275 lbs. Rotate the platform such that it can rotate by at least 180 degrees in both directions (CW and CCW), along the vertical axis once the Upper Suspension Frame and Top Mass is mounted on the platform.

3. Bring the Genie underneath the Upper Suspension Frame and Top Mass, which is still hanging from the Bosch frame. One point to check is the location of the Upper Suspension Frame and Top Mass on the 5-axis platform. You want to locate the 'front' of the Upper Suspension Frame and Top Mass as close to the edge of the platform as possible, as this will aid during the mounting on the ISI platform next to the ETM, which is fully assembled. The clearance between the TMS and the ETM is approximately 4 inch.

4. Mount the Upper Suspension Frame and Top Mass to the platform and unclamp the cube from the Bosch frame.

5. Lower the Upper Suspension Frame and Top Mass to knee height and cover with the clean-room cover.

6. The Upper Suspension Frame and Top Mass is now ready for transport to the ISI.

### Attach Upper Suspension Frame and Top Mass to Cartridge

The Suspension Frame will be placed on the ISI Optical Table at the local coordinates specified in D0900419 AdvLIGO SUS BSC6-H2, XYZ Local CS for TMS ETM Tel Assy. Refer to E1101131  BSC6 H2, Test Stand, Quad, TMs, Tooling Clearances, aLIGO

Protective covers for the ETM must be in place.

Note: D1300448 TMS Cartridge Swing Stop Tool is used to prevent TMS and ETM from Bumping

Use the Genie lift to transport and lift the Upper Suspension Frame and Top Mass up to the ISI table on the Cartridge placed on top of the Test Stand, as shown in Figure 40. This portion of the TMS Assembly weighs 273 lbs (124kg). Place the Frame against the preset location stops on the ISI Optical Table, as described in D1101260.

Upper Suspension Frame and Top Mass

Test Stand

Genie Lift

ISI Cartridge Assembly

Figure 40: Upper Frame and Top Mass Attached to Cartridge Assembly

### Attach TMS Telescope/ISC Table Assembly to the Top Mass Suspension Wires

The Transmon ISC Protective cover should be in place.

Use the Genie lift to transport and lift the TMS Telescope/ISC Table Assembly up to the suspension wires hanging from the Top Mass, and attach to the suspension wires following the procedure described in section 5.8.12. This portion of the TMS Assembly weighs 176 lbs (80 kg).

## Final Installation inside BSC

### Preparing the TMS System for Cartridge Flight into the BSC

After alignment of the TMS Assembly and before transporting or installing the TMS Assembly to the Cartridge, install the Transmon ISC Protective cover.

The Transportation-Installation Restraints must also be in place to protect the TMS from damage by securing its suspended components per the Transportation-Installation Restraint Procedure, E1100841. Refer to the restraint tooling assemblies D1001891, D1101506, D1101559, D1101565 and related hardware to the TMS Assembly, with reference to Tooling Layout D1101307 and Assembly Drawing D1000549.

D1300248 Swing Stop Tool is used to Prevent the TMS and ETM from Bumping,  NOTE: see p. 160

Figure 41: Restraint Tooling Layout (D1101307), when

Transporting the Cartridge

#### TMS Transportation-Installation Restraint Procedure, E1100841

##### Adjust & Install Eye-Bolts

Adjust the 2 Eye-Bolts of the TMS Assembly according to Tooling Layout D1101307, view D-D. Install the additional tooling eye-bolts, nuts & washers according to Tooling Layout D1101307 (Parts List items 1, 2, & 3).

##### Install Vertical Safety Chains

Install the 4 Vertical Safety Chains D1001891 according to Tooling Layout D1101307. Tighten the nuts on the eye-bolts to raise the TMS Telescope/ISC Table Assembly until the Top Mass blade springs rest against their stops and the suspension wires become slack.

##### Lift Telescope

Lift the suspended Telescope .35" to .5" according to Tooling Layout D1101307, using the 4 Vertical Safety Chains.

##### Install Turnbuckles

Install Turnbuckle assemblies D1101506 (2), D1101559 (2), & D1101565 (4), adjusting them snugly according to Tooling Layout D1101307 to secure the suspended Telescope.

##### Secure Upper Mass

Secure the upper mass using the Earthquake Stop Screws and jam nuts of the TMS Assembly (see Assembly Drawing D1000549, Parts List items 21 & 22). Adjust the lower 4 Earthquake Stop Screws to touch the bottom of the Upper Mass. Adjust the upper 4 Earthquake Stop Screws to clamp down on the top of the Upper Mass to fully restrain it. Tighten all 8 jam nuts.

### Cartridge Flight into the BSC

Fly the Cartridge with the TMS into the BSC.

Figure 42: TMS Assembly being Installed in BSC

### TMS Seismic Safety Stop Structure Attachment to TMS

Assemble the TMS Seismic Safety Stop Structure D1001781, as shown in Figure 43.

After the TMS is placed inside the BSC, attach the TMS Seismic Safety Stop Structure D1001781

To the Chamber wall brackets as shown in Figure 45, Figure 44, and Figure 45.

The Seismic Safety Stop Structure supports a rod that extends thru the eye-bolts of the TMS Telescope Safety Support Beam Assembly located at the end edge of the optical table; it will restrain the TMS from violent motion. In normal operation, the rod passes clearly through the inside diameter of the eye-bolts without making contact with the suspended TMS assembly; this arrangement limits travel motion but will not interfere with the normal operation of the suspended TMS.

Figure 43: TMS Seismic Safety Stop Structure (D1001781)

Figure 44 Earthquake Restraint

Figure 45: Seismic Stop Installation Procedure

## Pull Back Procedure for Maintenance of ETM

The TMS Telescope/ISC Table Assembly will be supported by the Vertical Safety Chains and pulled back during maintenance of the ETM Quad assembly, as shown in Figure 46, and held in place by the Seismic Safety Stop. The procedure is described below.

Disconnected Vertical Safety Wire Assembly

Vertical Seismic Safety Stop Assembly

Figure 46: TMS Telescope Pulled Back to Service the ETM, Elevation View

Figure 47: TMS Telescope Pulled Back to Service the ETM, Top View

### Disconnect the Vertical Safety Wire Assembly

Remove the screws that connect the Vertical Safety Wire Assembly to the Telescope Frame. Be careful not to drop the captured washer between the bracket and the Telescope Frame, and the captured washer underneath the screw head. Keep the hardware for subsequent re-attachment after the pull back procedure is undone.

Captured Washer

Figure 48: Vertical Safety Wire Assembly Connected to the TMS Telescope Frame

### Disengage Earthquake Restraint from the TMS

#### Attach the Seismic Safety Stop Chain

The Seismic Safety Stop Chains are hanging permanently on the Seismic Safety Stop Structure inside the BSC. Unhook the threaded mechanical link and wrap the chain around the TMS Stop Rod, as shown in Figure 49.

Figure 49: Seismic Safety Stop Chain attached to TMS Stop Rod

#### Free the TMS Stop Rod

Unscrew the attachment bolt that fastens the TMS Stop Rod to the TMS End Bracket Rod. Be careful not to drop the captured washer. Keep the hardware for subsequent re-attachment after the pull back procedure is undone.

### Attach Vertical Safety Chains and Pull Back

Install the Vertical Safety Chains as described in section 6.2.1.1.2.

Slide the TMS Horizontal Restrain Link to the center of the TMS Stop Rod against the Two Piece Collar, as shown in Figure 50.

Figure 50: Positioning of Horizontal Restrain Link on TMS Stop Rod, as Seen from Below

Two people will grab the TMS Stop Rod and pull the ISC Optical Table/ Telescope Assembly toward the Earthquake Stop Cross Bar. A third person will apply the bolt and nut from underneath to attach the Restrain Link to the Rod End, as shown in Figure 51.

Nut and Bolt Assembly

Horizontal Restrain Link

Rod End

Figure 51: Detail of Horizontal Restrain Link attachment to Earthquake Stop Cross Bar, View from Back of TMS Assembly

# Appendix A: Useful Balancing and Alignment Data

# The Telescope and Optical Table Assembly First Article Assembly: Total mass 80 kg

# The Telescope and Optical Table Assembly Production Assembly: Total mass 82 kg

1. TMS system is Weighed and Massed Properly Using the Telescope and Table Drawing D1002460 to roughly locate added Cylinder Masses. (this has been changed due to new mass properites)

2. The Cylinder Masses are positioned for Balance per the Balance Bridge (page 129) and instruction text, note the trim pitch and roll mass locations per instructions.

3. The assembly is lowered to the table and then attached to the D1100531 Bridge to align the Telescope to the ISC table per outlined procedures.

# Appendix C: Related Documents

T1300639 Primary Mirror Installation Procedure

D060310 Advanced LIGO, SUS, QUAD N-PType Tablecloth

D060324 SUS, Quad N-Ptype Top Stage, Blade Cartridge

D060370 Picture Book

D0900419 AdvLIGO SUS BSC6-H2, XYZ Local CS for TMS ETM Tel Assy

D1000484: ISC Transmon Assembly

D1000549 Upper Intermediate Mass Top Assembly

D1001781 TMS Seismic Safety Stop Structure

D1001891: aLIGO TMS VERTICAL SEISMIC SAFETY STOP ASSEMBLY

D1002097 aLIGO_TMS_TEST_MASS_ASEM

D1002097 TMS Test Mass Assembly

D1100460: aLIGO TMS TELE ALIGNMENT SUPPORT BRIDGE ASSY

D1100531: aLIGO TMS OPTICAL TABLE SUPPORT BRIDGE

D1100613: aLIGO TMS TELESCOPE ASSEMBLY ALIGNMENT STABILIZATION TOOL

D1100649: aLIGO TMS TELE-OPT TABLE BALANCE BRIDGE ASSY

D1100807:  aLIGO TMS Setup Table Bosch-Frame Assembly

D1100827: VERTICAL SAFETY WIRE ASSEMBLY

D1100841: aLIGO TMS TELE-OPT TABLE INSTALL TOOL ASSY

D1100886: MODIFIED GENIE LIFT FORKS

D1100933: aLIGO TMS TELESCOPE BENCH SUPPORT ASSEMBLY

D1101064: aLIGO TMS FIRST ARTICLE INSTALLATION AND TOOLING

D1101130: aLIGO TMS TELESCOPE SAFETY SUPPORT BEAM ASSY

D1101163: aLIGO TMS Telescope SUS Wire Assembly

D1101166 TMS Upper Suspension Wire Assembly

D1101307 aLIGO TMS Telescope Safety Support Assemblies

D1101361: TMS TOOLING TELESCOPE I-P ALIGNMENT MIRROR ASSEMBLY

D1101487: TMS TOOLING TELESCOPE O-P ALIGNMENT MIRROR ASSEMBLY

D1101510: aLIGO TMS TELE ALIGNMENT 6in MIRROR MOUNT ASSY

D1101526 aLIGO TMS Mass Balanced Assembly

D1102291: aLIGO TMS ISC Table Structural Assembly

D1102361: aLIGO TMS Telescope Frame Assembly

D1200453 aLIGO TMS Upper Structure Weldment

E0900047 Contamination Control Plan

E1000006 Advanced LIGO Quad Suspension Metal-Build Assembly Procedure

E1100841 Transportation-Installation Restraint Procedure (integrated into this doc.)

E1101131 BSC6 H2, Test Stand, Quad, TMs, Tooling Clearances, aLIGO

E1200049 aLIGO Transmission Monitor Suspension Installation Plan

F0900052 Advanced LIGO Inventory Database Part Import

T1000674-v2 TMS wire assembly procedure

T1100304 TMS ASSEMBLY and ALIGNMENT DOCUMENTATION

T1100603: TMS Telescope Alignment Procedure

T1200388 TMS Wire Jig Assembly

D1101624 Install Tooling (reference for other sub-systems)

# TMS Swing Stop Tool Used on Cartridge Installation w/ETM

D1300248 TMS Swing Stop Tool ( Used to Prevent the TMS and ETM from Bumping)

See Page 146

used when TMS and ETM are mounted on the Cartridge. Note counter weight disc's are fixed from sliding off with long cap screws, nuts are not required.

This Drawing needs minor Update as of 6-18-13

# TMS BEAM CENTER LINE DISTANCE FROM CARTRIDGE TABLE

# TMS TOOLING USED DURING VARIOUS STAGES OF INSTALL ON BOSCH FRAME AND CARTRIDGE

Genie Transfer tool D1100908 and Optional D1200355 spacers.  Ref. section 4.6 /4.10

| TOOLING USED DURING VARIOUS STAGES on INSTALL | DCC Numbers | Notes: Refer to E1201012 overall TMS installation procedure. | Additional Notes |
| --- | --- | --- | --- |
| BOSCH FRAME |  |  |  |
| ON BOSCH FRAME (Suspended) |  | OUT - SWING STOP TOOL, - DCC D1300248 IN - FRONT SAFETY WIRE, - DCC D1100827 OUT - Stabilization tooling, - DCC D1100613 IN - CHAINS, DCC D1001891 OUT - TURNBUCKLES, - DCC D1101307 OUT -ISC Cover, DCC D1102327 | Chains - chains and eye disks should not contact telescope during functional tests. , see D1101307 |
| ON BOSCH FRAME (Working on telescope or Away i.e. overnight / weekend) |  | OUT - SWING STOP TOOL - DCC D1300248 IN - FRONT SAFETY WIRE - DCC D1100827 IN - Stabilization tooling, - DCC D1100613 IN - CHAINS. - DCC D1001891 OUT - TURNBUCKLES, - DCC D1101307 IN - ISC Cover,DCC D1102327, (WHEN NEEDED) |  |
| TEST STAND 'Cartridge' |  |  |  |
| ON TEST STAND 'Cartridge' (working on telescope) |  | IN - SWING STOP TOOL, - DCC D1300248 IN - FRONT SAFETY WIRE, - DCC D1100827 IN - Stabilization tooling, DCC D1100613 OUT - CHAINS, - DCC D1001891 OUT - TURNBUCKLES, - DCC D1101307 OUT - ISC Cover, - DCC D1102327 |  |
| ON TEST STAND (working on BOSEM adjustment / blade adjustment) |  | IN - SWING STOP TOOL, - DCC D1300248 IN - FRONT SAFETY WIRE, - DCC D1100827 IN - Stabilization tooling, - DCC D1100613 OUT - CHAINS, - DCC D1001891 OUT - TURNBUCKLES, - DCC D1101307 IN - ISC Cover,DCC D1102327- Used to protect optics when e.g. working BOSEMS - then removed |  |
| ON TEST STAND 'Cartridge' (suspended - testing BOSEMS or transfer function) |  | IN - SWING STOP TOOL - DCC D1300248 IN - FRONT SAFETY WIRE - DCC D1100827 OUT - CHAINS, - DCC D1001891 OUT - TURNBUCKLES, - DCC D1101307 OUT - Stabilization tooling, - DCC D1100613 OUT - ISC Cover, DCC D1102327 |  |
| ON TEST STAND 'Cartridge' (Away i.e. overnight / weekend) |  | IN - SWING STOP TOOL - DCC D1300248 IN - FRONT SAFETY WIRE, - DCC D1100827 OUT - CHAINS, - DCC D1001891 OUT - TURNBUCKLES, - DCC D1101307 OUT - Stabilization tooling (Item #6 in D1101064) OUT - ISC Cover, DCC D1102327 |  |
| FLIGHT |  |  |  |
| FLIGHT |  | OUT - SWING STOP TOOL - DCC D1300248 IN - FRONT SAFETY WIRE, - DCC D1100827 IN - CHAINS DCC D1001891 Per Installation Procedure IN - TURNBUCKLES, - DCC D1101307 OUT - Stabilization tooling, - DCC D1100613 IN - ISC Cover, DCC D1102327 | Chains - chain and eye disks should be adjusted to lift telescope, Per Installation Procedure |
| IN CHAMBER |  |  |  |
| IN CHAMBER (work e.g. aligning) |  | OUT - SWING STOP TOOL - DCC D1300248 IN - FRONT SAFETY WIRE, - DCC D1100827 OUT - CHAINS, - DCC D1001891 OUT - TURNBUCKLES, - DCC D1101307 IN - Stabilization tooling, - DCC D1100613 OUT - ISC Cover, DCC D1102327 (UNLESS NEEDED) |  |
| IN CHAMBER (doors closed) |  | IN - *NEW* Seismic Safety Stop D1001781 OUT - SWING STOP TOOL - DCC D1300248 IN - FRONT SAFETY WIRE, - DCC D1100827 OUT - CHAINS, - DCC D1001891 OUT - TURNBUCKLES, - DCC D1101307 OUT - Stabilization tooling, - DCC D1100613 OUT - ISC Cover, DCC D1102327 | D1001781 is installed in the BSC, there is a description of the installation procedure for this in E1201012, 6.2.3, and noted on drawing, also use gage D1101486 To set clamping torque by Gauge gap. |

# TMS ROLL AND PITCH ADJUSTER PROCEDURE

Using D1002460 page 2 and more importantly D1102291 pages 5 and 7 you will see (along with the text below) i) the starting positions of the pitch and roll devices ii) how to un-lock these devices iii) how to move them and iv) how to lock them (same as #2).

ROLL ADJUSTER

i) Starting position of roll adjustment mass (page 5 of D1102291) - Ref dimension in zone B8.

ii) and iv) Blow out in zone D3 shows un-lock and re-lock screws

iii) Detail E shows the mechanism to adjust - use Allen key

PITCH ADJUSTER

i) Starting position of roll adjustment mass (page 7 of D1102291) - Ref dimension in zone D3.

ii) and iv) in zone A4 and A5 you will see fasteners in slots. These should be loosened prior to adjusting. They can also  be moved to different slots as needed

iii) Detail F shows the mechanism to adjust - use Allen key