LIGO Laboratory / LIGO Scientific Collaboration

LIGO- E1201074-v6	LIGO	8/26/2013

Output Faraday Isolator Assembly and Alignment  Procedure

Mike Smith

Distribution of this document:

LIGO Scientific Collaboration

This is an internal working note

of the LIGO Laboratory.

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 159 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

http://www.ligo.caltech.edu/

Table of Contents

1	Introduction	7

2	OFI Assembly	7

2.1	Overview of the D0900136 OFI Assembly	7

2.2	D0900048 Faraday Isolator Damper Holder Assembly	10

2.3	D0900614 PRISM MOUNT ASSY_LH	11

2.3.1	Step 1	13

2.3.2	Step 2	14

2.3.3	Step 4	15

2.3.4	Step 5	16

2.3.5	Step 6	17

2.4	D0900615 PRISM MOUNT ASSY_RH	18

2.4.1	Step 5	20

2.4.2	Step 6	21

2.5	D1002364 Faraday Isolator Beam Dump Assy	22

2.6	D0900353 Half Wave Plate Holder Assy	24

2.7	D0900440 TFP Polarizer Plate Assy	26

2.8	D1001958 Wire Support Block Assy	28

2.9	D1001963 Output Alignment Fixture Assy	30

2.10	D1200197 Wedged Prism Holder Assy	32

2.10.1	D0900359 Wedged Prism Assy	34

2.11	Faraday Isolator Table Assy	35

2.11.1	Step 1	37

2.11.2	Step 2	39

2.11.3	Step 3	39

2.11.4	Step 4	41

2.11.5	Step 5	43

2.11.6	Step 6	45

2.11.7	Step 7	47

2.11.8	Step 8	48

2.11.9	Step 9	50

2.11.10	Step 10	51

2.11.11	Step 11	53

2.11.12	Step 12	53

2.11.13	Step 13	54

2.11.14	Step 14	55

2.12	D0900586 Faraday Isolator Upper Wire Assembly	56

2.12.1	OFI Wire Length	59

2.13	D0900170 Earthquake Crossbar Assy	65

2.14	D1002256 Earthquake Crossbar_In Assy	66

2.15	D0900579 Blade Guard Assy	67

2.16	D0900655 Structural Weldment Assy, OMC	68

3	OFI Balancing	70

3.1	Install Faraday Isolator Damper Holder Assembly	70

3.2	Install Blade Springs	72

3.3	Install Blade Guard Assembly	72

3.4	Install Upper Wire Assembly	72

3.4.1	Clamp Wire Clamp to Tip of Blade Spring	72

3.5	Install Table Assy	73

3.6	Balance Table	73

3.6.1	Adjust Balance Weights	73

3.6.2	Setting Gap Between Blade Guard Assembly and Spring Blade	88

3.6.3	Adjust Wire Lengths	88

3.7	Install Earthquake Crossbar Assy	92

4	OFI Optical Alignment	94

4.1	OFI Preliminary Optical Alignment	94

4.1.1	Equipment List	94

4.1.2	Alignment Setup	94

4.1.3	Optical Alignment Procedure	96

4.2	OFI Shack-Hartmann Wavefront Test	96

4.2.1	Equipment List	96

4.2.2	Alignment Setup	98

4.2.3	Double Pass OFI Test Procedure	101

4.2.4	Double Pass OFI Test Results	102

Table of Figures

Figure 1: Output Faraday Isolator Assembly	9

Figure 2: Prism Mount Assy_LH	12

Figure 3: Prism Mount Assy_RH	19

Figure 4: 3.5	D1002364 Faraday Isolator Beam Dump Assy	23

Figure 5: Half Wave Plate Holder Assy	25

Figure 6: 2.7	D0900440 TFP Polarizer Plate Assy	27

Figure 7: 3.8	D1001958 Wire Support Block Assy	29

Figure 8: Output Alignment Fixture Assy	31

Figure 9: 3.10	D1200197 Wedged Prism Holder Assy	33

Figure 10: Assembly of Faraday Isolator Table	36

Figure 11: Magnet Locations, with Opposite Pole Orientations	56

Figure 12: OFI Wire Length Fixture with Wire Loop Tool	60

Figure 13: Right End of Wire Length Fixture with Recess for Inserting the Small End of the Upper Wire Clamp	61

Figure 14: Upper Wire Loop Clamp Assembly	62

Figure 15: Bottom Wire Clamp Assembly	64

Figure 16: 2.16	D0900655 Structural Weldment Assy, OMC	69

Figure 17: Mounting of the Damper Holder Assembly	70

Figure 18: Nominal Height of Damper Holder Assy	71

Figure 19: Split Clamp Clamp D1300321	72

Figure 24: Light Optical Table_D0900015-v4: Off-set Positions for Balance Weights CW1 and CW2	75

Figure 25: Heavy Optical Table_D0900015-v5: Off-set Positions for Balance Weights CW1 and CW2	77

Figure 26: Retract Spring Restraints 0.25in so that Blades Move Freely in Vertical Direction	88

Figure 27: Guitar Tuner Placed at Base of Blade Spring	89

Figure 28: Wire Tension Adjustment	90

Figure 29: Table Height Gauge, D1201449	91

Figure 30: Nominal Height of Optics Table above the Structure Base	92

Figure 31: Earthquake Stop Posts Attached to Four Corners of the Optics Plate	93

Figure 32: Optical Layout for OFI Alignment	95

Figure 33: Optical Layout for SH Test	98

Figure 34: Contour of the Input Beam on Shack-Hartmann Sensor	99

Figure 35: Reference Power Read by the Shack-Hartman Lens Array	100

Figure 36: Reference Wavefront	101

Figure 37: Typical OFI Double-pass Wavefront with Tilt Removed	102

Figure 38: Typical OFI Double-pass Zernike Test Results	103

Table of Tables

Table 1: Bill of Materials, D0900136 Faraday Isolator Assembly	8

Table 2: Bill of Materials, D0900614 Prism Mount Assy LH	11

Table 3: Bill of Materials, D0900615 Prism Mount Assy RH	18

Table 4: Bill of Materials, D1002364 Faraday Isolator Beam Dump Assy	22

Table 5: Bill of Materials, D0900353 Half Wave Plate Holder Assy	24

Table 6: Bill of Materials, D0900440 TFP Polarizer Plate Assy	26

Table 7: Bill of Materials, D1001958 Wire Support Block Assy	28

Table 8: Bill of Materials, D1001963 Output Alignment Fixture Assy	30

Table 9: Bill of Materials, D1200197 Wedged Prism Holder Assy	32

Table 10: Bill of Materials, D0900623 Faraday Isolator Table Assy	35

Table 11: Bill of Materials, D0900586 Faraday Isolator Upper Wire Assembly	56

Table 12: Bill of Material, D0900655 Structural Weldment Assy, OMC	68

Table 13: Light Optical Table_D0900015-v4: Balance Weights and Positions	74

Table 14: Heavy Optical Table_D0900015-v5: Balance Weights and Positions	76

CHANGE LOG

| Date, version | Summary of Changes |
| --- | --- |
| V3 2/1/13 | - Added Change Log - Changed to 4 magnets and added magnet orientation - Revised counter weights - Revised Earthquake Rods and Transport Locking Shim |
| V4 4/9/13 | - Added SH Test Procedure |
| V5 5/24/13 | - Added part number for earthquake stop post - 3.6.2 Setting Gap Between Blade Guard Assembly and Spring Blade - Added Split Clamp Clamp D1300321 to secure SUS wire to Blade Tip |
| V6 8/26/13 | - Added balance weight locations for heavy optical table D0900015-v5 |

# Introduction

This document details the assembly and alignment of the AOS Faraday Isolator (OFI) Assembly, D0900136. The OFI is one element of the Stray Light Control (SLC) subsystem.

The OFI will be assembled and aligned in a Class A environment that meets the following Clean room standards:

For a clean assembly all LIGO standards should be followed, as presented in the latest version of the LIGO Contamination Control Plan (E0900047). Clean room garb including UHV gloves should be worn when working with parts.

All tools that come in contact with assembly should be cleaned to class B standards.

Assembly will be done under a portable clean room. Any time a part of the assembly is not covered by the portable clean room or is not being actively worked on it should be covered with appropriate clean covers, e.g. C3 polyester or equivalent.

Alignment will be done on an optical bench inside a CLASS 100 enclosure.

All parts that will be included in the final assembly must be cleaned to LIGO standards, Class A. The list of parts to be Class A-cleaned includes screws, washers, inserts, and assorted other hardware. All tooling and other parts that are not included in the final assembly, but that contact Class A parts during assembly must be cleaned to LIGO standards, Class B.

This procedure must be read before beginning the assembly and alignment of the OFI.

# OFI Assembly

## Overview of the D0900136 OFI Assembly

The OFI is comprised of the following sub-assemblies, as shown in Figure 1.

D0900048 DAMPER HOLDER ASSEMBLY

D0900623 FARADAY ISOLATOR TABLE ASSY

D0900586 UPPER WIRE ASSY

D0900170 EARTHQUAKE CROSSBAR ASSY

D1002256 EARTHQUAKE CROSSBAR_IN ASSY

D0900579 BLADE GUARD ASSY

D0900655 STRUCTURAL WELDMENT ASSY, OMC

Note: Use torque values specified in T1100066 Torque Values, unless a different value is indicated on the sub-assembly drawing.

Table 1: Bill of Materials, D0900136 Faraday Isolator Assembly

STRUCTURAL

WELDMENT

EARTHQUAKE CROSSBAR ASSY

EARTHQUAKE CROSSBAR_IN ASSY

BLADE GUARD ASSY

UPPER WIRE ASSEMBLY

TABLE ASSY

DAMPER HOLDER ASSEMBLY

Figure 1: Output Faraday Isolator Assembly

## D0900048 Faraday Isolator Damper Holder Assembly

## D0900614 PRISM MOUNT ASSY_LH

Table 2: Bill of Materials, D0900614 Prism Mount Assy LH

Figure 2: Prism Mount Assy_LH

### Step 1

### Step 2

### Step 4

### Step 5

### Step 6

## D0900615 PRISM MOUNT ASSY_RH

Table 3: Bill of Materials, D0900615 Prism Mount Assy RH

The assembly of the Right Hand Prism Mount follows the same steps detailed in section D0900614 PRISM MOUNT ASSY_LH, with the addition of the “W” baffles, detailed in the following steps.

Figure 3: Prism Mount Assy_RH

### Step 5

### Step 6

## D1002364 Faraday Isolator Beam Dump Assy

Table 4: Bill of Materials, D1002364 Faraday Isolator Beam Dump Assy

Figure 4: 3.5	D1002364 Faraday Isolator Beam Dump Assy

## D0900353 Half Wave Plate Holder Assy

Table 5: Bill of Materials, D0900353 Half Wave Plate Holder Assy

| ITEM NO. | PART NUMBER | DESCRIPTION | MATERIAL | QTY |
| --- | --- | --- | --- | --- |
| 1 | D1100029 | ALIGO AOS HALF WAVE QUAARATZ ASSY | N/A | 1 |
| 2 | C-808-N | SCREW, SHC, 8-32 x 1/2, MS16995-26 | 300 SSTL | 1 |
| 3 | D0900352 | HALF WAVE PLATE HOLDER | 6061-T6 Al | 1 |

Figure 5: Half Wave Plate Holder Assy

## D0900440 TFP Polarizer Plate Assy

Table 6: Bill of Materials, D0900440 TFP Polarizer Plate Assy

Figure 6: 2.7	D0900440 TFP Polarizer Plate Assy

## D1001958 Wire Support Block Assy

Table 7: Bill of Materials, D1001958 Wire Support Block Assy

Figure 7: 3.8	D1001958 Wire Support Block Assy

## D1001963 Output Alignment Fixture Assy

Table 8: Bill of Materials, D1001963 Output Alignment Fixture Assy

Figure 8: Output Alignment Fixture Assy

## D1200197 Wedged Prism Holder Assy

Table 9: Bill of Materials, D1200197 Wedged Prism Holder Assy

Figure 9: 3.10	D1200197 Wedged Prism Holder Assy

### D0900359 Wedged Prism Assy

Note: The thick side of the Wedged Prism must be aligned with the scribed line on Part 1. When the wedged prism is installed on the optics table, the thick side of the wedged prism must be on top.

## Faraday Isolator Table Assy

Table 10: Bill of Materials, D0900623 Faraday Isolator Table Assy

Figure 10: Assembly of Faraday Isolator Table

### Step 1

### Step 2

Front Balance Weight Installation: see 3.6.1 Adjust Balance Weights.

### Step 3

### Step 4

### Step 5

### Step 6

### Step 7

### Step 8

### Step 9

### Step 10

### Step 11

Rear Balance Weight Installation: see 3.6.1 Adjust Balance Weights.

### Step 12

### Step 13

### Step 14

Place 2 magnets on each Magnetic Attachment Plate in pairs with opposite polarity, at the locations shown in Figure 11.

S

N

S

N

Figure 11: Magnet Locations, with Opposite Pole Orientations

## D0900586 Faraday Isolator Upper Wire Assembly

Table 11: Bill of Materials, D0900586 Faraday Isolator Upper Wire Assembly

### OFI Wire Length

#### Attaching the upper wire clamp

Select a piece of wire that is approximately 5 in longer than twice the length of the D1201506 OFI Wire Length Fixture, shown in Figure 12.

Press the Pin 5 into the D1002168 Music Wire Split Clamp 3, as shown in Figure 14: NOTE: the upper wire clamp pair, Clamp 3 and Clamp 4 must be used only for the upper end. Loop the wire at the midpoint once over Pin 5, and clamp the upper loop by attaching the other half of the wire clamp, D1002168 Music Wire Split Clamp 4, using the two SHCS, 2-56 x 3/16 long: tighten the screws iteratively to maintain a uniform gap between the clamping halves. The two halves of the wire should extend uniformly beyond the small ends of the assembled split wire clamp.

#### Making the Bottom Loop

Place the upper split wire clamp into the recess on the right edge of the OFI Wire Length Fixture and hold it in place. Grab one of the extending wire lengths with a pair of pliers and pull it across the top of the OFI Wire Length Fixture toward the opposite end. While maintaining tension on the wire, loop it half way around the small pin protruding from the surface at the left end of the OFI Wire Length Fixture. Place the D1201508 OFI Wire Loop Tool over the wire, mating it with the protruding pin, as shown in Figure 12. Rotate the wire loop tool clockwise 1 1/8 turn, while maintaining tension of the wire; the wire end should be pointing approximately toward the right end of the wire length fixture. Remove the wire loop tool, and lift the looped end from the pin. Cut the end of the wire to a length approximately ½ in beyond the formed double loop.

Repeat the above steps to form a loop on the other length of wire from the upper clamp.

Figure 12: OFI Wire Length Fixture with Wire Loop Tool

D1201508 OFI Wire Loop Tool ?

Figure 13: Right End of Wire Length Fixture with Recess for Inserting the Small End of the Upper Wire Clamp

Figure 14: Upper Wire Loop Clamp Assembly

#### Attaching the Bottom Clamp

Thread the double looped end of the wire through the D0900588 Wire Adjustable Adapter, with the machined flat sides facing toward the upper wire clamp as shown in Figure 15.

Assemble both halves of the bottom Music Wire Split Clamp, D0900582 and D0900583, and loosely attach the 2-56 SHCS. Wet the D0900582 press fit hole for the pin 5 with clean alcohol and insert the pin through the clearance hole in D0900583 until it mates with the press fit hole. Hold the Split Clamp assembly with pliers and use a hammer to lightly tap the pin into D0900582 Music Wire Split Clamp 1; the pin must be tapped until it is flush with the outer surface of D0900583.

Separate the two halves of the Music Wire Split Clamp. Place the double loop at the end of the wire over the pin in Music Wire Split Clamp 1, making sure that the double loop is facing away from the clamp so that it will fit into the recess in D0900583.

NOTE: The lower wire clamp pair, Clamp 1 and Clamp 2 must be used only for the lower end--the pin is purposely offset from the center of the split clamp; the double wire loop must be oriented so that the continuous wire passes through the center of the clamp, and the cut end is offset;

Align the long music wire with the groove in the Split Clamp 1. Clamp the wire by attaching the other half of the wire clamp, D0900583 Music Wire Split Clamp 2, using the two SHCS, 2-56 x 3/16 long; tighten the screws iteratively to maintain a uniform gap between the clamping halves. The wire should extend from the center of the small end of the assembled split wire clamp.

Figure 15: Bottom Wire Clamp Assembly

## D0900170 Earthquake Crossbar Assy

## D1002256 Earthquake Crossbar_In Assy

## D0900579 Blade Guard Assy

## D0900655 Structural Weldment Assy, OMC

Table 12: Bill of Material, D0900655 Structural Weldment Assy, OMC

Figure 16: 2.16	D0900655 Structural Weldment Assy, OMC

# OFI Balancing

## Install Faraday Isolator Damper Holder Assembly

Mount the Damper Holder Assembly to the bottom cross beams, as shown in Figure 17. Set the nominal height of the top of the mounting bracket to 0.62 in, as shown in Figure 18.

Figure 17: Mounting of the Damper Holder Assembly

Figure 18: Nominal Height of Damper Holder Assy

## Install Blade Springs

Mount the D0901514 BLADE CLAMP PLATFORM to the top end cross member of the structural Weldment, as shown in Figure 1: Output Faraday Isolator Assembly.

Clamp the D0900541 FI Upper Blade Spring with the D0900566 Up Blade Clamp Top.

Note: The blade pairs are listed in Table 13; place the stiffest blade of the pair on the input side of the OFI.

## Install Blade Guard Assembly

Mount two Blade Guard Assemblies, D0900579 Blade Guard Assy, to the OMC structure top, as shown in Figure 1. Adjust the length of the 8-32 bolt to extend 0.45in below the lower side of Blade Guard Crosspiece, as shown in D0900579.

Loosen the D0900566 Up Blade Clamp Top slightly so that the tip of blade spring tip can be moved laterally to align with the 8-32 bolt; then, firmly re-clamp the blade spring.

## Install Upper Wire Assembly

See 2.12 D0900586 Faraday Isolator Upper Wire Assembly.

### Clamp Wire Clamp to Tip of Blade Spring

Clamp the Music Wire Split Clamp to the tip of the blade spring with the Split Clamp Clamp D1300321, as shown in

Split Clamp Clamp

Figure 19: Split Clamp Clamp D1300321

## Install Table Assy

See 2.11 Faraday Isolator Table Assy.

Attach the Wire Adjustable Adapter to the ¼-20 bolt, as shown in Figure 28. Screw the adapter onto the bolt with approximately 0.25 in gap between the adapter bottom and the nut; this gap will be adjusted during the subsequent table balancing step.

## Balance Table

### Adjust Balance Weights

Two sets of balance weights, CW1 and CW2, are used to balance the suspended OFI optical table; the weights and their positions are specific to a particular pair of as-built matched blade springs. The matched blade springs for the light optical table_D0900015-v4 are listed Table 13 below, together with the magnitude and mounting positions for the weights; the matched blade springs for the heavy optical table_D0900015-v5 are listed Table 14. The off-set mounting positions are described in the Figure 24 and Figure 25. Load each spring to its measured suspended weight by placing the specified balance weight at the correct position; this will also place the center of gravity of the optics table approximately at the midpoint between the suspension wires, as indicated by the intersection of the lines drawn between the four suspension wires in Figure 24 and Figure 25.

Table 13: Light Optical Table_D0900015-v4: Balance Weights and Positions

| E1300043-v2 Blade pairs | S/N | Blade Balance Weight, kg | Blade Balance Weight, lbs | x-offset of CW1 from mounting bolt, in | CW1 (centered), lbs |  | CW1 assembly |  | y-offset of CW2 from mounting bolt, in | CW2 (offset), lbs |  | CW1 assembly |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  | 1# | 0.5# | 0.25# |  |  | 3# | 0.5# | 0.25# |
| 1, 1st article | 11 | 6.50 | 14.33 | -0.276 | 0.50 | 0 | 1 | 0 | 0.717 | 3.25 | 1 | 0 | 1 |
|  | 7 | 6.50 | 14.33 |  |  |  |  |  |  |  |  |  |  |
| 2, LLO | 5 | 6.45 | 14.22 | 0.374 | 0.75 | 0 | 1 | 1 | 0.770 | 3.00 | 1 | 0 | 0 |
|  | 9 | 6.45 | 14.22 |  |  |  |  |  |  |  |  |  |  |
| 3, LHO | 4 | 6.65 | 14.66 | 0.000 | 1 | 1 | 0 | 0 | 0.673 | 3.50 | 1 | 1 | 0 |
|  | 1 | 6.60 | 14.55 |  |  |  |  |  |  |  |  |  |  |
| 5, H2 | 2 | 6.55 | 14.44 | 0.000 | 0.75 | 0 | 1 | 1 | 0.717 | 3.25 | 1 | 0 | 1 |
|  | 6 | 6.45 | 14.22 |  |  |  |  |  |  |  |  |  |  |
| 6, spare | 12 | 6.80 | 14.99 | 0.294 | 1.50 | 1 | 1 | 0 | 0.635 | 3.75 | 1 | 1 | 1 |

Figure 24: Light Optical Table_D0900015-v4: Off-set Positions for Balance Weights CW1 and CW2

Table 14: Heavy Optical Table_D0900015-v5: Balance Weights and Positions

| Blade pairs | E1300043-v3 S/N | Location | Blade Balance Weight, kg | Blade Balance Weight, lbs | x-offset of CW1 from mounting bolt, in | CW1, lbs | CW1 weight assembly |  |  |  | y-offset of CW2 from mounting bolt, in | CW2, lbs | CW2 weight assembly |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  |  | 1# | 0.5# | 0.25# | 0.13# |  |  | 3# | 1# | 0.5# | 0.25# |
| 1 | 11 | Input | 9.22 | 20.33 | 0.218 | 1.63 | 1 | 1 | 0 | 1 | -0.575 | 3.75 | 1 | 0 | 1 | 1 |
|  | 12 | Output | 9.22 | 20.33 |  |  |  |  |  |  |  |  |  |  |  |  |
| 2 | 13 | Input | 9.15 | 20.18 | 0.296 | 1.63 | 1 | 1 | 0 | 1 | -0.575 | 3.75 | 1 | 0 | 1 | 1 |
|  | 14 | Output | 9.17 | 20.22 |  |  |  |  |  |  |  |  |  |  |  |  |
| 3 | 15 | Input | 9.17 | 20.22 | -0.296 | 1.50 | 1 | 1 | 0 | 0 | -0.480 | 3.75 | 1 | 0 | 1 | 1 |
|  | 19 | Output | 9.19 | 20.26 |  |  |  |  |  |  |  |  |  |  |  |  |
| 4 | 16 | Input | 9.07 | 20.00 | -0.030 | 1.38 | 1 | 0 | 1 | 1 | -0.413 | 3.50 | 1 | 0 | 1 | 0 |
|  | 18 | Output | 8.99 | 19.82 |  |  |  |  |  |  |  |  |  |  |  |  |
| 5 | 17 | Input | 9.20 | 20.29 | 0.046 | 1.75 | 1 | 1 | 1 | 0 | -0.627 | 4.00 | 1 | 2 | 0 | 0 |
|  | 20 | Output | 9.32 | 20.55 |  |  |  |  |  |  |  |  |  |  |  |  |
| total weights needed |  |  |  |  |  |  | 5 | 4 | 2 | 3 |  |  | 5 | 2 | 4 | 3 |

Figure 25: Heavy Optical Table_D0900015-v5: Off-set Positions for Balance Weights CW1 and CW2

#### D0900015-v5 Weight Offset Locations—Pair 1

#### D0900015-v5 Weight Offset Locations—Pair 2

#### D0900015-v5 Weight Offset Locations—Pair 3

#### D0900015-v5 Weight Offset Locations—Pair 4

#### D0900015-v5 Weight Offset Locations—Pair 5

### Setting Gap Between Blade Guard Assembly and Spring Blade

After the balance weights are added, retract the 8-32 screws in the Blade Guard Assembly until the spring blade tip is approximately 0.25in and is free to move vertically without any constraints. Lock the nuts to secure the 8-32 screws.

Figure 26: Retract Spring Restraints 0.25in so that Blades Move Freely in Vertical Direction

### Adjust Wire Lengths

- Place the guitar tuner against the base of the blade spring, as shown in Figure 27, so that it will pick up the sound of the suspension wire when strummed.

Figure 27: Guitar Tuner Placed at Base of Blade Spring

- Strum the wire as you would a guitar string and watch the frequency on the guitar tuner. Tension the wire by turning the Wire Adjustable Adapter while holding the Music Wire Split Clamp with a pair of pliers so that the Music Wire Split Clamp does not rotate within the Wire Adjustable Adapter, until the desired frequency is displayed on the guitar tuner. The objective is to match the frequencies of the pair of wires connected to the same blade spring. In general, the wire pairs from the front and back blade springs will have different frequencies because the blade springs may have different stiffness.

Wire Adjustable Adapter

Music Wire Split Clamp

Figure 28: Wire Tension Adjustment

- Place the Table Height Gauge (Figure 29) on the OMC Structure Base to set the nominal height of the suspended optical table 0.105 in above the base, as shown in Figure 30.

Error! Reference source not found.

Figure 29: Table Height Gauge, D1201449

This may result in the wire tuning changing. To reach a compromise between equal wire tension and equal table height on both sides, the balance weight CW2 may need to be shifted laterally to a position different from that listed in Table 13.

Figure 30: Nominal Height of Optics Table above the Structure Base

- Verify that the Damper Holder Assy bracket is 0.515 in above the Optics Table by placing the Table Height Gauge (Figure 29) on the Damper Holder Assy bracket and using the other end of the height gauge.

## Install Earthquake Crossbar Assy

Attach the Output Faraday Isolator Earthquake Stop Post D1300136 to the four corners of the optics table, as shown in Figure 31.

Earthquake Stop Post

Figure 31: Earthquake Stop Posts Attached to Four Corners of the Optics Plate

Assemble the Earthquake Crossbar input and output assemblies loosely--see D0900170 Earthquake Crossbar Assy and D1002256 Earthquake Crossbar_In Assy--and attach the crossbar sides to the structural Weldment at the input and output ends of the weldment, as shown in Figure 1 and Error! Reference source not found.. Move the crossbar plate within the slotted mounting holes to align the earthquake stop clearance holes with the earthquake stop rods; then firmly bolt the crossbar plate to the crossbar sides.

# OFI Optical Alignment

## OFI Preliminary Optical Alignment

The preliminary optical alignment of the Output Faraday Isolator will take place on an optical bench in a class A environment.

### Equipment List

- 250 mW YAG Laser (NPRO), 1064nm wavelength

- 1064 nm Photodetector on adjustable height, movable stand; sensitivity 0.25 W to 0.25 E-4 W

- HR steering mirror on adjustable height stand with tip-tilt mount

- 50% beam splitter on adjustable height stand with tip-tilt mount

- HR retro mirror on adjustable height stand with tip-tilt mount

### Alignment Setup

- Clamp the OFI structural frame to the middle of an optical table

- Mount the input laser, the input laser aperture, the steering mirror, and the beam splitter approximately 171.4 mm above the surface of the optical table in accordance with the optical layout shown in Figure 32. The height of the OFI optical centerline will be 167.8 mm above the table top, and the input laser beam will tilt down toward the input aperture, before being turned horizontal by the OFI wedge prism. The retro-mirror should be approximately 168 mm above the table.

- Adjust the height of the input laser, and position and tilt the steering mirror until the laser beam passes through both the input aperture and the output aperture of the OFI.

- Adjust the tilt of the retro-mirror so that the retro-reflected laser beam passes back through the OFI and through the input laser aperture—during this step, rotate the half-wave plate to maximize the retro-reflected power; then displace the beam slightly until the retro-reflected beam is blocked by the input laser aperture.

BEAM DUMP

PERISCOPE

INPUT LASER

HALF-WAVE PLATE

INPUT APERTURE

INPUT POLARIZER

PD 2

50 % BS

INPUT ALIGNMENT APERTURE

PD 4

OUTPUT ALIGNMENT APERTURE

OUTPUT FARADAY ISOLATOR

RETRO MIRROR

PD 3

PD 1

Figure 32: Optical Layout for OFI Alignment

### Optical Alignment Procedure

- Set the power of the laser to approximately 250 mW. Calibrate the transmissivity and reflectivity of the beam splitter by placing the photodetector alternately at position PD1 and PD2 and measure the power levels. The power at PD2 will be the reference power for calculating the Transmissivity and the power extinction ratio of the OFI.

- Place the photodetector at position PD4 and rotate the half-wave plate to minimize the power at PD4; measure the power.

- Place the photodetector at position PD3 and measure the power.

- Calculate the OFI Transmissivity as the ratio of PD3 to PD2; calculate the OFI power extinction ratio as the ratio of PD4 to PD2.

#### Test Data

## OFI Shack-Hartmann Wavefront Test

The OFI Shack-Hartmann Wavefront Test will take place on an optical bench in a class A environment.

### Equipment List

- WFS300-14AR--00-OFI1 Shack-Hartmann WFS

- S1FC635--00-OFI1 Fabry-Perot  Benchtop  Laser Source

- P1-630A-FC-5--00-OFI1 Single Mode Fiber Patch Cable

- LW-3-2037-C--00-OFI1 Large Wedged Windows

- KM200--00-OFI1 Kinematic  Mount

- F810FC-635--00-OFI1 FC/PC Collimation Package

- AOSSLCOFI4--00-OFI1 3" Mirror Mount

- AOSSLCOFI2--00-OFI1 XY Stage Assembly

- AOSSLCOFI1--00-OFI1 5 Axis Mirror Mount Assembly

### Alignment Setup

Set up the SH test apparatus according to the optical schematic shown in Figure 33. Align the SH beam to pass freely through the entrance and exit apertures of the OFI.

SH Sensor

Beam Dump

SH BS

Steering mirrors

SH LASER

OUTPUT FARADAY ISOLATOR

RETRO

MIRROR 2

RETRO

MIRROR 1

PD 3

Figure 33: Optical Layout for SH Test

#### Reference Beam Alignment

Place the lambda/10 3” mirror in front of the OFI entrance aperture and steer the mirror so that the reflected beam is centered on the SH sensor.

Iteratively align the retro-reflected beam and the SH sensor by adjusting the steering mirror, the lateral position of the SH sensor, and the tilt of the SH sensor to center the beam on the SH sensor and to reduce the wavefront tilt by minimizing the Zernike Z2 and Z3 tilt coefficients.

The beam centroid should be approximately centered on the SH sensor, as shown in Figure 34.

Figure 34: Contour of the Input Beam on Shack-Hartmann Sensor

The power measured by the central array of the SH lens sensor matrix is shown in Figure 35.

Figure 35: Reference Power Read by the Shack-Hartman Lens Array

#### Reference Beam Wavefront

The reconstructed wavefront of the reference beam is shown in Figure 36.

Figure 36: Reference Wavefront

The Zernike coefficients for reconstructing the reference wavefront will be saved in a software file and be used subsequently to subtract the reference wavefront from the double pass wavefront through the OFI.

### Double Pass OFI Test Procedure

Place the lambda/10 3” mirror beyond the exit aperture OFI entrance aperture and steer the mirror so that the reflected beam passes back through the OFI and is centered on the SH detector.

Iteratively align the retro-reflected beam and the SH sensor by adjusting the steering mirror, the lateral position of the SH sensor, and the tilt of the SH sensor to center the beam on the SH sensor and to reduce the wavefront tilt by minimizing the Zernike Z2 and Z3 tilt coefficients—as was done in section 4.2.2.

Remove the tilt Zernike coefficients by means of the SH software commands. A typical double-pass wavefront through the OFI with tilt removed is shown in Figure 37.

Figure 37: Typical OFI Double-pass Wavefront with Tilt Removed

### Double Pass OFI Test Results

Typical OFI double pass test results are shown in Figure 38. The Zernike coefficients are listed in the figure. Zernike Coefficients 01, 02, 03, and 05, in units of micron, describe the piston, Y-tilt, X-tilt, and defocus respectively; these coefficients do not affect the wavefront aberration and will be ignored.

The astigmatism coefficients 04 and 06 are the most significant because they are responsible for direct coupling loss into the OMC. The higher order coefficients describe coma and other irregularities, which may also cause coupling loss due to generation of higher order Gaussian modes in the OMC cavity.

Figure 38: Typical OFI Double-pass Zernike Test Results