
|  |  |  |  |
| --- | --- | --- | --- |
| **Test Date** | Apr. 26-29, 2013 | | |
| **Author(S)** | Liyuan Zhang, Margot Phelps, GariLynn Billingsley | | |
| **Approval(s)** |  | | |
| **Specification Doc.** | LIGO-E0900068 | Specification | < 500 ppm |
| **Procedure Doc.** | LIGO-E1000863 | Mean ± Error\* | 820 ± 10 ppm |
| **Conclusion** |  | | |

\* Error is the calibration error, which is ~1%.

**Discussions and Comments:**

**The scan was carried out over aperture of 160 mm in diameter with the beam and step sizes of 1 mm. The calibration is done by normalizing the AR reflection signal to the signal with a 1” HR mirror (T=70 ppm) and the variation of power during scan is monitored and corrected, the result is summarized in Fig.1.**

**The shallow fringes seem to be caused by interference between the AR reflection and the tail of strong reflection of HR. Since the wedge angle is only 0.07~0.1º (D080658), the separation of the two beam was not large enough at the measurement.**

****

**Fig. 1 ETM08 S2 AR reflection over an aperture of 160 mm in diameter.**