## 50-50 splitter
Optics calibration |  |  | Filters for power calibration
Before | 0.356 | W | ND 1 transmittivity | 0.04 | Filters used
After | 0.292 | W | ND2 transmittivity | 0.08 | ND = 2.0
Transmissivity | 0.82022 |  | Before | 0.01546 | ND = 1.0
|  |  | After | 2.149e-06
Beam splitter calibration |  |  | Transmissivity | 0.000139
Before |  | mW | Total transmissivity | 0.0032
After |  | mW
Transmissivity | 0.5
Total transmissivity | 0.41011
Power meter diode calibration
Diode voltage (uV) | Calibrated power meter (uW)
0.966 | 15.2
1.644 | 8.6
2.142 | 2.7
4.073 | 0.41
3.01 | 0.321
3.674 | 0.373
Enter Voltage | Power hitting sample in W:
3.77 | 0.38429
WARNING: No Calibration data available for BRDF
Lock-in Calibration
Lock-in with filters (mV) | Power stick (uW) | Power hitting diode due to filter (uW)
68 | 116 | 0.3712
122 | 201 | 0.6432
162 | 270 | 0.864
27.6 | 46 | 0.1472
9.1 | 16.4 | 0.05248
Enter Voltage (mV) | Power hitting filter is (uW)
0.071 | 0.0046256
Solid Angle Calculation
Distance from
sample to aperture (cm) | 30
Aperture radius (cm) | 0.635
mirror hole radius (cm) | 0.3
Steradian is (cm^2) | 900
Steradian for mirror is (cm^2) | 400
Aperture area is (cm^2) | 1.2661
Hole area is (cm^2) | 0.2826
Solid Angle is: | 0.00070031
Power incident on sample | 0.375 | watts
Power scattered | 0.011279 | microwatts
Solid Angle | 0.00070031 | steradians
Power scattered
per watt*steradian | 4.2948e-05
IMPOTANT INFO
1064 nm splitter |  | Normal is ---->
CVI made
BRDF for P-polarization
Angle | Angle corrected | Voltage (mV) | Background corrected Voltage | Power on photodiode (uW) | Power scattered (uW) | BRDF
14 | 4 | 12.8 | 12.799 | 0.071957 | 0.17546 | 0.00066811
20 | 10 | 5 | 4.999 | 0.030695 | 0.074845 | 0.000285
25 | 15 | 1.4 | 1.399 | 0.011651 | 0.028409 | 0.00010818
30 | 20 | 1.4 | 1.399 | 0.011651 | 0.028409 | 0.00010818
35 | 25 | 0.8 | 0.799 | 0.0084767 | 0.020669 | 7.8705e-05
40 | 30 | 0.51 | 0.509 | 0.0069426 | 0.016929 | 6.4461e-05
45 | 35 | 0.37 | 0.369 | 0.006202 | 0.015123 | 5.7585e-05
50 | 40 | 0.3 | 0.299 | 0.0058317 | 0.01422 | 5.4147e-05
55 | 45 | 0.25 | 0.249 | 0.0055672 | 0.013575 | 5.1691e-05
60 | 50 | 0.26 | 0.259 | 0.0056201 | 0.013704 | 5.2182e-05
64 | 54 | 0.28 | 0.279 | 0.0057259 | 0.013962 | 5.3165e-05
66 | 56 | 0.27 | 0.269 | 0.005673 | 0.013833 | 5.2673e-05
BRDF for S-polarization
Angle | Angle corrected | Voltage (mV) | Background corrected Voltage | Power on photodiode (uW) | Power scattered (uW) | BRDF
14 | 4 | 9.4 | 9.399 | 0.053971 | 0.1316 | 0.00050111
20 | 10 | 5.6 | 5.599 | 0.033869 | 0.082584 | 0.00031447
25 | 15 | 1.75 | 1.749 | 0.013502 | 0.032923 | 0.00012537
30 | 20 | 1.9 | 1.899 | 0.014296 | 0.034858 | 0.00013273
35 | 25 | 0.95 | 0.949 | 0.0092702 | 0.022604 | 8.6073e-05
40 | 30 | 0.5 | 0.499 | 0.0068897 | 0.0168 | 6.397e-05
45 | 35 | 0.33 | 0.329 | 0.0059904 | 0.014607 | 5.562e-05
50 | 40 | 0.17 | 0.169 | 0.005144 | 0.012543 | 4.7762e-05
55 | 45 | 0.13 | 0.129 | 0.0049324 | 0.012027 | 4.5797e-05
60 | 50 | 0.093 | 0.092 | 0.0047367 | 0.01155 | 4.398e-05
65 | 55 | 0.071 | 0.07 | 0.0046203 | 0.011266 | 4.2899e-05

## Laser Black 3188
Optics calibration |  |  | Filters for power calibration
Before | 0.356 | W | ND 1 transmittivity | 0.04 | Filters used
After | 0.292 | W | ND2 transmittivity | 0.08 | ND = 2.0
Transmissivity | 0.82022 |  | Before | 0.01546 | ND = 1.0
|  |  | After | 2.149e-06
Beam splitter calibration |  |  | Transmissivity | 0.000139
Before |  | mW | Total transmissivity | 0.0032
After |  | mW
Transmissivity | 0.5
Total transmissivity | 0.41011
Power meter diode calibration
Diode voltage (uV) | Calibrated power meter (uW)
0.966 | 15.2
1.644 | 8.6
2.142 | 2.7
4.073 | 0.41
3.01 | 0.321
3.674 | 0.373
Enter Voltage | Power hitting sample in W:
3.77 | 0.38429
Lock-in Calibration
Lock-in with filters (mV) | Power stick (uW) | Power hitting diode due to filter (uW)
68 | 116 | 0.3712
122 | 201 | 0.6432
162 | 270 | 0.864
27.6 | 46 | 0.1472
9.1 | 16.4 | 0.05248
Enter Voltage (mV) | Power hitting filter is (uW)
45 | 0.2423
Solid Angle Calculation
Distance from
sample to aperture (cm) | 30
Aperture radius (cm) | 0.635
mirror hole radius (cm) | 0.3 |  | BRDF for P-polarization
Steradian is (cm^2) | 900
Steradian for mirror is (cm^2) | 400 |  | Angle | Angle corrected | Voltage (mV) | Background corrected Voltage | Power on photodiode (uW) | Power scattered (uW) | BRDF
Aperture area is (cm^2) | 1.2661 |  | 9 | 0 | 149 | 149 | 0.79246 | 1.9323 | 0.069854
Hole area is (cm^2) | 0.2826 |  | 15 | 6 | 127 | 127 | 0.67608 | 1.6485 | 0.059595
Solid Angle is: | 0.00070031 |  | 20 | 11 | 140 | 140 | 0.74485 | 1.8162 | 0.065657
|  |  | 25 | 16 | 153 | 153 | 0.81362 | 1.9839 | 0.071719
Power incident on sample | 0.0395 | watts | 30 | 21 | 138 | 138 | 0.73427 | 1.7904 | 0.064724
Power scattered | 0.59081 | microwatts | 35 | 26 | 144 | 144 | 0.76601 | 1.8678 | 0.067522
Solid Angle | 0.00070031 | steradians | 40 | 31 | 131 | 131 | 0.69724 | 1.7001 | 0.06146
Power scattered |  |  | 45 | 36 | 143 | 143 | 0.76072 | 1.8549 | 0.067056
per watt*steradian | 0.021358 |  | 50 | 41 | 137 | 137 | 0.72898 | 1.7775 | 0.064258
|  |  | 55 | 46 | 127 | 127 | 0.67608 | 1.6485 | 0.059595
|  |  | 60 | 51 | 128 | 128 | 0.68137 | 1.6614 | 0.060061
Normal is ---->
BRDF for S-polarization
|  |  | Angle | Angle corrected | Voltage (mV) | Background corrected Voltage | Power on photodiode (uW) | Power scattered (uW) | BRDF
|  |  |  | -9 |  | -1 | -0.00104 | -0.0025359 | -9.1674e-05
|  |  |  | -9 |  | -1 | -0.00104 | -0.0025359 | -9.1674e-05
|  |  |  | -9 |  | -1 | -0.00104 | -0.0025359 | -9.1674e-05
|  |  |  | -9 |  | -1 | -0.00104 | -0.0025359 | -9.1674e-05
|  |  |  | -9 |  | -1 | -0.00104 | -0.0025359 | -9.1674e-05
|  |  |  | -9 |  | -1 | -0.00104 | -0.0025359 | -9.1674e-05
|  |  |  | -9 |  | -1 | -0.00104 | -0.0025359 | -9.1674e-05
|  |  |  | -9 |  | -1 | -0.00104 | -0.0025359 | -9.1674e-05

## SS #4 Rough Side
Optics calibration |  |  | Filters for power calibration
Before | 0.356 | W | ND 1 transmittivity | 0.04 | Filters used
After | 0.292 | W | ND2 transmittivity | 0.08 | ND = 2.0
Transmissivity | 0.82022 |  | Before | 0.01546 | ND = 1.0
|  |  | After | 2.149e-06
Beam splitter calibration |  |  | Transmissivity | 0.000139
Before |  | mW | Total transmissivity | 0.0032
After |  | mW
Transmissivity | 0.5
Total transmissivity | 0.41011
Power meter diode calibration
Diode voltage (uV) | Calibrated power meter (uW)
0.966 | 15.2
1.644 | 8.6
2.142 | 2.7
4.073 | 0.41
3.01 | 0.321
3.674 | 0.373
Enter Voltage | Power hitting sample in W:
3.77 | 0.38429
Lock-in Calibration
Lock-in with filters (mV) | Power stick (uW) | Power hitting diode due to filter (uW)
68 | 116 | 0.3712
122 | 201 | 0.6432
162 | 270 | 0.864
27.6 | 46 | 0.1472
9.1 | 16.4 | 0.05248
Enter Voltage (mV) | Power hitting filter is (uW)
45 | 0.2423
Solid Angle Calculation
Distance from
sample to aperture (cm) | 30
Aperture radius (cm) | 0.635
mirror hole radius (cm) | 0.3 |  | BRDF for P-polarization
Steradian is (cm^2) | 900
Steradian for mirror is (cm^2) | 400 |  | Angle | Angle corrected | Voltage (mV) | Background corrected Voltage | Power on photodiode (uW) | Power scattered (uW) | BRDF
Aperture area is (cm^2) | 1.2661 |  | 25 | 16 | 122 | 121.999 | 0.64962 | 1.584 | 0.059997
Hole area is (cm^2) | 0.2826 |  | 30 | 21 | 114 | 113.999 | 0.6073 | 1.4808 | 0.056089
Solid Angle is: | 0.00070031 |  | 35 | 26 | 100 | 99.999 | 0.53324 | 1.3002 | 0.049249
|  |  | 40 | 31 | 74 | 73.999 | 0.3957 | 0.96487 | 0.036546
Power incident on sample | 0.0377 | watts | 45 | 36 | 73 | 72.999 | 0.39041 | 0.95197 | 0.036057
Power scattered | 0.59081 | microwatts | 50 | 41 | 67.3 | 67.299 | 0.36026 | 0.87845 | 0.033272
Solid Angle | 0.00070031 | steradians | 55 | 46 | 62 | 61.999 | 0.33222 | 0.81008 | 0.030683
Power scattered |  |  | 60 | 51 | 58 | 57.999 | 0.31106 | 0.75849 | 0.028729
per watt*steradian | 0.022378
BRDF for S-polarization
|  | Normal is ----> | Angle | Angle corrected | Voltage (mV) | Background corrected Voltage | Power on photodiode (uW) | Power scattered (uW) | BRDF
|  |  |  | -9 |  | -1 | -0.00104 | -0.0025359 | -9.6051e-05
|  |  |  | -9 |  | -1 | -0.00104 | -0.0025359 | -9.6051e-05
|  |  |  | -9 |  | -1 | -0.00104 | -0.0025359 | -9.6051e-05
|  |  |  | -9 |  | -1 | -0.00104 | -0.0025359 | -9.6051e-05
|  |  |  | -9 |  | -1 | -0.00104 | -0.0025359 | -9.6051e-05
|  |  |  | -9 |  | -1 | -0.00104 | -0.0025359 | -9.6051e-05
|  |  |  | -9 |  | -1 | -0.00104 | -0.0025359 | -9.6051e-05
|  |  |  | -9 |  | -1 | -0.00104 | -0.0025359 | -9.6051e-05

## SS Polished
Optics calibration |  |  | Filters for power calibration
Before | 0.356 | W | ND 1 transmittivity | 0.04 | Filters used
After | 0.292 | W | ND2 transmittivity | 0.08 | ND = 2.0
Transmissivity | 0.82022 |  | Before | 0.01546 | ND = 1.0
|  |  | After | 2.149e-06
Beam splitter calibration |  |  | Transmissivity | 0.000139
Before |  | mW | Total transmissivity | 0.0032
After |  | mW
Transmissivity | 0.5 |  | Power meter filter calibration
Total transmissivity | 0.41011 |  | Before | 1.01e-05
|  |  | After | 0.00588
Power meter diode calibration
Diode voltage (uV) | Calibrated power meter (uW)
0.966 | 15.2
1.644 | 8.6
2.142 | 2.7
4.073 | 0.41
3.01 | 0.321
3.674 | 0.373
Enter Voltage | Power hitting sample in W:
3.77 | 0.38429
|  |  | Transmissivity | 0.0017177
Lock-in Calibration |  |  | Power measured | 0.000518
Lock-in with filters (mV) | Power stick (uW) | Power hitting diode due to filter (uW) | Actual Power | 0.30157
68 | 116 | 0.3712
122 | 201 | 0.6432
162 | 270 | 0.864
27.6 | 46 | 0.1472
9.1 | 16.4 | 0.05248 |  | WARNING: No calibration data available for BRDF
Enter Voltage (mV) | Power hitting filter is (uW)
45 | 0.2423
Solid Angle Calculation
Distance from
sample to aperture (cm) | 30
Aperture radius (cm) | 0.635
mirror hole radius (cm) | 0.3 |  | BRDF for P-polarization
Steradian is (cm^2) | 900
Steradian for mirror is (cm^2) | 400 |  | Angle | Angle corrected | Voltage (mV) | Background corrected Voltage | Power on photodiode (uW) | Power scattered (uW) | BRDF
Aperture area is (cm^2) | 1.2661 |  | 25 | 16 | 5 | 4.999 | 0.030695 | 0.074845 | 0.00035389
Hole area is (cm^2) | 0.2826 |  | 30 | 21 | 3.3 | 3.299 | 0.021702 | 0.052916 | 0.0002502
Solid Angle is: | 0.00070031 |  | 35 | 26 | 3.6 | 3.599 | 0.023289 | 0.056786 | 0.0002685
|  |  | 40 | 31 | 1.68 | 1.679 | 0.013132 | 0.03202 | 0.0001514
Power incident on sample | 0.302 | watts | 45 | 36 | 1.66 | 1.659 | 0.013026 | 0.031762 | 0.00015018
Power scattered | 0.59081 | microwatts | 50 | 41 | 1.68 | 1.679 | 0.013132 | 0.03202 | 0.0001514
Solid Angle | 0.00070031 | steradians | 55 | 46 | 1.03 | 1.029 | 0.0096934 | 0.023636 | 0.00011176
Power scattered |  |  | 60 | 51 | 0.63 | 0.629 | 0.0075774 | 0.018476 | 8.7362e-05
per watt*steradian | 0.0027935
BRDF for S-polarization
|  | Normal is ----> | Angle | Angle corrected | Voltage (mV) | Background corrected Voltage | Power on photodiode (uW) | Power scattered (uW) | BRDF
|  |  | 25 | 16 | 4 | 3.999 | 0.025405 | 0.061946 | 0.0002929
|  |  | 30 | 21 | 2.25 | 2.249 | 0.016147 | 0.039373 | 0.00018617
|  |  | 35 | 26 | 1.9 | 1.899 | 0.014296 | 0.034858 | 0.00016482
|  |  | 40 | 31 | 1.22 | 1.219 | 0.010699 | 0.026087 | 0.00012335
|  |  | 45 | 36 | 0.79 | 0.789 | 0.0084238 | 0.02054 | 9.712e-05
|  |  | 50 | 41 | 0.63 | 0.629 | 0.0075774 | 0.018476 | 8.7362e-05
|  |  | 55 | 46 | 0.38 | 0.379 | 0.0062549 | 0.015252 | 7.2115e-05
|  |  | 60 | 51 | 0.25 | 0.249 | 0.0055672 | 0.013575 | 6.4186e-05

## SS #4
Optics calibration |  |  | Filters for power calibration
Before | 0.356 | W | ND 1 transmittivity | 0.04 | Filters used
After | 0.292 | W | ND2 transmittivity | 0.08 | ND = 2.0
Transmissivity | 0.82022 |  | Before | 0.01546 | ND = 1.0
|  |  | After | 2.149e-06
Beam splitter calibration |  |  | Transmissivity | 0.000139
Before |  | mW | Total transmissivity | 0.0032
After |  | mW
Transmissivity | 0.5
Total transmissivity | 0.41011
Power meter diode calibration
Diode voltage (uV) | Calibrated power meter (uW)
0.966 | 15.2
1.644 | 8.6
2.142 | 2.7
4.073 | 0.41
3.01 | 0.321
3.674 | 0.373
Enter Voltage | Power hitting sample in W:
3.77 | 0.38429
Lock-in Calibration
Lock-in with filters (mV) | Power stick (uW) | Power hitting diode due to filter (uW)
68 | 116 | 0.3712
122 | 201 | 0.6432
162 | 270 | 0.864
27.6 | 46 | 0.1472
9.1 | 16.4 | 0.05248
Enter Voltage (mV) | Power hitting filter is (uW)
45 | 0.2423
Solid Angle Calculation
Distance from
sample to aperture (cm) | 30
Aperture radius (cm) | 0.635
mirror hole radius (cm) | 0.3 |  | BRDF for P-polarization
Steradian is (cm^2) | 900
Steradian for mirror is (cm^2) | 400 |  | Angle | Angle corrected | Voltage (mV) | Background corrected Voltage | Power on photodiode (uW) | Power scattered (uW) | BRDF
Aperture area is (cm^2) | 1.2661 |  | 25 | 16 | 111 | 110.999 | 0.59143 | 1.4421 | 0.054623
Hole area is (cm^2) | 0.2826 |  | 30 | 21 | 102.1 | 102.099 | 0.54435 | 1.3273 | 0.050275
Solid Angle is: | 0.00070031 |  | 35 | 26 | 80.2 | 80.199 | 0.4285 | 1.0448 | 0.039575
|  |  | 40 | 31 | 78 | 77.999 | 0.41686 | 1.0165 | 0.0385
Power incident on sample | 0.0377 | watts | 45 | 36 | 54.2 | 54.199 | 0.29096 | 0.70947 | 0.026872
Power scattered | 0.59081 | microwatts | 50 | 41 | 45.7 | 45.699 | 0.246 | 0.59983 | 0.022719
Solid Angle | 0.00070031 | steradians | 55 | 46 | 36.8 | 36.799 | 0.19892 | 0.48503 | 0.018371
Power scattered |  |  | 60 | 51 | 31 | 30.999 | 0.16823 | 0.41022 | 0.015538
per watt*steradian | 0.022378
BRDF for S-polarization
|  | Normal is ----> | Angle | Angle corrected | Voltage (mV) | Background corrected Voltage | Power on photodiode (uW) | Power scattered (uW) | BRDF
|  |  | 25 | 16 | 78 | 77.999 | 0.41686 | 1.0165 | 0.0385
|  |  | 30 | 21 | 71 | 70.999 | 0.37983 | 0.92617 | 0.03508
|  |  | 35 | 26 | 59 | 58.999 | 0.31635 | 0.77139 | 0.029217
|  |  | 40 | 31 | 49 | 48.999 | 0.26345 | 0.6424 | 0.024332
|  |  | 45 | 36 | 40 | 39.999 | 0.21584 | 0.52631 | 0.019935
|  |  | 50 | 41 | 34 | 33.999 | 0.1841 | 0.44891 | 0.017003
|  |  | 55 | 46 | 29.5 | 29.499 | 0.1603 | 0.39087 | 0.014805
|  |  | 60 | 51 | 26.6 | 26.599 | 0.14496 | 0.35346 | 0.013388

## SS #3
Optics calibration |  |  | Filters for power calibration |  | Filters used
Before | 0.356 | W | ND 1 transmittivity | 0.04 | ND = 2.0
After | 0.292 | W | ND2 transmittivity | 0.08 | ND = 1.0
Transmissivity | 0.82022 |  | Before | 0.01546
|  |  | After | 2.149e-06
Beam splitter calibration |  |  | Transmissivity | 0.000139
Before |  | mW | Total transmissivity | 0.0032
After |  | mW
Transmissivity | 0.5
Total transmissivity | 0.41011
Power meter diode calibration
Diode voltage (uV) | Calibrated power meter (uW)
0.966 | 15.2
1.644 | 8.6
2.142 | 2.7
4.073 | 0.41
3.01 | 0.321
3.674 | 0.373
Enter Voltage | Power hitting sample in W:
3.77 | 0.38429
Lock-in Calibration
Lock-in with filters (mV) | Power stick (uW) | Power hitting diode due to filter (uW)
68 | 116 | 0.3712
122 | 201 | 0.6432
162 | 270 | 0.864
27.6 | 46 | 0.1472
9.1 | 16.4 | 0.05248
Enter Voltage (mV) | Power hitting filter is (uW)
45 | 0.2423
Solid Angle Calculation
Distance from
sample to aperture (cm) | 30
Aperture radius (cm) | 0.635
mirror hole radius (cm) | 0.3 |  | BRDF for P-polarization
Steradian is (cm^2) | 900
Steradian for mirror is (cm^2) | 400 |  | Angle | Angle corrected | Voltage (mV) | Background corrected Voltage | Power on photodiode (uW) | Power scattered (uW) | BRDF
Aperture area is (cm^2) | 1.2661 |  | 25 | 16 | 127 | 126.999 | 0.67607 | 1.6485 | 0.06244
Hole area is (cm^2) | 0.2826 |  | 30 | 21 | 82 | 81.999 | 0.43802 | 1.0681 | 0.040454
Solid Angle is: | 0.00070031 |  | 35 | 26 | 79 | 78.999 | 0.42215 | 1.0294 | 0.038989
|  |  | 40 | 31 | 47 | 46.999 | 0.25287 | 0.6166 | 0.023355
Power incident on sample | 0.0377 | watts | 45 | 36 | 46.3 | 46.299 | 0.24917 | 0.60757 | 0.023013
Power scattered | 0.59081 | microwatts | 50 | 41 | 35 | 34.999 | 0.18939 | 0.46181 | 0.017492
Solid Angle | 0.00070031 | steradians | 55 | 46 | 26.5 | 26.499 | 0.14443 | 0.35217 | 0.013339
Power scattered |  |  | 60 | 51 | 23 | 22.999 | 0.12591 | 0.30702 | 0.011629
per watt*steradian | 0.022378
BRDF for S-polarization
|  | Normal is ----> | Angle | Angle corrected | Voltage (mV) | Background corrected Voltage | Power on photodiode (uW) | Power scattered (uW) | BRDF
|  |  | 25 | 16 | 81 | 80.999 | 0.43273 | 1.0552 | 0.039966
|  |  | 30 | 21 | 61.5 | 61.499 | 0.32958 | 0.80363 | 0.030439
|  |  | 35 | 26 | 51.8 | 51.799 | 0.27827 | 0.67851 | 0.0257
|  |  | 40 | 31 | 37.2 | 37.199 | 0.20103 | 0.49019 | 0.018567
|  |  | 45 | 36 | 32.7 | 32.699 | 0.17723 | 0.43214 | 0.016368
|  |  | 50 | 41 | 26.2 | 26.199 | 0.14284 | 0.3483 | 0.013192
|  |  | 55 | 46 | 21.5 | 21.499 | 0.11798 | 0.28768 | 0.010896
|  |  | 60 | 51 | 15.2 | 15.199 | 0.084653 | 0.20641 | 0.0078182

## SS #2
Optics calibration |  |  | Filters for power calibration
Before | 0.356 | W | ND 1 transmittivity | 0.04 | Filters used
After | 0.292 | W | ND2 transmittivity | 0.08 | ND = 2.0
Transmissivity | 0.82022 |  | Before | 0.01546 | ND = 1.0
|  |  | After | 2.149e-06
Beam splitter calibration |  |  | Transmissivity | 0.000139
Before |  | mW | Total transmissivity | 0.0032
After |  | mW
Transmissivity | 0.5
Total transmissivity | 0.41011
Power meter diode calibration
Diode voltage (uV) | Calibrated power meter (uW)
0.966 | 15.2
1.644 | 8.6
2.142 | 2.7
4.073 | 0.41
3.01 | 0.321
3.674 | 0.373
Enter Voltage | Power hitting sample in W:
3.77 | 0.38429
Lock-in Calibration
Lock-in with filters (mV) | Power stick (uW) | Power hitting diode due to filter (uW)
68 | 116 | 0.3712
122 | 201 | 0.6432
162 | 270 | 0.864
27.6 | 46 | 0.1472
9.1 | 16.4 | 0.05248
Enter Voltage (mV) | Power hitting filter is (uW)
45 | 0.2423
Solid Angle Calculation
Distance from
sample to aperture (cm) | 30
Aperture radius (cm) | 0.635
mirror hole radius (cm) | 0.3 |  | BRDF for P-polarization
Steradian is (cm^2) | 900
Steradian for mirror is (cm^2) | 400 |  | Angle | Angle corrected | Voltage (mV) | Background corrected Voltage | Power on photodiode (uW) | Power scattered (uW) | BRDF
Aperture area is (cm^2) | 1.2661 |  | 25 | 16 | 103.3 | 103.299 | 0.5507 | 1.3428 | 0.050861
Hole area is (cm^2) | 0.2826 |  | 30 | 21 | 87 | 86.999 | 0.46447 | 1.1326 | 0.042897
Solid Angle is: | 0.00070031 |  | 35 | 26 | 75 | 74.999 | 0.40099 | 0.97777 | 0.037034
|  |  | 40 | 31 | 66.5 | 66.499 | 0.35603 | 0.86813 | 0.032882
Power incident on sample | 0.0377 | watts | 45 | 36 | 43.7 | 43.699 | 0.23542 | 0.57403 | 0.021742
Power scattered | 0.59081 | microwatts | 50 | 41 | 37.6 | 37.599 | 0.20315 | 0.49535 | 0.018762
Solid Angle | 0.00070031 | steradians | 55 | 46 | 27.6 | 27.599 | 0.15025 | 0.36636 | 0.013876
Power scattered |  |  | 60 | 51 | 30.7 | 30.699 | 0.16665 | 0.40635 | 0.015391
per watt*steradian | 0.022378
BRDF for S-polarization
|  | Normal is ----> | Angle | Angle corrected | Voltage (mV) | Background corrected Voltage | Power on photodiode (uW) | Power scattered (uW) | BRDF
|  |  | 25 | 16 | 80 | 79.999 | 0.42744 | 1.0423 | 0.039477
|  |  | 30 | 21 | 62 | 61.999 | 0.33222 | 0.81008 | 0.030683
|  |  | 35 | 26 | 50 | 49.999 | 0.26874 | 0.6553 | 0.02482
|  |  | 40 | 31 | 43.3 | 43.299 | 0.2333 | 0.56887 | 0.021547
|  |  | 45 | 36 | 31.6 | 31.599 | 0.17141 | 0.41796 | 0.015831
|  |  | 50 | 41 | 25.6 | 25.599 | 0.13967 | 0.34056 | 0.012899
|  |  | 55 | 46 | 22.8 | 22.799 | 0.12486 | 0.30445 | 0.011531
|  |  | 60 | 51 | 22.1 | 22.099 | 0.12115 | 0.29542 | 0.011189

## SS #1
Optics calibration |  |  | Filters for power calibration
Before | 0.356 | W | ND 1 transmittivity | 0.04 | Filters used
After | 0.292 | W | ND2 transmittivity | 0.08 | ND = 2.0
Transmissivity | 0.82022 |  | Before | 0.01546 | ND = 1.0
|  |  | After | 2.149e-06
Beam splitter calibration |  |  | Transmissivity | 0.000139
Before |  | mW | Total transmissivity | 0.0032
After |  | mW
Transmissivity | 0.5
Total transmissivity | 0.41011
Power meter diode calibration
Diode voltage (uV) | Calibrated power meter (uW)
0.966 | 15.2
1.644 | 8.6
2.142 | 2.7
4.073 | 0.41
3.01 | 0.321
3.674 | 0.373
Enter Voltage | Power hitting sample in W:
3.77 | 0.38429
Lock-in Calibration
Lock-in with filters (mV) | Power stick (uW) | Power hitting diode due to filter (uW)
68 | 116 | 0.3712
122 | 201 | 0.6432
162 | 270 | 0.864
27.6 | 46 | 0.1472
9.1 | 16.4 | 0.05248
Enter Voltage (mV) | Power hitting filter is (uW)
45 | 0.2423
Solid Angle Calculation
Distance from
sample to aperture (cm) | 30
Aperture radius (cm) | 0.635
mirror hole radius (cm) | 0.3 |  | BRDF for P-polarization
Steradian is (cm^2) | 900
Steradian for mirror is (cm^2) | 400 |  | Angle | Angle corrected | Voltage (mV) | Background corrected Voltage | Power on photodiode (uW) | Power scattered (uW) | BRDF
Aperture area is (cm^2) | 1.2661 |  | 25 | 16 | 84 | 83.999 | 0.4486 | 1.0939 | 0.041432
Hole area is (cm^2) | 0.2826 |  | 30 | 21 | 59 | 58.999 | 0.31635 | 0.77139 | 0.029217
Solid Angle is: | 0.00070031 |  | 35 | 26 | 46 | 45.999 | 0.24758 | 0.6037 | 0.022866
|  |  | 40 | 31 | 38 | 37.999 | 0.20526 | 0.50051 | 0.018958
Power incident on sample | 0.0377 | watts | 45 | 36 | 39 | 38.999 | 0.21055 | 0.51341 | 0.019446
Power scattered | 0.59081 | microwatts | 50 | 41 | 37.6 | 37.599 | 0.20315 | 0.49535 | 0.018762
Solid Angle | 0.00070031 | steradians | 55 | 46 | 23.7 | 23.699 | 0.12962 | 0.31605 | 0.011971
Power scattered |  |  | 60 | 51 | 24.5 | 24.499 | 0.13385 | 0.32637 | 0.012362
per watt*steradian | 0.022378
BRDF for S-polarization
|  | Normal is ----> | Angle | Angle corrected | Voltage (mV) | Background corrected Voltage | Power on photodiode (uW) | Power scattered (uW) | BRDF
|  |  | 25 | 16 | 67 | 66.999 | 0.35867 | 0.87458 | 0.033126
|  |  | 30 | 21 | 44.8 | 44.799 | 0.24124 | 0.58822 | 0.02228
|  |  | 35 | 26 | 36.2 | 36.199 | 0.19574 | 0.47729 | 0.018078
|  |  | 40 | 31 | 28.5 | 28.499 | 0.15501 | 0.37797 | 0.014316
|  |  | 45 | 36 | 24.3 | 24.299 | 0.13279 | 0.32379 | 0.012264
|  |  | 50 | 41 | 24.6 | 24.599 | 0.13438 | 0.32766 | 0.012411
|  |  | 55 | 46 | 14.7 | 14.699 | 0.082008 | 0.19996 | 0.0075739
|  |  | 60 | 51 | 14 | 13.999 | 0.078305 | 0.19093 | 0.0072319

## Brewster Window
Optics calibration |  |  | Filters for power calibration
Before | 0.356 | W | ND 1 transmittivity | 0.04
After | 0.292 | W | ND2 transmittivity | 0.08
Transmissivity | 0.82022 |  | Before | 0.01546
|  |  | After | 2.149e-06
Beam splitter calibration |  |  | Transmissivity | 0.000139
Before |  | mW | Total transmissivity | 4.4481e-07
After |  | mW
Transmissivity | 0.5
Total transmissivity | 0.41011
Power meter diode calibration
Diode voltage (uV) | Calibrated power meter (uW)
0.966 | 15.2
1.644 | 8.6
2.142 | 2.7
4.073 | 0.41
3.01 | 0.321
3.674 | 0.373
Enter Voltage | Power hitting sample in W:
3.77 | 0.38429
Lock-in Calibration
Lock-in with filters (uV) | Power stick (uW) | Power hitting diode due to filter (uW)
24 | 345 | 0.00015346
4.3 | 65.8 | 2.9269e-05
11 | 164 | 7.2949e-05
3 | 44.5 | 1.9794e-05
1.5 | 18.7 | 8.318e-06
Enter Voltage (uV) | Power hitting filter is (uW)
71 | 0.00047708
Solid Angle Calculation
Distance from
sample to aperture (cm) | 30
Aperture radius (cm) | 0.635
mirror hole radius (cm) | 0.3 |  | BRDF for P-polarization
Steradian is (cm^2) | 900
Steradian for mirror is (cm^2) | 400 |  | Angle | Angle corrected | Voltage (uV) | Background corrected Voltage | Power on photodiode (uW) | Power scattered (uW) | BRDF
Aperture area is (cm^2) | 1.2661 |  | 34 | 54 | 7 | 6 | 3.7749e-05 | 9.2046e-05 | 3.4139e-07
Hole area is (cm^2) | 0.2826 |  | 36 | 56 | 7 | 6 | 3.7749e-05 | 9.2046e-05 | 3.4139e-07
Solid Angle is: | 0.00070031 |  | 40 | 60 | 17 | 16 | 0.00010175 | 0.0002481 | 9.2019e-07
|  |  | 32 | 52 | 8.3 | 7.3 | 4.6069e-05 | 0.00011233 | 4.1664e-07
Power incident on sample | 0.385 | watts | 30 | 50 | 10 | 9 | 5.6949e-05 | 0.00013886 | 5.1503e-07
Power scattered | 0.0011633 | microwatts | 25 | 45 | 6 | 5 | 3.1349e-05 | 7.644e-05 | 2.8351e-07
Solid Angle | 0.00070031 | steradians | 20 | 40 | 23.5 | 22.5 | 0.00014335 | 0.00034954 | 1.2964e-06
Power scattered |  |  | 15 | 35 | 10 | 9 | 5.6949e-05 | 0.00013886 | 5.1503e-07
per watt*steradian | 4.3146e-06 |  | 10 | 30 | 16 | 15 | 9.5349e-05 | 0.00023249 | 8.6231e-07
|  |  | 5 | 25 | 12 | 11 | 6.9749e-05 | 0.00017007 | 6.3079e-07
|  |  | 0 | 20 | 20 | 19 | 0.00012095 | 0.00029492 | 1.0938e-06
|  | Normal is ----> | 355 | 15 | 11 | 10 | 6.3349e-05 | 0.00015447 | 5.7291e-07
|  |  | 350 | 10 | 22 | 21 | 0.00013375 | 0.00032613 | 1.2096e-06
BRDF for S-polarization
|  |  | Angle | Angle corrected | Voltage (uV) |  | Power on photodiode (uW) | Power scattered (uW) | BRDF
|  |  | 40 | 60 | 24 | 23 | 0.00014655 | 0.00035734 | 1.3254e-06
|  |  | 30 | 50 | 14 | 13 | 8.2549e-05 | 0.00020128 | 7.4655e-07
|  |  | 20 | 40 | 16 | 15 | 9.5349e-05 | 0.00023249 | 8.6231e-07
|  |  | 10 | 30 | 13 | 12 | 7.6149e-05 | 0.00018568 | 6.8867e-07
|  |  | 0 | 20 | 15 | 14 | 8.8949e-05 | 0.00021689 | 8.0443e-07
|  |  | 350 | 10 | 16 | 15 | 9.5349e-05 | 0.00023249 | 8.6231e-07

## Lambertian full profile
Optics calibration |  |  | Filters calibration
Before | 0.356 | W | ND 1 transmissivity | 0.08
After | 0.292 | W | ND2 transmissivity | 0.043
Transmissivity | 0.82022 |  | Before | 0.01546
|  |  | After | 2.149e-06
Beam splitter calibration |  |  | Transmissivity reflective filter | 0.000139
Before | 358 | mW | Total transmissivity | 4.7817e-07
After | 187 | mW
Transmissivity | 0.52235 |  | Attenuation in front of photodiode
Total transmissivity | 0.42844 |  | 1.112e-05 | ND=1 and reflective filter
Power meter diode calibration
Diode voltage (uV) | Calibrated power meter (uW)
0.966 | 15.2
1.644 | 8.6
2.142 | 2.7
4.073 | 0.41
3.01 | 0.321
3.674 | 0.373
Enter Voltage | Power hitting sample in W:
3.77 | 0.38429
THEORETICAL BRDF
Lock-in Calibration | [10 empty] | Angle |  | Radians | BRDF
Lock-in with filters (uV) | Power stick (uW) | Power hitting diode due to filter (uW) | [8 empty] | 0 | 0 | 0 | 0.31831
39 | 560 | 0.00026778 | [8 empty] | 1.8 | 1 | 0.031416 | 0.31815
24 | 345 | 0.00016497 | [8 empty] | 3.6 | 2 | 0.062832 | 0.31768
4.3 | 65.8 | 3.1464e-05 | [8 empty] | 5.4 | 3 | 0.094248 | 0.3169
3 | 44.5 | 2.1279e-05 | [8 empty] | 7.2 | 4 | 0.12566 | 0.3158
1.5 | 18.7 | 8.9418e-06 | [8 empty] | 9 | 5 | 0.15708 | 0.31439
94 | 1350 | 0.00064553 | [8 empty] | 10.8 | 6 | 0.1885 | 0.31267
Enter Voltage (uV) | Power hitting photodiode is (uW) | [9 empty] | 12.6 | 7 | 0.21991 | 0.31064
23 | 0.00015812 |  | 1.5812e-10 | [7 empty] | 14.4 | 8 | 0.25133 | 0.30831
[11 empty] | 16.2 | 9 | 0.28274 | 0.30567
Solid Angle Calculation | [10 empty] | 18 | 10 | 0.31416 | 0.30273
Distance from | [10 empty] | 19.8 | 11 | 0.34558 | 0.29949
sample to aperture (cm) | 30 | [9 empty] | 21.6 | 12 | 0.37699 | 0.29596
Aperture radius (cm) | 0.635 | [9 empty] | 23.4 | 13 | 0.40841 | 0.29213
mirror hole radius (cm) | 0.3 | [9 empty] | 25.2 | 14 | 0.43982 | 0.28802
Steradian is (cm^2) | 900 | [9 empty] | 27 | 15 | 0.47124 | 0.28362
Steradian for mirror is (cm^2) | 400 | [9 empty] | 28.8 | 16 | 0.50265 | 0.27894
Aperture area is (cm^2) | 1.2661 |  |  | Angle | Angle corrected | Voltage (uV) | Background corrected Voltage | Power on photodiode (uW) | Power scattered (uW) | BRDF | 30.6 | 17 | 0.53407 | 0.27398
Hole area is (cm^2) | 0.2826 |  |  | 0 | 0 | 59 | 57 | 0.00039136 | 82.1434 | 0.38458 | 32.4 | 18 | 0.56549 | 0.26876
Solid Angle is: | 0.00070031 |  |  | 10 | 10 | 55 | 53 | 0.00036392 | 76.384 | 0.35761 | 34.2 | 19 | 0.5969 | 0.26327
|  |  |  | 20 | 20 | 53 | 51 | 0.0003502 | 73.5043 | 0.34413 | 36 | 20 | 0.62832 | 0.25752
Power incident on sample | 0.305 | watts | 63.9212 | 30 | 30 | 51 | 49 | 0.00033648 | 70.6246 | 0.33065 | 37.8 | 21 | 0.65973 | 0.25151
Power scattered | 33.1886 | microwatts |  | 40 | 40 | 46 | 44 | 0.00030218 | 63.4254 | 0.29694 | 39.6 | 22 | 0.69115 | 0.24526
Solid Angle | 0.00070031 | steradians |  | 50 | 50 | 39 | 37 | 0.00025416 | 53.3465 | 0.24976 | 41.4 | 23 | 0.72257 | 0.23877
Power scattered |  |  |  | 60 | 60 | 32 | 30 | 0.00020614 | 43.2676 | 0.20257 | 43.2 | 24 | 0.75398 | 0.23204
per watt*steradian | 0.15538 |  | 0.31847 | 70 | 70 | 23 | 21 | 0.0001444 | 30.3089 | 0.1419 | 45 | 25 | 0.7854 | 0.22508
[11 empty] | 46.8 | 26 | 0.81681 | 0.2179
[11 empty] | 48.6 | 27 | 0.84823 | 0.2105
[11 empty] | 50.4 | 28 | 0.87965 | 0.2029
[11 empty] | 52.2 | 29 | 0.91106 | 0.19509
[5 empty] | MEASURED BRDF FOR LAMBERTIAN |  | Data calculated using slightly larger solid angle |  |  |  | 54 | 30 | 0.94248 | 0.1871
[5 empty] | 0 | 0.39811 | 0.3 |  |  |  | 55.8 | 31 | 0.97389 | 0.17892
[5 empty] | 10 | 0.37114 | 0.279 |  |  |  | 57.6 | 32 | 1.0053 | 0.17056
[5 empty] | 20 | 0.35766 | 0.269 |  |  |  | 59.4 | 33 | 1.0367 | 0.16203
[5 empty] | 30 | 0.34418 | 0.259 |  |  |  | 61.2 | 34 | 1.0681 | 0.15335
[5 empty] | 40 | 0.31048 | 0.234 |  |  |  | 63 | 35 | 1.0996 | 0.14451
[5 empty] | 50 | 0.26329 | 0.198 |  |  |  | 64.8 | 36 | 1.131 | 0.13553
[5 empty] | 60 | 0.2161 | 0.163 |  |  |  | 66.6 | 37 | 1.1624 | 0.12642
[5 empty] | 70 | 0.15543 | 0.117 |  |  |  | 68.4 | 38 | 1.1938 | 0.11718
[11 empty] | 70.2 | 39 | 1.2252 | 0.10782
[11 empty] | 72 | 40 | 1.2566 | 0.098363
[11 empty] | 73.8 | 41 | 1.2881 | 0.088806
[11 empty] | 75.6 | 42 | 1.3195 | 0.07916
[11 empty] | 77.4 | 43 | 1.3509 | 0.069437
[11 empty] | 79.2 | 44 | 1.3823 | 0.059645
[11 empty] | 81 | 45 | 1.4137 | 0.049795
[11 empty] | 82.8 | 46 | 1.4451 | 0.039895
[11 empty] | 84.6 | 47 | 1.4765 | 0.029956
[11 empty] | 86.4 | 48 | 1.508 | 0.019987
[11 empty] | 88.2 | 49 | 1.5394 | 0.0099984
[11 empty] | 90 | 50 | 1.5708 | 1.9499e-17

## Black glass 1
Optics calibration |  |  | Filters calibration
Before | 0.356 | W | ND 1 transmittivity | 0.04
After | 0.292 | W | ND2 transmittivity | 0.08
Transmissivity | 0.82022 |  | Before | 0.01546
|  |  | After | 2.149e-06
Beam splitter calibration |  |  | Transmissivity | 0.000139
Before | 2.13 | mW | Total transmissivity | 4.4481e-07
After | 1.25 | mW
Transmissivity | 0.58685
Total transmissivity | 0.48135
Power meter diode calibration
Diode voltage (uV) | Calibrated power meter (uW)
0.966 | 15.2
1.644 | 8.6
2.142 | 2.7
4.073 | 0.41
3.01 | 0.321
3.674 | 0.373
Enter Voltage | Power hitting sample in W:
3.77 | 0.38429
Lock-in Calibration
Lock-in with filters (uV) | Power stick (uW) | Power hitting diode due to filter (uW)
4.3 | 65.8 | 2.9269e-05
11 | 164 | 7.2949e-05
3 | 44.5 | 1.9794e-05
1.5 | 18.7 | 8.318e-06
Enter Voltage (uV) | Power hitting filter is (uW)
1.5 | 9.35e-06
Solid Angle Calculation
Distance from
sample to aperture (cm) | 30
Aperture radius (cm) | 0.635
mirror hole radius (cm) | 0.3
Steradian is (cm^2) | 900
Steradian for mirror is (cm^2) | 400 |  |  | Angle | Voltage | Noise corrected | BRDF
Aperture area is (cm^2) | 1.2661 |  | 25 | 70.6 | 9.8 | 8.9 | 4.8743e-07
Hole area is (cm^2) | 0.2826 |  | 20 | 65.6 | 3.1 | 2.2 | 1.1587e-07
Solid Angle is: | 0.00070031 |  | 16 | 61.6 | 2.3 | 1.4 | 7.15e-08
|  |  | 14 | 59.6 | 2.2 | 1.3 | 6.6e-08
Power incident on sample | 0.36 | watts | 12.6 | 58.2 | 2.4 | 1.5 | 7.7e-08
Power scattered | 1.9424e-05 | microwatts | 8 | 53.6 | 2.4 | 1.5 | 7.7e-08
Solid Angle | 0.00070031 | steradians | 0 | 45.6 | 1.6 | 0.7 | 3.27e-08
Power scattered |  |  | -10 | 35.6 | 1.5 | 0.6 | 2.71e-08
per watt*steradian | 7.7047e-08 |  | -20 | 25.6 | 1.7 | 0.8 | 3.82e-08
|  |  | -25 | 20.6 | 1.9 | 1 | 4.93e-08
|  |  | -30 | 15.6 | 2.4 | 1.5 | 7.7e-08
|  | Normal is ----> | -45.6

## Black Glass 2
Optics calibration |  |  | Filters for power calibration
Before | 0.356 | W | ND 1 transmittivity | 0.04 | Filters used
After | 0.292 | W | ND2 transmittivity | 0.08 | ND = 2.0
Transmissivity | 0.82022 |  | Before | 0.01546 | ND = 1.0
|  |  | After | 2.149e-06 | Reflective
Beam splitter calibration |  |  | Transmissivity | 0.000139
Before |  | mW | Total transmissivity | 4.4481e-07
After |  | mW
Transmissivity | 0.5 |  | Power meter filter calibration
Total transmissivity | 0.41011 |  | Before | 1.01e-05
|  |  | After | 0.00588
Power meter diode calibration
Diode voltage (uV) | Calibrated power meter (uW)
0.966 | 15.2
1.644 | 8.6
2.142 | 2.7
4.073 | 0.41
3.01 | 0.321
3.674 | 0.373
Enter Voltage | Power hitting sample in W:
3.77 | 0.38429
|  |  | Transmissivity | 0.0017177
Lock-in Calibration |  |  | Power measured | 0.000518
Lock-in with filters (uV) | Power stick (uW) | Power hitting diode due to filter (uW) | Actual Power | 0.30157
24 | 345 | 0.00015346
4.3 | 65.8 | 2.9269e-05
11 | 164 | 7.2949e-05
3 | 44.5 | 1.9794e-05
1.5 | 18.7 | 8.318e-06
Enter Voltage (uV) | Power hitting filter is (uW)
10 | 6.4651e-05
Solid Angle Calculation
Distance from
sample to aperture (cm) | 30
Aperture radius (cm) | 0.635
mirror hole radius (cm) | 0.3 |  | BRDF for P-polarization
Steradian is (cm^2) | 900
Steradian for mirror is (cm^2) | 400 |  | Angle | Angle corrected | Voltage (uV) | Background corrected Voltage (uV) | Power on photodiode (uW) | Power scattered (uW) | BRDF
Aperture area is (cm^2) | 1.2661 |  | 25 | 15 | 8.2 | 7.2 | 4.6731e-05 | 0.00011395 | 4.3389e-07
Hole area is (cm^2) | 0.2826 |  | 30 | 20 | 4.6 | 3.6 | 2.3691e-05 | 5.7767e-05 | 2.1997e-07
Solid Angle is: | 0.00070031 |  | 35 | 25 | 4.7 | 3.7 | 2.4331e-05 | 5.9328e-05 | 2.2591e-07
|  |  | 40 | 30 | 3.7 | 2.7 | 1.7931e-05 | 4.3722e-05 | 1.6649e-07
Power incident on sample | 0.375 | watts | 45 | 35 | 5 | 4 | 2.6251e-05 | 6.4009e-05 | 2.4374e-07
Power scattered | 0.00015764 | microwatts | 50 | 40 | 5.2 | 4.2 | 2.7531e-05 | 6.713e-05 | 2.5562e-07
Solid Angle | 0.00070031 | steradians | 55 | 45 | 5.5 | 4.5 | 2.9451e-05 | 7.1812e-05 | 2.7345e-07
Power scattered |  |  | 60 | 50 | 6 | 5 | 3.2651e-05 | 7.9615e-05 | 3.0316e-07
per watt*steradian | 6.0028e-07 |  | 65 | 55 | 7.6 | 6.6 | 4.2891e-05 | 0.00010458 | 3.9824e-07
BRDF for S-polarization
Normal is ----> 10 degrees
|  |  | Angle | Angle corrected | Voltage (uV) | Background corrected Voltage (uV) | Power on photodiode (uW) | Power scattered (uW) | BRDF
|  |  | 25 | 15 | 4.6 | 3.6 | 2.3691e-05 | 5.7767e-05 | 2.1997e-07
|  |  | 30 | 20 | 3.5 | 2.5 | 1.6651e-05 | 4.0601e-05 | 1.546e-07
|  |  | 35 | 25 | 3 | 2 | 1.3451e-05 | 3.2798e-05 | 1.2489e-07
|  |  | 40 | 30 | 2.6 | 1.6 | 1.0891e-05 | 2.6556e-05 | 1.0112e-07
|  |  | 45 | 35 | 3 | 2 | 1.3451e-05 | 3.2798e-05 | 1.2489e-07
|  |  | 50 | 40 | 3.6 | 2.6 | 1.7291e-05 | 4.2162e-05 | 1.6055e-07
|  |  | 55 | 45 | 2.9 | 1.9 | 1.2811e-05 | 3.1238e-05 | 1.1895e-07
|  |  | 60 | 50 | 2.8 | 1.8 | 1.2171e-05 | 2.9677e-05 | 1.1301e-07