## Slide 1: First Lessons from the Advanced LIGO Integration Testing
May 20, 2013
Daniel Sigg
LIGO Hanford Observatory
GWDAW, Elba, 2013

## Slide 2: Advanced LIGO Sensitivity
Standard quantum limit
Initial LIGO
Radiation pressure noise
Shot noise

## Slide 3: The Advanced LIGO Detector
Laser
Laser:
200W - 1064nm
Input mode cleaner:
stabilizes frequency
and cleans laser mode
Arm cavities:
Fabry-Perrot cavities store light to effectively increase length
Homodyne Readout
Signal recycling mirror: amplifies readout signal
Input
Test
Mass
End
Test
Mass
4 km
4 km
800kW
800kW
Power recycling mirror: reflects light back coming from the beam splitter, increasing power in the arm cavities

## Slide 4: Global Timeline
October 2010: Hand-off of Observatories to Advanced LIGO for installation
February 2012: Both Observatories have decommissioned initial LIGO detectors, started in-vacuum Installation and subsystem Integration
April 2012: Recommendation to the NSF to place one interferometer in India
Aug 2014: LLO ‘L1’ Interferometer accepted (internal plan date)
Sep 2014 LHO ‘H1’ interferometer accepted (internal plan date)
LHO ‘H2’ detector was on schedule to be accepted in March 2014, but instead will go to India pending NSF Approval
Mar 2015: Data Analysis computer system completed, planned Project end

## Slide 5: Sequence of Installation and Integration Testing
Install
PSL/
IO table
Install
IMC
Install
DRMI
Full interferometer
L1
Squeezing Test
H1
Squeezer install
Feb 2014
Install
PSL/
IO table
Install
IMC
Install
(non-project, risk reduction)
Install
H2
Single arm cavity
(H2 OAT)
PSL/IO table
Install
IMC
Rec’ld vertex MICH
Full interferometer
Deinstall/Store 3rd Interferometer
HIFO-X
HIFO-Y
Oct 2013
completed
HIFO-Y
HIFO-X
DRMI
Full interfere.
completed
Feb 2014
Oct 2010

## Slide 6: Subsystem Testing
Subsystem testing
  All components are tested before installation
  All subsystems have a test and verification phase before commissioning
  Paid off big time: Much faster startup of commissioning
PSL: Accepted and working
  Lasers delivers stable 180W, currently running at 35W
  Excess frequency, intensity and jitter noise due to water cooling flow (avoid 90 degree turns)
  Lifetime of laser diodes factor 2 below specs of manufacturer
  Unknown contamination reduced AR coating performance of PMC tank windows, windows could be cleaned, since tank is open no accumulation of stuff anymore

## Slide 7: Core Optics Coatings
ETM spiral pattern generates scattered ring
  Back scatter from beam tube baffles can effect <30Hz sensitivity
Spherical aberration acceptable (two ETMs are nearly identical)
Arm Cavity Loss is within budget (50 ppm achieved vs. <75 ppm spec)

## Slide 8: One Arm Test Components
BSC ISI
HEPI
HEPI
Quad SUS
TrMon SUS
ACB
Phase locked to fiber-delivered PSL light
1064 / 532 nm NPRO
BSC ISI
HEPI
HEPI
Quad SUS
ITM
Triple SUS
FM
ACB
ETM
TM
LVEA:
No PSL
Test Mass optical lever (not shown)
Y-End
Arm Cavity length & alignment sensing

## Slide 9: One Arm Test Summary and Actions
Verified the basic functionality of many subsystems:
  Two-stage active seismic isolation system (BSC ISI)
  Quadruple suspension
  Initial Alignment system/procedure
  Thermal compensation ring heater
  Green beam cavity locking
Actions:
  ALS wavefront sensors eliminated from design: alignment sufficiently stable
  PZT steering control of ALS beam incorporated into design
  Additional hardware was identified to support automation
  Usability of various systems needs to be improved to be accessible to non-experts

## Slide 10: Intermediate and Quantitative Goals of the One Arm Test
Initial alignment: Sustained flashes of optical resonance in the arm cavity | Achieved, within one week of operation
Cavity locking/ISC: Green laser locked to cavity for 10 minutes or more | 
TransMon/ALS: Active beam pointing error on the TransMon table below 1 urad rms in angle and below 100 um rms in transverse motion | Achieved
Calibration: ETM displacement calibration at the 20% level | Achieved
Thermal Compensation: Ring heater wavefront distortion, measured by Hartmann sensor, in agreement with model at the 10 nm rms level | Achieved
Optical levers: Long term drift below 1 urad | Diurnal motion about twice this level, possibly actual test mass motion

## Slide 11: Intermediate and Quantitative Goals of the One Arm Test
Controls/SUS: Decoupling of length-to-angle drive of the quad suspension | Achieved for TOP stage
Seismic isolation: Relative motion between two SEI platforms below 250 nm rms (w/o global feedback) | Achieved
Cavity alignment fluctuations: Relative alignment fluctuations below 100 nrad rms | Achieved
Controls/ISC: Long term cavity locking; fully automated locking sequence | Long term locking achieved; automation was rudimentary
Cavity length control: Relative test mass longitudinal motion below 10 nm rms | Not possible to assess with OAT. These have become objectives for the HIFO-Y test.
ALS: Ability to control frequency offset between 1064 nm and 532 nm resonances at the 10 Hz level | 
ALS: Relative stability of the 1064/532 nm resonances at the 10 Hz level | 

## Slide 12: IMC Test
Ifo reflected port in-vac readout
Input optics in-air optical table 1
Input optics in-air optical table 2

## Slide 13: IMC Test Summary and Actions
Locking was as easy and reliable as expected
  Seismic isolation controls for HAM ISI are straightforward
  Angular stability quite good; wavefront sensor alignment control only for long term drifts
  New design for low-noise Voltage-Controlled Oscillator validated
  No major problems with high power operation
Issues and actions
  Excess laser noise (frequency, amplitude, beam pointing). No show stoppers, but room for improvement (some already made)
  PSL Intensity servo (outer stage) found to need re-engineering
  Absorption in IMC mirrors. Two of the three mirrors found to absorb 2 ppm, vs 0.6-0.7 ppm nominal – relevant to contamination control

## Slide 14: Intermediate and Quantitative Goals of the IMC Test
IMC availability: Locked duty cycle of >90% | Achieved, would remain locked indefinitely
Mean lock duration: > 4 hours | Achieved
Optical efficiency: Transmission from PSL output to Interferometer input (O-PRM), > 75% | Achieved, 86%
IMC visibility: > 95% (include mode-matching) | Achieved, 97-98%
IMC length/frequency control bandwidth: Goal of 40 kHz or higher | Achieved, 60 kHz
IMC frequency/length crossover: ~10 Hz | Achieved
IMC transmitted power stability: relative rms fluctuations of 1% or less | Achieved, 0.5% RIN
Pointing stability: angular motion of transmitted beam, < 1.6 urad rms | Achieved, 0.4 urad
Intensity noise: transmitted light RIN <10-7/Hz1/2 | Not achieved
Faraday isolation: > 30 dB at full power | Not measured in-situ

## Slide 15
So far so good!