## Slide 1: aLIGO BSFM “Level 2” Damping Loop Design (Supplemental to LHO aLOG 6392)
J. Kissel

## Slide 2
aLIGO BSFM "Level 2" Damping Loop Design
Mission Statement
The damping loops installed during the SUS testing phase
merely to prove that the suspensions could be damped
damped quickly and robustly
little-to-no regard to re-injection of sensor noise
very aggressive, but poorly placed elliptic filters to rolloff noise
The mission here was to design a set of loops, that
doesn’t take you years to design and tweak
isn’t on the hairy edge of instability
doesn’t require any “Brett Shapiro” trickery (damping in Modal, Global bases)
doesn’t require and new infrastructure (which Modal and Global damping would),
but still
designed with what modeling experience we’ve gained
gets us close to what we’ll need for aLIGO, primarily focusing on Longitudinal
will be sufficient for the first several stages of integrated testing
Level 1
Level 2

## Slide 3
Stability: Bode plots of Open and Closed Loop Gain Transfer Functions
Cross-Coupling: The above, Modeled both as SISO and MIMO systems, the below as MIMO
Modeled Performance: Compute all DOF’s of Top Mass sensor noise contribution to Optic degree of freedom of interest
Compare: with other noise sources, requirements, coupling to DARM, etc.
Measured Performance: With what we can: Closed Loop TOP2TOP Open and Closed Loop TFs, TOP Sensor ASDs, TOP Control Signal ASDs
  (Check out G1300537 and LLO aLOG 6949 for a more thorough description; I assume from here on that you’ve seen and understand what they mean, so I can get right to the points.)
Damping Loop Design Model Figures of Merit

## Slide 4: BSFM vs QUAD
Important differences between BSFM and QUAD:
BSFM is a long, triple suspension, so (L/P) and (T/R) mode frequencies are, in general, lower
  Plenty more phase to play with between the highest frequency resonance and 10 [Hz] requirements
BSFM’s lower blades are aligned with the T - V plane, so there’s no fundamental coupling between (L/P) and (T/R) like there is for the QUADs
  Don’t have to consider sensor noise of 4 different loops to improve L
Though BSFM (L and P) or (T and R) are (independently) coupled, respectively, damping a resonance in one DOF damps it in both
  Can again play the “take advantage of the MIMO” game to relax the design where needed.
  E.g. Highest to T/R modes at 2.1 and 3.2 [Hz] can be damped in R, where there are no noise requirements
These three points mean that the L, T, R, and P loops are significantly easier to design and meet requirements (not to mention the requirements are less stringent, of course)
M1
TOP
UIM
+T
+L
PUM
M2
M1’s principle axes of inertia are aligned with Euler Basis
TOP and UIM’s principle axes of inertia NOT aligned with Euler Basis

## Slide 5: Damping Loop Design Stabililty (New Filters -- L)
I started out by using the same general design features as the Level 2 QUAD filters,
but
because resonances are lower, I had more phase with which to play.
So
Elliptic filters start at lower frequency
Qs of boosts elliptic notches are a little higher
More dBs of isolation on elliptic’s stop-band
Unconditionally Stable
Elliptic first, then rolloff
Boost on low-frequency modes to reduce RMS
Not-so-high-Q Elliptic Filter
Not just velocity damping anymore:
[z:p]=[~1:~10] Hz pair for more phase with which to play

## Slide 6: Damping Loop Design Compare with other Noises (New Filters -- L)
Residual Seismic Noise
Dominates below 0.8 Hz
M2 Actuator Noise
Dominates above 9 Hz
M1 BOSEM Sensor Noise only dominates between 0.8-9 Hz, and is well out of the way before detection band!
And sensor noise easily comes in a factor of 10 below the requirements at 10 Hz.
but it’s not all puppies and rainbows…

## Slide 7: Damping Loop Design Coupling to DARM (OLD Filters)
Even with the Level 1 filters, we see it’s
VERTICAL that’s the worst offender when it comes to contribution to DARM

## Slide 8: Damping Loop Design Stabililty (New Filters -- V)
M2 and M2 response to M1 force shows the highest V mode at 17.5 [Hz]. But M1 to M1 doesn’t.
BSFM is pretty stiff in vertical, so all modes are relatively high compared to 10 [Hz] requirement, so
  Boost for extra gain at lowest-mode has to be high in frequency
  Very little phase left with which to play
But must roll of loop gain really fast *and* get it extra low around 17 [Hz]
Solution:
Forced to use more complex filter design (i.e. 4th order elliptic instead of 3rd order)
Get 2nd elliptic notch as close to 17.5 Hz as possible

## Slide 9: Damping Loop Design Compare with other Noises (New Filters -- V)
Resulting design is a compromise between:
Loop Stability
Resulting Qs of 1st and 2nd V modes
Sensor noise level between 10 Hz and 3rd V mode
More gradual rolloff at high-frequency
Misses V requirement for BSFMs by at most a factor of 5.
But…

## Slide 10: Damping Loop Design Coupling to DARM (NEW Filters)
… assuming a V2L coupling factor of 0.001, the resulting total VERTICAL noise is still well below any of the expected DARM sensitivities!

## Slide 11: Level 1 vs.Level 2

## Slide 12: Level 1 vs.Level 2
Can’t measure Optic noise improvement directly yet…
If sensor noise dominates above ~0.5 [Hz] both before and after, then M1 control signal should be ~100 times less at 10 Hz.
And it is!

## Slide 13: Concluding Remarks
BSFM Level 2 damping filters beat almost all aLIGO requirements
BSFM design, in general was easier than the QUAD; Vertical is the toughest
Design choices
  Chose to absorb over all gain into boost filter
  Chose to move boost filters up in frequency on some DOFs
Proof of design measurements
  I didn’t measure the open loop gain transfer functions
    the experience with the QUAD has shown
      the measurements to be more confusing than they’re worth
      the measurements confirm the cross-coupled MIMO model works
    So in the interest of time, I’ve nixed them
  Will get closed loop transfer functions and spectra over the course of phase 3a testing

## Slide 14: Bonus Material For the Curious
Remember, for even more text, plots, details see LHO aLOG 6392

## Slide 15: Level 1 vs.Level 2

## Slide 16: Seismic Input Motion
In the absence of real, best-possible performance data from the BSC-ISIs, there are a few choices for Residual Ground Input Motion to the QUAD:
- Use the Requirements for all DOFs
- Use M. Evans’ Model of the “Translation” (same for X, Y, Z) and “Rotational” (same RX, RY, RZ)
- Use not-yet-awesome, but real H2OAT data (different for every DOF, and even between ISIs)
We know every degree of freedom will be different, so I don’t like using the Reqs.
We know the very low frequency data of the H2OAT data is all tilt (if not sensor noise), and we know the mid-frequency band will be better
So I went with Matt’s Model since it seems to fold in the most (optimistic?) realism

## Slide 17: Damping Loop Design MIMO Games (New Filters – T/R)
Highest two T/R modes at 2.1 and 3.2 Hz would be trouble if T plant was SISO
Turns out they can be damped in R
So, push up the boost frequency in R (and subsequently the elliptic cutoff frequency), “skipping” the first three T/R modes, since they don’t couple well to the R plant
Then ignore them in the T design, since R has squashed them below unity gain
Higher frequency boost
Higher frequency elliptic cutoff
Stable by a large margin
2.1 and 3.2 Hz modes don’t cross unity
but still get damped
Stable by a large margin