## PRM-02 checked 30Aug2013
DCC: T1300740-v1
PRM-02 ROC Tolerances
R. Martin and G. Billingsley
5 Sep. 2013 |  |  | Average Files |  |  | Random check of individual files |  |  | Effect of ROC-TS uncertainty |  |  | Effect of Gap |  | Effect of Zernike Terms
|  |  | 0 deg (arrow up) | 90 deg (arrow right) | 225 deg | 11m_PRM-02-0_006 | 11m_PRM-02-90_061 | 11m_PRM-02-225_047 | 0 deg ROC_TS ave | 0 deg ROC_TS max | 0 deg ROC_TS min | +2 mm | -2 mm | Z[3]max=0.002778 | Z[3]min=0.001839
Radius of curvature of TS | Assumes converging TS and cvx part | m | -11.4933 | -11.4933 | -11.4933 | -11.4933 | -11.4933 | -11.4933 | -11.4933 | -11.5058 | -11.4808 | -11.4933 | -11.4933 | -11.4933 | -11.4933
Gap between TS and part to be measured | Vertex to vertex, no tilt | m | 0.48962 | 0.48962 | 0.48962 | 0.48962 | 0.48962 | 0.48962 | 0.48962 | 0.48962 | 0.48962 | 0.49162 | 0.48762 | 0.48962 | 0.48962
Zernike power coefficient | Assumes "Fringe" Zernikes | nm | 2.4153 | 2.0769 | 2.4951 | 1.8897 | 1.9078 | 2.5791 | 2.511 | 2.511 | 2.511 | 2.511 | 2.511 | 2.9558 | 1.9567
Part measured aperture |  | mm | 30.324 | 30.324 | 30.324 | 30.324 | 30.324 | 30.324 | 30 | 30 | 30 | 30.324 | 30.324 | 30.324 | 30.324
Wavefront "sag" at part plus power |  | mm | -0.010441 | -0.010442 | -0.010441 | -0.010442 | -0.010442 | -0.010441 | -0.010219 | -0.010207 | -0.01023 | -0.010443 | -0.010439 | -0.01044 | -0.010442
[14 empty] | 22 | 25
Part radius |  | m | -11.0088 | -11.0081 | -11.0089 | -11.0077 | -11.0077 | -11.0091 | -11.0091 | -11.0216 | -10.9966 | -11.007 | -11.011 | -11.0099 | -11.0078
|  |  | -11.0086 |  |  | -11.0082 |  |  | Diff. (mm): | -12.5123 | 12.5123 | 2.1169 | -1.8869 | -0.82299 | 1.2839
Specified Radius |  | m | -11 | [8 empty] | -15.2222 | mm
|  | (+/-0.11) | -10.89 | [8 empty] | 15.9131
-11.11
Comments:
-Measured on Sep 5-6, 2011 at CIT by R. Martin and G. Billingsley with the Wyko interferometer
- Wave and small central features are from the Test Sphere (C1203095-v1)
One indication that ROC of PRM-02 is significantly different than PRM-01 (measured one after another on Sep 5, 2011), is that the Gap had to be changed from 50.5 cm (PRM-01) to 48.5 cm (PRM-02).