## Slide 1: LIGO Characterization ITM09 AR “milky” appearance

## Slide 2: ITM09 AR Dark Field Examination
Circular; some are not resolvable by our microscope so ~1 µm or smaller
Larger defects, ~ 2 to 4 micrometers appeared with a density of roughly 30/mm2
Smaller defects: density from 50x103/mm2 to 150x103/mm2
There were no areas observed to be free of defects.
The uncoated edge was not examined.

## Slide 3: AR Reflection

## Slide 4: Looking for a correlation with the milky pattern and AR Reflection

## Slide 5: AR Absorption

## Slide 6: Scatter
We had no schedule opportunity to measure scatter
We have arranged with Joshua Smith to use his imaging system to measure the scatter when ITM06 arrives.

## Slide 7: For the Record HR Absorption

## Slide 8: For the Record HR Transmission