## ROC calculator
|  | Min | Avg | Max | Avg+Stdev
| Radius of curvature of TS | 2100 | 2100 | 2100 | 2100 | m | Assumes diverging TS and concave part
| Gap between TS and part to be measured | 0.08 | 0.08 | 0.08 | 0.08 | m | Vertex to vertex, no tilt
| Measured Power | 126.32 | 126.989 | 127.38 | 127.168 | nm | Assumes "Fringe" Zernikes
| Standard deviation in power (from avg) |  | 0.179
| Part measured aperture | 160.4 | 160.4 | 160.4 | 160.4 | mm | Enter Data into shaded cells
| range from Process stats |  | 1.061
| Wavefront "sag" at part plus power | 0.0016577 | 0.0016584 | 0.0016588 | 0.0016585 | mm
| Part radius | 1940.0498 | 1939.2671 | 1938.81 | 1939.0578 | m
| total spread (m) |  | -1.2398
| environmental uncertainty (m) |  | -0.2093
Power stats
160mm diameter
Useful numbers
DCC number |  | Aperture diameter | ROC
C1102035 | R36 Reference Sphere | 260 | 35.751
C1102247 | R34 Reference Sphere | 260 | 33.74
C1002794 | Test Mass Reference Sphere | all | 2100
Power Trend
300mm diameter
Transmitted Wavefront
|  | Min | Avg | Max | Avg+Stdev
| Radius of curvature of TS | 2100 | 2100 | 2100 | 2100
| Gap between TS and part to be measured | 0.08 | 0.08 | 0.08 | 0.08
| Measured Power | -3976.55 | -3982 | -3987.45 | -3979.3 |  | 160mm diameter plot
| Standard deviation in power (from avg) |  | 2.7
| Part measured aperture | 160.4 | 160.4 | 160.4 | 160.4
| range from Process stats |  | 10.9
| Wavefront "sag" at part plus power | -0.0024452 | -0.0024506 | -0.0024561 | -0.0024479
| Part radius | -1315.254 | -1312.329 | -1309.4169 | -1313.7765
| total spread (m) |  | 5.8371
| environmental uncertainty (m) |  | -1.4475
300mm

## Pixel Sizes
|  | Magnification | Full Aperture size
Full | 0.4 | 1 | 409.3617
2x | 0.192 | 2.0833 | 196.4936
small | 0.047 | 8.5106 | 48.1
s2x | 0.023 | 17.3913 | 23.5383

## power fit
Find a corresponding power term over different aperture
INPUT
d1 - initial (mm OR pix) | 200.823 | 160 mm aperture, in pixels
d2 - final (mm OR pix) | 401.622 | full aperture size in pixels
Zernike1 (nm) | -2012.725 | Focus fit for 160 mm aperture
OUTPUT
Zernike2 (nm) | -8049.9379 | Focus term to remove over full aperture
For Best fit on tilt and piston subtract 2x the residual
first use tilt values from full aperture fit
then subtract the generated file from the raw data - note the residual on 160mm
in this case I added 2X the positive tilt residual to the negative full aperture tilt