LIGO Laboratory / LIGO Scientific Collaboration

LIGO-E1400033-v1	Advanced LIGO	1/21/14

Test Procedure for IQ Demodulator Filtering

Yuta Michimura

Distribution of this document:

LIGO Scientific Collaboration

This is an internal working note

of the LIGO Laboratory.

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 159 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

http://www.ligo.caltech.edu/

# Introduction

This test procedure describes the test of proper filtering of the Quad I&Q Demodulator Module (D0902796) modified with RF splitter and filter based on E1300899. Basic operation of the Quad I&Q Demodulator Module should be tested beforehand by following the procedure described in T1100062.

# Test Equipment

- Function generator (e.g. Tektronix AFG3102)

- Oscilloscope

# Test

1) Apply LO input from the rear of the chassis and RF input to Ch1 or Ch3 (D0902745 with a low pass filter PLP-21.4+). Set the amplitude of LO and RF input to 10 dBm and 3 dBm, respectively. Set the frequency of LO and RF input to 18.2 MHz and 18.201 MHz, respectively. Read the peak to peak amplitude of 1 kHz components of I-Mon and Q-Mon output using an oscilloscope and check that they are within the typical values. Repeat the same thing with RF input frequency of 91.001 MHz.

| LO frequency | RFin frequency | Ch1/Ch3 I-Mon | Ch1/Ch3 Q-Mon | Typ. |
| --- | --- | --- | --- | --- |
| 18.2 MHz | 18.201 MHz |  |  | 4.8 ±0.5 V |
| 18.2 MHz | 91.001 MHz |  |  | < 15 mV |

2) Repeat 1) for Ch2 or Ch4 (D0902745 with a high pass filter PHP-100+) with LO frequency of 91 MHz.

| LO frequency | RFin frequency | Ch2/Ch4 I-Mon | Ch2/Ch4 Q-Mon | Typ. |
| --- | --- | --- | --- | --- |
| 91 MHz | 18.201 MHz |  |  | < 14 mV |
| 91 MHz | 91.001 MHz |  |  | 4.3±0.4 V |