## Slide 1: TCS installation & results at LLO
Update: Systems Telecon
8th Jan 2014
Aidan Brooks (for the TCS Team)
G1400018-v1

## Slide 2: CO2 Laser Status
Enclosures installed
Lasers plumbed
Phase I (central heating) optics to be installed in January

## Slide 3: HWS cartoon layout
Hartmann in-air table
  Next to HAM4
Probe beams injected
  Into main optical axis
  Retro-reflected from ITMX and ITMY
  Accumulated wavefront distortion is readout by HWS on return beam
In-vacuum optics in HAM4
Must match in-vacuum and in-air optical axes

## Slide 4: HWS installation
All optics installed
In-air and in-vacuum optics aligned separately
In-air and in-vacuum axes mated
  Took some effort
  Currently out of alignment following December vent
Calibrated
  HWS lever arm
  Magnification
In-air optics
In-vacuum optics

## Slide 5: Injected HWS-X beam BS baffle is limiting aperture
Earthquake stop
HWS-X beam

## Slide 6: Transmitted HWS-X beam Mostly clean
Edge of baffles
Earthquake stop shadow
Diffraction from dust on an optic

## Slide 7: With Hartmann Plate HWS-Y

## Slide 8: Stray reflections from in-vacuum optics
We see stray reflections from optics within HAM4
These can be mostly spatially filtered out.
Example of unfiltered reflection is shown here
Possibly require tweaking alignment of in-vacuum optics at a later date
  If determined that these introduce significant noise

## Slide 9: HWS-Y background noise measurements
Shows background defocus over time
  Equivalent Power Absorbed ~ 1.2mW
  0.3% of maximum expected
  Fluctuation is correlated with temperature variation by HWS
  Expect to be able to remove some of temperature fluctuations

## Slide 10: Recall: Indirect Ring Heater measurement Measured 1064nm beam size

## Slide 11: Directly observed RH thermal lens
Step 1: Apply ring heater
Step 2: See thermal lens form at glacial speed
Still need to characterize this lens. Need to measure astigmatism.
http://youtu.be/GcoqCn8Ytds
HWS beam
ITM
Ring heater
No data points outside ring
-105 mm 0mm +105mm
X coordinate

## Slide 12: HWS beam position on test mass
Located the center of the test mass
  Relative to the center of the HWS beam
  Used the ring heater thermal lens
HWS beam
ITM
Ring heater

## Slide 13: Appendix: HWS alignment
Located return beams by swinging SR3
Identified 2 reflections from ITM and 2 from CP
Once correct beam is identified:
  Misalign SR3 to return beam to HWS
  Adjust HWS input periscope until no misalignment is required on SR3