LIGO Laboratory / LIGO Scientific Collaboration

LIGO-E1300865-v1 S1400080	Advanced LIGO	9/17/12

Test Procedure for RF diplexer amplifier

Koji Arai, Daniel Sigg

Distribution of this document:

LIGO Scientific Collaboration

This is an internal working note

of the LIGO Laboratory.

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 159 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

http://www.ligo.caltech.edu/

# Introduction

The following Test Procedure describes the test of proper operation of the RF diplexer amplifier (E1300852). The unused outputs should always be properly terminated.

# Test Equipment

- Voltmeter

- Oscilloscope

- Agilent Network/Spectrum analyzer (e.g. AG4395A)

- Agilent ESG-1000 signal generator (or similar signal source with fmax>150MHz)

- RF Power Meter (e.g. HP E4418A or N1914A)

- Board Schematics--D1300989

# Tests

The RF Preamplifier comes with the Low Noise Power Module (D0901846.

- Verify the proper current draw.  Using a bench DC supply apply ±24Volts to P7 and 
±17 Volts to P6 of the low noise power Module (D0901846).  Measure the current draw of the board.

+24 Volt current ________0.02A__	0.02 A Nom.

–24 Volt current  ________0.02A_	0.02 A Nom.

+17 Volt current ________0.30A_	0.2~0.3 A Nom.

–17  Volt current _______ 0.03A_	0.03 A Nom.

- On the low noise power module check the voltage on TP 1-13.

TP1 ( +17V ) ___+17.27_V____		TP2 (-17V )    ___-17.27 V______

TP3 , 4 ( GND )				TP5 (+ 5V)    ____+5.05 V_______

TP6 (-15V ) ____-14.97 V_____		TP7 (+24V ) ___+24.17 V__________

TP8 ( GND )				TP9 ( -24V ) ___-24.09 V__________

TP10 ( GND )				TP11 (+15V ) __+14.98 V__________

TP12 (+VREF ) ___+9.99V_____	TP13  (-VREF ) ______-9.99V______

- If  TP 1 , 2, 7 , 9 and 8 are correct then pin 5 on U1 and U7,  TP14 (OK) should be

Logic high ~3Volts. The front panel LED should be on.

Confirm.________OK__________

- The noise on TP 12, 13, 11 and 6 should be measured with a SR785 using an

rms power spectrum.

TP12  noise __8.38 nVrms/sqrtHz___ less than 200 nVrms/sqrt Hz at 140 Hz

TP13 noise  __17.13 nVrms/sqrtHz___less than 900 nVrms/sqrt Hz at 140 Hz

TP11 noise ___13.14 nVrms/sqrtHz___ less than 1 uVrms/sqrt Hz at 140 Hz

TP6 noise	 ___22.3 nVrms/sqrtHz___ less than 1 uVrms/sqrt Hz at 140 Hz.

- Test the output powers of each port by applying the RF signal to the “In” port at the below specified power and frequency. With an RF power meter measure the power at the three output ports (i.e. “Direct”, “LF”, HF). Compare the output with the specification.

Frequency: 27.3MHz, Input power -20dBm

| Output port | Measured Power | Specification (dBm) | Judgement (Pass/NG) |
| --- | --- | --- | --- |
| Direct | -32.16 dBm | -32 +/- 1 dBm | Pass |
| LF | -0.59 dBm | > -2 dBm | Pass |
| HF | -29.67 dBm | < -27 dBm | Pass |

Frequency: 136.5MHz, Input power -40dBm

| Output port | Measured Power (dBm) | Specification (dBm) | Judgement (Pass/NG) |
| --- | --- | --- | --- |
| Direct | -52.43 dBm | -52 +/- 1 dBm | Pass |
| LF | Upper limit: < -87 dBm dBm | < -72† dBm | Pass |
| HF | 8.80 dBm | > +3 dBm | Pass |

† This power level may be too small to be measured with the power meter. An RF spectrum analyzer may be needed.

- Measure -0.5dB pass and -50dB rejection bandwidth by changing the input frequency from the above setting. Change the signal frequency, starting from the above specified signal frequency. Record the upper and lower frequency where the output is reduced by 0.5dB and 50dB. Confirm the pass and rejection band with the specification.

LF port, center frequency: 27.3MHz, Input power -20dBm

| Lower boundary of the -0.5dB band | Upper boundary of the -0.5dB band | Pass band width df=f0.5dBHI - f0.5dBLO | Spec. | Judgement (Pass/NG) |
| --- | --- | --- | --- | --- |
| f0.5dBLO = 26.4 MHz | f0.5dBHI = 28.3 MHz | df = 1.9 MHz | df > 1MHz | Pass |

| Lower boundary of the -50dB band | Upper boundary of the -50dB band | Rejection band specification | Judgement (Pass/NG) |
| --- | --- | --- | --- |
| f50dBLO= 10.6 MHz | f50dBHI= 41 MHz | f50dBLO > 18.2MHz f50dBLO < 45.5MHz | Marginal pass; lower -50dB boundary is too low |

HF port, center frequency: 135.6MHz, Input power -40dBm

| Lower boundary of the -0.5dB band | Upper boundary of the -0.5dB band | Pass band width df=f0.5dBHI - f0.5dBLO | Spec. | Judgement (Pass/NG) |
| --- | --- | --- | --- | --- |
| f0.5dBLO = 134.7 MHz | f0.5dBHI = 139.9 MHz | df = 5.2 MHz | df > 2MHz | Pass Peak transmission at 137.5MHz |

| Lower boundary of the -50dB band | Upper boundary of the -50dB band | Rejection band specification | Judgement (Pass/NG) |
| --- | --- | --- | --- |
| f50dBLO= MHz Stops at -30dBm | f50dBHI= MHz Stops at -30dBm | f50dBLO > 127.4MHz f50dBLO < 145.6MHz | NG; rejection never reaches -50dB |

- Measure the input power level to give the harmonic distortion of -30dBc. Connect the LF or HF ports to an RF spectrum analyzer. Setup the harmonic distortion measurement (the function “Multi Peak search” is useful). Increase the signal amplitude with a 0.1dB step until the largest harmonics is 30dB relative to the main signal peak (“Delta Marker” is useful for this measurement). Record the power of the harmonic (up to 5th harmonics for 27.3MHz and 3rd harmonics for 136.5MHz). Confirm the output power level is larger than the specification.

LF port, Signal frequency: 27.3MHz

| Input power for 30dBc harmonics | Power level of the main signal | Specification | Judgement (Pass/NG) |
| --- | --- | --- | --- |
| -2.66 dBm | 16.1 dBm | > 16 dBm | Pass |

| Harmonics |
| --- |
| 2nd (54.6MHz) | 3rd (81.9MHz) | 4th (109.2MHz) | 5th (136.5MHz) |
| -24.6dBm, -40.7dBc | -13.0dBm, -29.1dBc | -25.5dBm, -41.6dBc | -20.3, -36.4dBc |

HF port, Signal frequency: 136.5MHz

| Input power for 30dBc harmonics | Power level of the main signal | Specification | Judgement (Pass/NG) |
| --- | --- | --- | --- |
| -32.2 dBm | 16.3 dBm | > 16 dBm | Pass |

| Harmonics |
| --- |
| 2nd (273.0MHz) | 3rd (409.5MHz) |  |  |
| -26.7dBm, -43dBc | -13.4dBm, -29.7dBc |  |  |

- Measure the output noise level of the LF and HF output ports. A

| Output Port | Measured noise level | Specification | Judgement (Pass/NG) |
| --- | --- | --- | --- |
| LF (27.3MHz) | 15 nV/rtHz | < 20 nV/rtHz |  |
| HF (136.5MHz) | 230 nV/rtHz | < 300 nV/rtHz |  |