## Slide 1: Hanford Contamination Control Update 2 Back to Basics
LIGO G1400378-v2 Contamination Control Working Group March 21, 2014

## Slide 2: Why do we care?
Sensitivity loss
Absorption
Scatter
Damage
Contamination will limit performance.
Notes: aLIGO optics are vulnerable to particulate contamination. Surface particles absorb and scatter power from the interferometer mode. Such losses limit recycling and signal cavity finesse, degrading shot-noise limited strain sensitivity.
ven low-level contaminants can result in laser damage to optics during the operation the interferometers, and/or the unacceptable reduction of overall detector sensitivity.

## Slide 3
Laser Induced Damage due to Particulate Contamination
Billingsley, Gushwa, Phelps, Torrie, Zhang
LVC meeting March 2014
LIGO-G1400209

## Slide 4
Contamination Control Working Group:
Analyze and mitigate contamination

## Slide 5: Contamination
Complex mix of hydrocarbons
Biological
Mixed silicates (soil)
Fibers
Metals
Contamination comes from us and everything we do.

## Slide 6: Understanding PCLs
Chart: Particle diameter (µm)
| 3.2866548386940648
65 | 1 | 10000000
100 | 1 | 10000000
150 | 1 | 10000000
300 | 1 | 10000000
500 | 1 | 10000000
750 | 1 | 10000000
1,000 | 1 | 10000000
424 | 1 | 10000000
65 | 1105.2038 | 389.9749 | 131.0516 | 10.5442 | 1
100 | 5058.2466 | 1784.8195 | 599.7911 | 48.2583 | 4.5768 | 1
150 | 24263.6176 | 8561.4999 | 2877.104 | 231.4877 | 21.954 | 4.7968 | 1
300 | 480901.4782 | 169687.7206 | 57023.7944 | 4588.0545 | 435.1247 | 95.0728 | 19.8199 | 1
log vs. log2 graph
424
Goal for vertical surfaces
Goal for horizontal surfaces
Scan area = 4.2 cm2
Total particles = 10,229
Notes: Sample T1300454

## Slide 7: Telling a story…

## Slide 8: LLO HAM2 History
Chart: 
| 1 | Pump down #1 (doors close) | Vent (doors open) | 1//2/3/4 | Pump down #2 | Vent | 2 | 3/4/5 | 5 | 6 | Pump down #3 | Vent | 6/7 | Pump down #4 | Vent | 7
POSITION | -50 | 30 | 20 | -50 | 10 | 10 | -50 | -50 | -50 | -50 | 25 | 20 | -50 | 50 | 40 | -50
DATE | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
First clean – Basic wipe down.
~ 1 hour total
Big clean – Multiple wipe downs over 2 week period.
Big clean – Wiped horizontal & vertical surfaces, inside of ISI, underneath support tubes, etc.
1-1.5 hrs per chamber
Pump down #1
(doors close)
Pump down #2
Vent (doors open)
Vent
Pump down #3
Vent
Pump down #4
Vent
1 cycle
Medium work
(PRM, ISS PD array, PR3 gluing)
Medium work
(PRM, ISS PD array, PR3 gluing)
1 cycle
Minimal work
1 cycle
Light work
(PR2/PR3 baffles)

## Slide 9: LLO HAM2 Story
Chart: 
| 1 | Pump down #1 (doors close) | Vent (doors open) | 1//2/3/4 | Pump down #2 | Vent | 2 | 3/4/5 | 5 | 6 | Pump down #3 | Vent | 6/7 | Pump down #4 | Vent | 7
POSITION | -50 | 30 | 20 | -50 | 10 | 10 | -50 | -50 | -50 | -50 | 25 | 20 | -50 | 50 | 40 | -50
DATE | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
Chart: Particle diameter (µm)
| 3.2866548386940648
65 | 1 | 10000000
100 | 1 | 10000000
150 | 1 | 10000000
300 | 1 | 10000000
500 | 1 | 10000000
750 | 1 | 10000000
1,000 | 1 | 10000000
1,500 | 1 | 10000000
65 | 1105.2038 | 389.9749 | 131.0516 | 10.5442 | 1
100 | 5058.2466 | 1784.8195 | 599.7911 | 48.2583 | 4.5768 | 1
150 | 24263.6176 | 8561.4999 | 2877.104 | 231.4877 | 21.954 | 4.7968 | 1
300 | 480901.4782 | 169687.7206 | 57023.7944 | 4588.0545 | 435.1247 | 95.0728 | 19.8199 | 1
Chart: Particle diameter (µm)
| 3.2866548386940648
65 | 1 | 10000000
100 | 1 | 10000000
150 | 1 | 10000000
300 | 1 | 10000000
500 | 1 | 10000000
750 | 1 | 10000000
1,000 | 1 | 10000000
1,500 | 1 | 10000000
65 | 1105.2038 | 389.9749 | 131.0516 | 10.5442 | 1
100 | 5058.2466 | 1784.8195 | 599.7911 | 48.2583 | 4.5768 | 1
150 | 24263.6176 | 8561.4999 | 2877.104 | 231.4877 | 21.954 | 4.7968 | 1
300 | 480901.4782 | 169687.7206 | 57023.7944 | 4588.0545 | 435.1247 | 95.0728 | 19.8199 | 1
Chart: Particle diameter (µm)
| 3.2866548386940648
65 | 1 | 10000000
100 | 1 | 10000000
150 | 1 | 10000000
300 | 1 | 10000000
500 | 1 | 10000000
750 | 1 | 10000000
1,000 | 1 | 10000000
1,500 | 1 | 10000000
65 | 1105.2038 | 389.9749 | 131.0516 | 10.5442 | 1
100 | 5058.2466 | 1784.8195 | 599.7911 | 48.2583 | 4.5768 | 1
150 | 24263.6176 | 8561.4999 | 2877.104 | 231.4877 | 21.954 | 4.7968 | 1
300 | 480901.4782 | 169687.7206 | 57023.7944 | 4588.0545 | 435.1247 | 95.0728 | 19.8199 | 1
In front of PR3
In front of PR3
817
280
In front of PR3
In front of PR3
218
268
Near PR3, N door (Vertical)
NW corner of table
N side of table (Vertical)
105
218
192
  Cleaning as you go works!

## Slide 10: LHO BSC Stories
Chart: Particle diameter (µm)
| 3.2866548386940648
65 | 1 | 10000000
100 | 1 | 10000000
150 | 1 | 10000000
300 | 1 | 10000000
500 | 1 | 10000000
750 | 1 | 10000000
1,000 | 1 | 10000000
1,500 | 1 | 10000000
65 | 1105.2038 | 389.9749 | 131.0516 | 17.1325 | 1
100 | 5058.2466 | 1784.8195 | 599.7911 | 78.411 | 4.5768 | 1
150 | 24263.6176 | 8561.4999 | 2877.104 | 376.1255 | 21.954 | 4.7968 | 1
300 | 480901.4782 | 169687.7206 | 57023.7944 | 7454.7542 | 435.1247 | 95.0728 | 19.8199 | 1
Feb – Oct 2013
1 cycle
Work in adjacent chambers
WBSC1 – Floor below ITMY
230
Oct – Nov 2013
1 cycle
No work?
WBSC3 – Floor below ITMX
316
April – Oct 2013
1 cycle
Heavy work over wafer
WBSC2 – Floor below baffle
598
  Need more samples to determine trends.

## Slide 11: Total Wafers Collected
Chart: Number of wafers
| HAM2 | HAM3 | HAM4 | BSC1 | BSC2 | BSC3 | BSC6 | Other
Hanford | 11 | 0 | 0 | 1 | 1 | 1 | 2 | 1
Livingston | 18 | 18 | 2 | 5 | 6 | 5 | 0 | 1
  Need more samples from Hanford.
  LIGO-T1300014: Aligo, BSC Flooring + HAM ISI, Witness Sample Placement Requirements

## Slide 12: WBSC10 – March 2014

## Slide 13
Contamination Control 101

## Slide 14: Scrubs
Why:
  At least ½ class improvement
How:
  Wear instead of street clothes
  No special laundering requirements
  Tailor if needed
When:
  Required for all cleanroom work
  Strongly encouraged for VEA work
* NOT a magical barrier against contamination! Refer to LIGO-E0900047 aLIGO Contamination Control Plan

## Slide 15: Glove Washing
OR
Why:
  Covered in particulates
How:
  High purity IPA + wipe
  Think peanut butter hands
When*:
  After donning
  If dirty
* Change if ripped or excessively dirty

## Slide 16: Clean as you go policy
Why:
  Everyone contributes to problem
  Everyone MUST contribute to solution
How (basic):
  Inspect
  Wipe
  Vacuum
When:
GROSS
CHAMBER
End of Work
Assembly
End of Assembly
Post-Install
Pre-Chamber Close
Pre-Install

## Slide 17: Contamination Control Supplies

## Slide 18
Common mistakes
What’s wrong with these pictures…

## Slide 20
As shown ready for door removal.
If in to e.g. "clean" or other then skinny bunny okay but only when door or door cover is on.

## Slide 26: Questions or Concerns?
Local CC Czar:
  Jeff Bartlett
Caltech:
  Calum Torrie
  Kate Gushwa

## Slide 27: References
Contamination Control on the DCC and Wiki
LIGO-E1300953 aLIGO, Particulate Contamination
https://dcc.ligo.org/wiki/index.php/Engineering_for_LIGO/ContaminationControlWorkingGroup
Requirements
LIGO-T1300511 Some thoughts regarding Particulate Contamination Requirements
Talks
LIGO-G1301249 Hanford and Livingston Contamination Control Update Dec 2013-Jan2014
LIGO-G1400142: Summary of 4" Wafer Data
LIGO-G1300777 Contamination Control Requirements – Gloves, Cleaning on the Go and The Plan
Key Documents
LIGO-E0900047 aLIGO Contamination Control Plan
LIGO-E1201035 aLIGO Chamber Entry & Exit Procedures.
LIGO-E1201096 Contamination Sample Handling – How to receive, use, send, buy and store samples.
  LIGO-T1300014: Aligo, BSC Flooring + HAM ISI, Witness Sample Placement Requirements
LIGO-L1300086 Contamination Control Guidance on Cleanroom Scrubs

## Slide 28: Backup slides

## Slide 29: Defining Clean
Cleanroom Class
Volume of air
ISO 5 (Class 100)
Particles per:
Requirement:
Tool:
Particulate Cleanliness Level (PCL)
Area of surface
Level 65 (optics at full power)
Level 100 (chambers)
Notes: Cleanroom Class:
What we have been using all along to determine cleanliness.
Surface Cleanliness Level:
What we should have been looking at (in addition to Cleanroom Class).
ONE of the tools, but not the only one.