## S6
S6 Run Downtime Statistics on H1
2009-09-19
http://ilog.ligo-wa.caltech.edu/ilog/pub/ilog.cgi?group=detector&date_to_view=09/19/2009&anchor_to_scroll_to=2009:09:19:16:47:19-jgarcia
Operating Mode | % time | Notes
Science Mode | 48%
hardware | 6% | hardware issues
software | 4% | reboots
maintenance day | 2%
seismic | 5% | Seismic time was given for each time seismic noise was responsible for breaking the ifo lock. If it was an earthquake, a longer time was given for ringdown. If a truck or other anthropogenic noise caused lock loss, just a few minutes (<30min) was given for ringdown.
locking | 16% | when the operator is actively trying to get the ifo to lock
commissioning | 9%
calibrations | 1%
alignments | 5% | IFO full alignments, downtime when wfs were tweaked, TCS align, or op lev alignments
unknown | 2%
other | 1%
wind | 1%
100%

## S5
S5 Run Downtime Statistics on L1
2007-10-05
last 2 months of S5 (8/2/2007 - 10/1/2007)
http://ilog.ligo-la.caltech.edu/ilog/pub/ilog.cgi?group=detector&date_to_view=10/05/2007&anchor_to_scroll_to=2007:10:05:12:44:04-dhoak
Operating Mode | % time | Notes
Science Mode | 76.1%
Hardware & Software Failures | 2.7% | The 'Hardware & Software Failures' category is due to the trouble with the laser chiller on Aug15, the RBS PZT failure on Aug31, and trouble with the HEPI pumps on Aug9.
Maintenance & Commissioning | 6.1%
Up/Down Scripts | 2.2%
Weather | 0.2%
Trains | 2.3%
Earthquakes | 6.5% | There were seven earthquakes in the past two months of magnitude 7.0 or greater. (Typically there's about one earthquake of this size or larger per month).
Unknown/Other | 3.8% | includes a small amount of downtime due to logging
99.9%

## T080160
S5 Downtime Statitsics (entire run) from T080160
Period: 11/14/2005 through 5/27/2007
Category | L1 | H1 | Comments
Uptime | 67.2% | 76.9%
Calibration | 0.5% | 0.5%
Hardware & Software Failures | 3.7% | 3.9%
Maintenance & Commissioning | 7.9% | 10.5%
maint., plus single events/mysteries |  | 6.2%
commissioning |  | 4.3%
Up/Down Scripts | 2.7% | 2%
Wind & Storms | 1.6% | 2%
Seismic | 8.6% | 2.2%
(Microseism) | 2.3% | --
(Trains) | 2.3% | --
(Earthquakes) | 4% | --
Unknown | 2.6% | 2%
SEC Construction | 4.1% | --
ITMY | 1% | --
| 99.9% | 100%

## aLIGO
Category |  | S5 H1 (T080160) | S5 L1 * (T080160) | S5 L1 ** last 2 months | aLIGO Estimation | aLIGO comments
Uptime |  | 76.9% | 70.89% | 76.1% | 78.55%
Calibration |  | 0.5% | 0.5274% | 0.5% | 0.5% | same as for S5
Hardware Failures |  | 3.9% | 3.903% | 2.7% | 5.065% | Rate for aLIGO is based on failure rates (from T1300519, Sparing Analysis) for prompt (unscheduled) attention and LLO power outage rate for last year with recovery WAG
Software Failures |  |  |  |  | 0.5% | WAG; should be low after commissioning
Maintenance |  | 6.2% | 8.333% | 6.1% | 8.128% | corresponds to 4 hr/wk, plus one 3 wk vent&pump period per yr Note that for aLIGO the ISC in-vacuum output electronics/detectors are in vacuum volumes separate from the main vacuum volume
Commissioning |  | 4.3% |  |  | 0% | projection to steady-state operations; no commissioning
Transition to Low Noise Operation State (Up/Down Scripts) |  | 2% | 2.848% | 2.2% | 1.053% | 15 min every 24 hr. with Guardian automation system (less rms motion and better automation)
Wind |  | 2% | 1.688% | 0.2% | 2% | assume that we don't improve on sensitivity to wind-induced ground tilt senstivity
Seismic | microseism | 2.2% | 2.426% |  | 2.2% | should have better isolation from earthquakes and microseism assume better for L1, but same as H1 S5
| trains |  | 2.426% | 2.3%
| earthquakes |  | 4.219% | 6.5%
Unknown |  | 2% | 2.743% | 3.4% | 2% | same as H1 S5
|  | 100% | 100% | 100% | 100%
* renormalized to remove the exceptional time on L1 for SEC construction and ITMy repair due to SEC construction
** calibration time assumed to be part of the unknown/other category. Unusually high earthquake activity in this period, but no microseism induced downtime reported.

## Maintenance & Repair
# | Subsys | Event | Frequency (#/yr) | Duration (hrs each) | Downtime | Comments
1 | All | In-vacuum maintenance | 1 | 504 | 5.753% | once/year vent, parallel activities; 1 wk vent plus 2 wk pumpdown
2 | All | weekly maintenance period | 52 | 4 | 2.374% | All deferred/planned maintenance
8.128%
# | Subsys | Maintenance/Repair | Frequency (#/yr) | Duration (hrs each) | Downtime | Event | Comments
1 | COC | Clean optics accumulated hydrocarbon film off of optics | 1 | 8 | 0.09132% | 1 | duration is set by vertex optics set ETMs cleaned (if needed) in parallel
2 | DAQ | electronics failures, incl. all of the I/O chassis and real-time/DAQ computers used by aLIGO | 154 | 1 | 1.758% | NA | individual, separate events; no early detection
3 | IO | clean or replace degraded opto-electronics on the PSL/IO optics table | 8.19 | 4 | 0.374% | 2 | clean or replace when convenient during weekly maintenance period
4 | IO | replace, align and test EOM when it fails | 0.365 | 4 | 0.01667% | NA
5 | IO | replace, align and test slow & fast PDs when they fail | 0.526 | 4 | 0.02402% | NA
6 | IO | replace, align and test HP beam dump when it fails | 0.4 | 4 | 0.01826% | NA
7 | IO | replace/clean, align and test degraded HAM2 & HAM3 optics | 15 | 24 | 4.11% | 1 | very pessimistic MTBF estimates for degradation of the HAM2/3 IO optics; should revise based on iLIGO/eLIGO experience
8 | IO | replace/clean, align and test degraded IO table optics | 13.538 | 4 | 0.6182% | 2
9 | ISC | replace, align & test ISC in-vacuum electronics | 0.04 | 24 | 0.01096% | NA | HAM1 & HAM6 can be vented & pumped down separately from the main vacuum and turned around in 1 day (3 shifts) if necessary
10 | ISC | replace, align & test ISC in-air electronics | 6.973 | 4 | 0.3184% | NA
11 | ISC | replace, align & test ISC in-vacuum optomechanics, picomotors, detectors | 1.802 | 4 | 0.08228% | NA
12 | ISC | replace, align & test ISC in-air optomechanics, picomotors, detectors | 8.635 | 4 | 0.3943% | NA
13 | OptLev | replace, align & test diode laser, optical fiber or Quad PD | 6.526 | 8 | 0.596% | 2 | Optical levers are fly-wheels & diagnostic equipment, not critical to operation after commissioning (no downtime hit)
14 | PSL | repairs of failures without early detection: FrontEnd: EOM, AOM, shutter, PDs, pump fibers, NPRO driver | 0.7 | 4 | 0.03196% | NA
15 | PSL | scheduled repair/replacement of degraded components: FrontEnd: Faraday Isolator, Amplifier, Diode set, NPRO | 1.814 | 4 | 0.08283% | 2
16 | PSL | repairs of failures without early detection: HPO: PD Amp, PD int, PD Iso, PD BP, lock PD, Heads | 4.13 | 4 | 0.1886% | NA
17 | PSL | scheduled repair/replacement of degraded components: HPO: Diode set, etc. | 4.904 | 4 | 0.2239% | 2
18 | PSL | Diagnostic Bread Board (DBB) | NA | NA |  | NA | diagnostic function only
19 | PSL | ISS, PMC, FSS | 1.9 | 2 | 0.04338% | NA
20 | SEI | in-vacuum sensors | 13.193 | 16 | 2.41% | 1 | ISI can perform (albeit degraded) when a sensor fails due to redundancy (control algorithms/laws and software for re-configuration for graceful recovery are still in development)
21 | SEI | SEI repairs of failures without early detection: prompt attention to in-air electronics | 12.554 | 2 | 0.2866% | NA
22 | SEI | SEI scheduled repair/replacement or maintenance | 223.083 | 2 | 5.093% | 2
23 | SLC | NA |  |  | 0% |  | no maintenance or credible expected failures other than PDs, but they are not essential to operation
24 | SUS | in-vacuum OSEM repair/replacement | 1.839 | 4 | 0.08397% | 1 | Can sustain a single OSEM failure on a stage generally (some redundancy), so schedule replacement in yearly vent
25 | SUS | SEI repairs of failures without early detection: prompt attention to in-air electronics | 10.022 | 2 | 0.2288% | NA
26 | TCS | in-vacuum CO2 laser beam relay mirror coating failures | 2.608 | 8 | 0.2382% | 1 | pessimistic MTBF for CO2 beam coating life note that RH glass former is not a credible failure & RH RTD is not essential
27 | TCS | in-air TCS prompt repair | 15.516 | 4 | 0.7085% | NA | NOTE HWS IS NOT ESSENTIAL
28 | TCS | in-air TCS scheduled reapir | 3.827 | 4 | 0.1747% | 2 | NOTE HWS IS NOT ESSENTIAL
29 | TMS | in-vacuum | 0.876 | 4 | 0.04% | 1
30 | TMS | in-air, prompt | 0.291 | 2 | 0.006644% | NA
31 | Power | power outage | 16 | 5.1875 | 0.9475% | NA | based on LLO for the period 6/1/2013 thru 6/1/2014 time to recover is a WAG
19.2%
[5 empty] | 6.973% | event 1 - in-vacuum maintenance
[5 empty] | 7.163% | event 2 - weekly maintenance |  | 3.0167
[5 empty] | 5.065% | NA - prompt attention

## LLO Power Outages
LLO Power Outages in the last year (6/1/2013 through 6/1/2014)
as reported in the LLO elog entries
Date | Duration | elog # | duration of impact (Hr)
2013-06-16 | 2 sec | 7445 | 4
2013-07-06 | > 1hr | 7860 | 5
2013-07-11 | 1 hr | 7860 | 5
2013-07-13 | 6 sec | 7847 | 4
2013-07-16 | glitch | 9761 | 4
2013-08-30 | 2 sec | 8544 | 4
2013-09-13 | 6 sec | 8736 | 4
2013-11-24 |  | 9877 | 4
2014-01-06 | 21 min | 10397 | 4.5
2014-03-19 | 14.5 hr | 11579 | 20
2014-04-07 |  | 11933 | 4
2014-04-13 | 14 min | 12055 | 4.5
2014-04-14 | 17 sec | 12090 | 4
2014-04-14 | seconds | 12090 | 4
2014-04-14 | seconds | 12096 | 4
2014-04-22 | 17 sec | 12235 | 4
| # events | 16 | 83
|  |  | 5.1875 | avg

## LHO Power Outages
LHO Power Outages in the last year (6/1/2013 through 6/1/2014)
as reported in the LHO elog entries
Date | Duration | elog #
2013-06-15 |  | 6775
2013-08-25 |  | 7569
2013-09-05 |  | 7663
2013-09-29 |  | 7905
|  | 4 | events