## Slide 1: Control system in Gravitational Wave Detectors
Koji Arai – LIGO Laboratory / Caltech
LIGO-G1401146-v2

## Slide 2: Introduction ~ Control?
Gravitational wave detection Laser displacement sensor Requires linear displacement detection
Control for measurement Laser interferometer = nonlinear device Feedback control => linearization

## Slide 3: Introduction ~ Control?
What is the feedback control?
  A scheme to monitor and modify output(s) of a system by changing the input(s) depending on the output(s)
Examples
  Shower temperature
  Car driving
  Tight rope walking
Imagine what happens
  If the response is too slow?
  If the response is too fast?
  Air conditioning
  Bike riding
  Inverted bar on a hand

## Slide 4: Introduction ~ Control?
Elements of a feedback loop
  Stabilized
  motion
  Disturbance
  Error
  signal
  Sensor
  Servo
  Filter
  Actuator
  Feedback
  signal
  m or m/Hz1/2
  m or m/Hz1/2
  V or V/Hz1/2
  V or V/Hz1/2
  V/m
  V/V
  m/V
  Transducer
  Filter

## Slide 5: Introduction ~ Control?
Sensor: Transducer for displacement-to-voltage conversion
If the sensor is completely linear (and has or no frequency dependence)
We don’t need feedback control!

## Slide 6: Introduction ~ Control?
In reality: Sensors, laser interferometers in particular, are nonlinear!
Enclose the operating point in the linear region => The system recovers linearity
Was the displacement modified by the feedback? => Precise knowledge of the control system
for signal reconstruction

## Slide 7: Introduction ~ Control?
Elements of a feedback loop
disturbance
error signal
sensor
servo filter
feedback
signal
actuator
Open loop transfer function
stabilized disturbance

## Slide 8: Introduction ~ Control?
When G is small:
disturbance
error signal
sensor
servo filter
feedback
signal
actuator
Open loop transfer function
stabilized disturbance

## Slide 9: Introduction ~ Control?
When G is big: e.g. G = 10, 100, or 1000
dx
dxs= dx/(1+G) = 0.09 dx, 0.01 dx, 0.001 dx
Verr= dx H /(1+G)
Vfb = G/(1+G) /A dx
= 0.91 dx/A, 0.99 dx/A, 0.999 dx/A
0.91 dx, 0.99 dx,
0.999 dx
Actuator range
Filter frequency response / loop stability
Residual error
vs
Sensor linear range

## Slide 10: Introduction ~ Control?
When the openloop gain G is >>1, the error signal gets suppressed
“Wow! our sensor signal became smaller!” Is our system more sensitive now?
  No. We are just moving the actuator so that the error signal looks smaller. The signal and noise are equally suppressed in the error signal. Thus the SNR does not change.
OK... So can we still measure gravitational waves even if the error signal is almost zero?
  Yes. We should be able to recover the original signal by compensating the effect of the control i.e. (1+G)
  And we can also use the feedback signal in order to reconstruct the original signal with appropriate compensation i.e. (1+G)/G

## Slide 11: Introduction ~ Control?
Important difference between
  “Feedback control for stabilization” and “Feedback control for measurement”
  Feedback control changes the stabilized motion but reconstructed Disturbance is not modified by the loop* (*if everything is linear)
  Stabilized
  motion
  Disturbance
  Error
  signal
  Sensor
  Servo
  Filter
  Actuator
  Feedback
  signal

## Slide 12: Linear systems and their stability

## Slide 13: Linear systems and their stability
A deterministic and time-invariant system: H
The system H is LTI (linear & time-invariant) when
We can deal with such a system using Laplace transform (or almost equivalently Fourier Transform)
H
  System Input
  x(t)
  System Output
  y(t)

## Slide 14: Linear systems and their stability
Time domain vs Laplace (or Fourier) frequency domain
http://en.wikipedia.org/wiki/Linear_system
http://en.wikipedia.org/wiki/LTI_system_theory
Impulse response
transfer function

## Slide 15: Linear systems and their stability
It is easy to convert from an ordinary differential equation to a transfer function
e.g. Damped oscillator
  Laplace Transform
  Fourier Transform

## Slide 16: Linear systems and their stability
e.g. Damped oscillator
Bode diagram
  log-log
  log-lin

## Slide 17: Linear systems and their stability
e.g. RC filter
cut-off freq
w = 1/(R*C)

## Slide 18: Linear systems and their stability
In most cases, a system TF can be expressed as:
The roots of the numerator are called as “zeros” and the roots of the denominator are called as “poles”
Zeros (szi) and poles (spi) are real numbers (single zeros/poles) or pairs of complex conjugates (complex zeros/poles)

## Slide 19: Linear systems and their stability
Poles rule the stability of the system! H(s) can be rewritten as
Each term imposes exponential time impulse response
Therefore, if there is ANY pole with the response of the system diverges

## Slide 20: Linear systems and their stability
Poles rule the stability of the system!
http://nupet.daelt.ct.utfpr.edu.br/_ontomos/paginas/AMESim4.2.0/demo/Misc/la/SecondOrder/SecondOrder.htm
  Re(s)
  Im(s)
  Unstable
  response

## Slide 21: Linear systems and their stability
Requirement: All the roots for 1+G should be
in the left hand side of Laplace plane
  dx
  dxs = dx/(1+G)
  Open loop TF:
  G = H F A
  Closed loop TF:
  GCL = 1/(1+G)
  (!)
Now we eventually came back to this diagram

## Slide 22: Linear systems and their stability
Remarks
Requirement:
All the roots for 1+G should be in the left hand side of Laplace plane
  This does not mean all H, F, A needs to be stable. e.g. Unstable mechanical system A can be stabilized by a control loop. (cf. An inverted Rod)
  We usually play with F to tune the result. It is awkward to evaluate the stability of 1/(1+G) every time. There is a way to tell the stability only from G
  Nyquist’s stability criterion
  Open loop TF:
  G = H F A
  Closed loop TF:
  GCL = 1/(1+G)

## Slide 23: Linear systems and their stability
Nyquist stability criterion
  Plot openloop gain G in a complex plane (i.e. Nyquist diagram)
  If the locus of G(f) from f=0 to ∞, goes to 0 looking at the point (-1 + 0 i) at the left side => Stable
  If the locus sees the point (-1+o i) at the right side => Unstable
  Unity gain frequency fUGF : for |G(fUGF)| = 1
  Phase margin θ : θ = Arg(G(fUGF))
  Gain margin g: g = 1/|G(f0)| where Arg(G(f0)) = -π

## Slide 24: Linear systems and their stability
Phase Margin / Gain Margin in Bode diagram
  Most of the case, a bode diagram of G is enough to see the stability
  Nearly
  unstable
  Gain Margin
  4.3dB@420kHz
  -180deg
  Phase Margin
  48deg
  Unity Gain Freq
  (176kHz)
  A rough standard of a stable servo loop:
  Phase Margin > 40deg
  Gain Margin > 10dB

## Slide 25: Linear systems and their stability
Building blocks (“zpk” representation)
  Single pole
  Single zero
  A pair of complex poles
  A pair of complex zeros
  Gain

## Slide 26: Linear systems and their stability
Relationship between pole/zero locations and w0&Q
To be compared with

## Slide 27: Linear systems and their stability

## Slide 28: Linear systems and their stability
Summary
  Classical control theory
  Design locations of poles and zeros
  Stability: tuning of open loop transfer function is important

## Slide 29: Control system components in GW detectors

## Slide 30: Control systems
Elements of a feedback loop (again)
  Sensor
  Servo
  Filter
  Actuator

## Slide 31: Interferometer control system
Local control vs global control

## Slide 32: Local Sensors
L Carbone, Class. Quantum Grav. 29 (2012) 115005
Shadow sensor (relative displacement sensor)
  For suspension damping control, mirror attitude monitor
  Typical linear range ~1mm for 0-10V => dV/dx = 10 kV/m Typical noise level: ~ 100 pm/sqrtHz
  aLIGO: Birmingham Optical Sensor and Electro-Magnetic actuator (BOSEM)

## Slide 33: Local Sensors
Shadow sensor (relative displacement sensor)
  Linear range (~0.7 mm) / displacement noise sensor locations
L Carbone, Class. Quantum Grav. 29 (2012) 115005

## Slide 34: Local Sensors
Linear Variable Differential Transducer (relative disp. sensor)
  For low freq pendulum control (inverted pendulum), larger range VIRGO Super attenuator, KAGRA Seismic Attenuation System
  Typical linear range ~10mm for 0-10V => dV/dx = 1 kV/m Typical noise level: 10~100 nm/sqrtHz
H. Tariq, Nuclear Instruments and Methods in Physics Research A 489 (2002) 570–576
  10 nm/sqrtHz

## Slide 35: Local Sensors
Optical Lever (relative angular sensor)
  Angle local control
  Typical linear range ~beam side (0.1~1 mm) => dV/dθ = 1 ~10 kV/rad Typical noise level: 0.01~1 nrad/sqrtHz
  A B
  C D
  SUM = A+B+C+D
  X = A–B+C-D
  Y = A+B-C-D
  Optical Lever Calibration

## Slide 36: Local Sensors
Piezo Accelerometer (Inertial sensor)
  Vibration measurement
  Typical linear range ~ 100~1000 m/s2 Typical noise level: 0.5 ~ 50 (μm/s2)/sqrtHz

## Slide 37: Local Sensors
Servo Accelerometer (Inertial sensor)
  Seismic platform control (f>0.1Hz), Vibration measurement
  Apply force to the suspended mass
  => Keep the distance from a reference
  When the control gain G>>1
  => a = F / m

## Slide 38: Local Sensors
Servo Accelerometer (Inertial sensor)
  Above the resonant freq: Limited by the sensor noise
  Below the resonant freq: Steep rise of the noise as the mass does not move in relative to the ground => Low resonant freq is beneficial
  1 pm/sqrtHz
A. Bertolini et al, Nuclear Instruments and Methods in Physics Research A 564 (2006) 579–586
  Positive
  spring
  Negative
  spring

## Slide 39: Acutuators
Mechanical actuators
  Coil Magnet actuator
  Electro Static Driver (ESD)
  Piezo (PZT) actuator
Optical actuators
  Acousto-Optic Modulator
  Electro-Optic Modulator
  Laser Frequency

## Slide 40: Acutuators (Mechanical)
Coil-magnet actuator
  Coil current induces force on a magnet attached to a mass
  Contactless
  aLIGO coil-magnet actuator is integrated in BOSEM
  Actuator response (coupling) has position dependence. Preferable to use it at its maximum in order to avoid vibration coupling
  1 gw = 9.8e-3 N

## Slide 41: Acutuators (Mechanical)
Electro Static Driver (ESD)
  Apply potential close to the mirror => induces surface charge (or polarization) and attractive force
  In practice, comb patterns are used => strengthen the electric field, but less force range
  Can produce only attractive force. => Need DC Bias.
  Stray surface charging may cause problems.

## Slide 42: Acutuators (Mechanical)
Piezo (PZT) actuator
  Apply potential to a feroelectric material => cause internal polarization and induces strain
  To increase displacement, laminated piezo is often used => displacement 3~10 μm
  Requires a bias voltage and HV amplifier, but has wide applications
  OMC cavity mirror

## Slide 43: Acutuators (Optical)
Acousto-Optic Modulator
  Phonon-Photon scattering (or bragg diffraction) in AOM crystal
  Effect: Beam deflection / Frequency shift
  Application: Laser frequency actuator, Laser intensity actuator Beam angle scanner

## Slide 44: Acutuators (Optical)
Electro-Optic Modulator
  Pockels Cell effect: Refractive index changes linearly to the applied E-field
  Application: Laser phase modulation Phase actuation (= frequency actuation)
  LiNbO3 crystal

## Slide 45: Acutuators (Optical)
Laser frequency actuation (YAG NPRO laser)
  We often control laser frequency with multiple actuators
  1) Thermal actuator Thermo-Electric Cooler attached to the laser crystal. Huge response (1GHz/K or 1GHz/V) but slow (f<0.1Hz)
  2) Fast piezo actuator A piezo attached on the laser crystal induces stress induced refractive index change. Response (~1MHz/V). Bandwidth 10~100kHz
  3) External EOM Response (~10 mrad/V), Bandwidth ~1MHz

## Slide 46: Servo Controller
Analog servo filters
  High dynamic range (~1nV/sqrtHz, +/-10V), High bandwidth
  Pole/zero placement with active op-amp filters
  Until the end of the 20th century, analog filters have been commonly used for servo filters in our field
  Analog servos are still in action for the feedback loops with bandwidth >1kHz. (cf. frequency stabilization, intensity stabilization)
  LIGO 40m prototype (1998)

## Slide 47: Servo Controller
Digital servo filters
  Process digitized signals in a computer
  Large flexibility High compatibility with detector automation and management
  Limited dynamic range (~0.1mV/sqrtHz, +/-10V for 16bit)
  Limited bandwidth
    Each sample needs to be processed before the next sampled data comes
    Inevitable sampling delay
    Additional phase delays due to analog filters for analog-digital interface
    e.g. 16kHz sampling, control bandwidth ~200Hz

## Slide 48: Analog/Digital interface
Restriction of signal digitization
  Voltage quantization: quantization noise => limited dynamic range => Requires whitening/dewhitening filters
  Temporally discrete sampling: aliasing problem => limited signal bandwidth => Requires anti-aliasing (AA) / anti-imaging (AI) filters
  Typical signal chain

## Slide 49: Control room
Comparison of the control room in the analog and digital eras
  TAMA300 (2001)
  aLIGO (2014)