## Slide 1: “You can look but you can't touch” I. First Contact wrt Charge II. Experiments to back up ETM / ERM Discharge procedure III. Experiments associated with Static Residue Systems meeting 22nd April 2015 Kate Gushwa, Calum Torrie and Kyle Baric (with Travis Sadecki and Betsy Weaver at LHO and previously with GariLynn Billingsley and Norna Robertson at CIT and input from Danny Sellars and Gary Traylor)

## Slide 2: I. First Contact wrt Charge
I. Experimental Setup
ETM (single optic on V-block)
On a flow bench in clean-room at LHO
Grounded metal box
All other items also grounded including table, stand etc …
Electrometer positioned 1” from surface being measured.
Always peeled FC from AR side
Measured on both HR and AR side.
Zeroed electrometer with cap
Refer to LIGO-T1500170: Charge Experiments at LHO - March 2015

## Slide 3
Chart: Top Gun vs No Top Gun During Peel of First Contact (Measured from HR Side)

## Slide 4
Chart: Test #3 (Measured from HR Side)

## Slide 5
Chart: Test #3 (Measured from HR Side)

## Slide 6
Chart: Top Gun During & After Peel (Measured from Both HR & AR)

## Slide 7: II. Experiments to back up ETM / ERM Discharge procedure
II. Experimental Setup
ETM and PUM [ERM] each on a V-block
On a flow bench in clean-room at LHO
Grounded metal box
All other items also grounded including table, stand etc …
Electrometer positioned 1” from surface being measured.
10mm gap
Applied charge in gap
Measurements made from both behind PUM [ERM] and also in gap at ETM (after separated optics).
Refer to LIGO-T1500101: ETM / ERM Gap discharge procedure

## Slide 8: II. Experiments to back up ETM / ERM Discharge procedure
ERM
ETM
1 in
1 mm
AR
ETM
1 in
AR

## Slide 9
Refer to LIGO-T1500170:
Charge Experiments at LHO - March 2015
For Video of flappy panel technique refer to DCC number of this doc.

## Slide 10
Chart: Charging with First Contact & Discharging with Flappy Panel (Days 4 & 5 Part A, Measured from ERM AR)
Refer to LIGO-T1500170: Charge Experiments at LHO - March 2015 for full details of all steps / tweaks carried out here
  Consistency & repeatability
ERM
ETM

## Slide 11
Chart: Long Duration Experiments (Days 4 & 5 Part A Measured from ERM AR)
ERM
ETM

## Slide 12
Chart: Long Duration Experiments (Days 4 & 5 Part B, Measured from ETM AR)
ETM

## Slide 13
Chart: Long Duration Experiments (Days 4 & 5 Part C, Measured from ETM AR)
ETM

## Slide 14
Chart: 300 Seconds After Flappy Panel Removed (Days 4 & 5 Part A, Measured from ERM AR)
| 300s after FP
4.4 | -0.4
4.5 | -1
4.6 | -0.8
4.7 | -0.4
5.1 | -1.8
5.2 | 0.6
5.4 | -1.6
5.6 | 4.8
5.8 | 4.4
ERM
ETM

## Slide 15
Chart: 800 Seconds After Flappy Panel Removed (Days 4 & 5 Part B, Measured from ETM AR)
| Final Value
4.4 | 10.7
4.5 | 7.6
4.6 | 5.9
4.7 | 10
4.7 next morning | -1.6
5.1 | 12.8
5.2 | 7.4
5.3 | -11.8
5.4 | 13.7
5.6 | 13.3
5.8 | 9.2
5.8 After lunch | 4.6
ETM

## Slide 16: III. Experiments associated with Static Residue
III. Experimental setup
1” thick x 34cm diameter glass plate (one with gold pattern, as shown).
All experiments carried out in a grounded box.
Electrometer (USSVM II) used for measurements.
Electrometer position 1” from front face of plate.
Electrometer positioned at 2 o’clock position.
Various items used to charge optic!
  Touching or wiping in the general vicinity of the electrometer
Top Gun used at 10-20 PSI to discharge
  CONC: Can always recover with TG!
LIGO-T1500157: Static residue tests Gloves and plastics

## Slide 17: Qualakate Vs. Latex (current gloves we use in chamber)
Chart: Seconds
5 hold
8 release
11 hold
14 release
17 hold
20 release and walk a way
Qualakate
glove

## Slide 18: Teflon- similar to that used in ETM /ITM
Chart: 
5s-15s rub with Teflon
30s-45s “
60s-75s “
90s End

## Slide 19: Wipe (believe Contec wipe, similar to ones used at site)
Chart: 
5s-15s rub with wipe
30s-45s “
60s-75s “
90s End

## Slide 20: Alpha Swab (with stabilization)
Chart: 
5s-15s wipe with cloth
30s-40s “
55s-65s “
80s-85s Blow with Top Gun
100s End
  CONC: Can always recover with TG!

## Slide 21: Berkshire wipes (with stabilization)
Chart: 
5s-15s wipe with cloth
30s-40s “
55s-65s “
80s-85s Blow with Top Gun
100s End
  CONC: Can always recover with TG!

## Slide 22: Summary / Actions / Conclusions
Discharge procedure updated, LIGO-T1500101: ETM / ERM Gap discharge procedure. (version 2)
This work has taken us back to the 30s post peel recommendation with top gun.
Also recommending use of electrometer to check for charge & use top gun to “clean up”.
As a result we have a large number of documents updated
  LIGO-E1300142: DCN Associated with a collection of documents related to in chamber work in particular cleaning of optics (Unapproved)
Top Gun is our friend
  So far shown can always (so far) recover with Top Gun
ACTIONS
  Deliver TMDS to LLO, LHO and MIT (high priority)
  Go over docs including updated discharge procedure with Detector Engineering and Operator teams
  Work on whether Discharge procedure can be used on ITM (low priority)