## Slide 1: Noises in Gravitational Wave Detectors
Koji Arai (新井宏二) – LIGO Laboratory / Caltech
GW mini-school: Beijing Normal University 2016/9/15~18
LIGO-G1601927-v1
引力波

## Slide 2: Introduction ~ Noise?
GW detection Data stream of differential arm strain
Once recorded: Signals and noises are indistinguishable What we can do is to catch “likely” features
Reduce any kind of noises!

## Slide 3: Introduction ~ Noise?
Time domain vs frequency domain
Noise
Reduction
Time domain: transient noises Frequency domain: stationary noises

## Slide 4: Introduction ~ Noise?
Power Spectral Density (PSD) Double sided PSD (-Infinity < f < Infinity)
Single sided PSD (0 <= f < Infinity) [xunit2 / Hz]
Linearized PSD or Amplitude Spectral Density (ASD): [xunit/ sqrtHz]

## Slide 5: Introduction ~ Noise?
Parseval’s Theorem for signal RMS and PSD
Root Mean of x(t): average signal power density (per sec)
(cf. variance, std deviation)
PSD Sx(f):
power density per frequency (per sec)

## Slide 6: Introduction ~ Noise?
Example
PSD [fm/sqrtHz] in log-log scale, RMS [fm] ~ 50fm = 0.05pm
PSD [fm/sqrtHz] in log-lin scale
RMS [fm]
Time series [pm]

## Slide 7: Components of the interferometer
3 fundamental elements of GW detectors
Mechanics
Optics
Electronics

## Slide 8: Components of the interferometer
3 fundamental elements of GW detectors
Mechanics
Optics
Electronics

## Slide 9: Noise categories
3 fundamental elements of GW detectors
Mechanics -> Displacement noises
Optics -> Optical noises
Electronics -> Electrical noises
Optical
Optical
Displacement
Displacement
Electrical
Electrical
Electrical

## Slide 10: Noise coupling
Noise source / Noise coupling
Reduce this or this
Noise Source
GW
Channel
Coupling
Mechanism

## Slide 11: Sensitivity and noise
Sensitivity (=noise level) of Advanced LIGO
Current sensitivity
h= 8x10-24 /rtHz

## Slide 12: Sensitivity and noise
Sensitivity (=noise level) of Advanced LIGO
Noise budget
h= 2x10-23 /rtHz

## Slide 13: Displacement noises

## Slide 14: Displacement noise
Mechanical displacement sensed by a laser interferometer
The longer the arm length, the smaller the strain noise
  Seismic noise
  Thermal noise
  Newtonian Gravity noise

## Slide 15: Displacement noise
Seismic noise
  Even when there is no noticeable earth quake…
http://link.aps.org/doi/10.1103/RevModPhys.86.121 (http://arxiv.org/abs/1305.5188)
Target disp. noise
10-20m/rtHz
Ground motion @10Hz 10-10m/rtHz

## Slide 16: Displacement noise
Vibration isolation ~ utilize a harmonic oscillator
  A harmonic oscillator provides vibration isolation above its resonant frequency
w0=1Hz
|x/X|~1/1000@30Hz

## Slide 17: Displacement noise
How to get more isolation?
Damping
Lower the peak height
Worse isolation
Lower resonant freq
Better isolation
Complex to realize
Multi stage
Steeper isolation curve
More peaks
f -2
f -4
In practice: employ combination of these measures

## Slide 18: Displacement noise
iLIGO vibration isolation
Hydraulic active isolation / Isolation stack / Single Pendulum

## Slide 19: Displacement noise
aLIGO vibration isolation
Hydraulic active isolation / Invacuum Active Isolation Platforms / Multiple Pendulum

## Slide 20: Displacement noise
http://link.aps.org/doi/10.1103/RevModPhys.86.121 (http://arxiv.org/abs/1305.5188)
10-10

## Slide 21: Displacement noise
Thermal noise:
Thermally excitation of the motion of a system
  Each d.o.f. has <E> = ½ kb T (Equipartition theorem)
能量均分定理

## Slide 22: Displacement noise
Thermal noise:
Fluctuation dissipation theorem (FDT)
涨落耗散定理

## Slide 23: Displacement noise
cf Johnson noise:
Resistor R: Thermal voltage noise
Measurement of R: R = V / I
Wikipedia: Johnson noise

## Slide 24: Displacement noise
Suspension thermal noise Measurement of loss: -> Q factor
http://www.giangrandi.ch/electronics/ringdownq/ringdownq.shtml
Suspension thermal noise
Frequency [Hz]
Displacement noise [m/sqrtHz]
Q = 101 ~ 106
m = 40kg T = 300K
Monolithic suspension
for high pendulum Q

## Slide 25: Displacement noise
Mirror thermal noise
Sensing of the mirror surface deformation with a laser beam (with intensity profile of f(r))
Apply periodic pressure with profile of f(r)
Calculate the rate of dissipation Wdiss
analytically, using FEA, or etc
Put this into the formula
We don’t need to
know Z(w)!
Y. Levin PRD 57, 659-663 (1998)

## Slide 26: Displacement noise
Mirror substrate thermal noise
  Brownian motion Mechanical loss associated with the internal friction Thermally excited body modes Optical coating (high mechanical loss) will be limiting noise source in aLIGO
  Thermo elastic noise Elastic strain & thermal expansion coefficient => cause heat distribution & flow in the substrate Temperature fluctuation causes mirror displacement
  Thermo-refractive noise Temp. fluctuation causes fluctuation of refractive index
We don’t need to
know Z(w)!

## Slide 27: Displacement noise
Newtonian Gravity noise
  Mass density fluctuations around the test masses => test mass motion via gravitational coupling
  Dominant source of Newtonian noise = Seismic surface wave
  Mitigation 1) Going to quiet place (underground) 2) Feedforward subtraction 3) Passive reduction by shaping local topography
NN
aLIGO
sensitivity
J Driggers, et al, PRD 86, 102001 (2012)
J Harms, et al, Class. Quantum Grav. 31 185011 (2014)

## Slide 28: Optical noises

## Slide 29: Optical noises
Noises that contaminate the readout signal
  Quantum noises (shot noise, radiation pressure noise)
  Laser technical noises (frequency/intensity noise)
  Modulation noises
  Scattered light noise
Optical
Optical

## Slide 30: Optical noises
Quantum noises: Shot noise
  Photon shot noise associated with photodetection
  Michelson interferometer
  Michelson response (@DC)
at the limit of d𝜙->0
  iDC : DC Photocurrent
  η : PD Quantum Efficiency
  ν : Optical Frequency
Shot-noise limit of the Michelson phase sensitivity
  1.3x10-20 1/sqrtHz
  @1W, 1064nm, 4km
Michelson
Strain Sensitivity

## Slide 31: Optical noises
Quantum noises ~ Radiation pressure noise
  Photon number fluctuation in the arm cavity => Fluctuation of the back action force
  Vacuum fluctuation injected from the dark port => Differential power fluctuation => Cause the noise in the GW signal

## Slide 32: Optical noises
Quantum noises
  Standard Quantum Limit (SQL)
- Trade-off Between Shot Noise and Radiation-Pressure Noise
- Uncertainty of the test mass position due to observation
=> Quantum Optical techniques to overcome SQL

## Slide 33: Optical noises
Laser frequency noise
  Laser wavelength (λ = c / ν) = reference for the displacement measurement
  Optical phase 𝜙 = 2 pi ν L / c d𝜙 = 2 pi / c (L dν + ν dL) <= indistinguishable
  Laser frequency stabilization

## Slide 34: Optical noises
Laser intensity noise
  Relative Intensity Noise (RIN): dP/P
  Sensor output V = P x => dV = P dx + x dP <= indistinguishable
  Laser intensity stabilization
P. Kwee et al, Optics Express 20 10617-10634 (2012)
  In-vacuum 8-branch Photodiode array

## Slide 35: Electrical noises

## Slide 36: Noise in photodetectors
P-layer
I-layer
N-layer
Photocurrent
Photodiodes
aLIGO InGaAs PD (𝜙3mm)
Anode
Cathode
“Photodiode Amplifiers”, J. Graeme (McGrawHill 1995)
input referred noise current
  Photodiode equivalent circuit
    Shunt Capacitance RD
    Junction Capacitance CD
    Series Resistance RS

## Slide 37: Digitization (Quantization) noise
Analog signals (~+/-10V) -> Digital signal
  Digitized to a discrete N bit integer number
  Quantization causes a white noise e.g. +/-10V 16bit => Δ = 0.3mV => Vn ~ 100 μV/sqrtHz cf. Input noise of a typical analog circuit 10nV/sqrtHz
http://www.analog.com/static/imported-files/tutorials/MT-229.pdf
quantization
quantization error

## Slide 38: Digitization (Quantization) noise
Whitening
  Amplify a signal in the freq band where the signal is weak
Dewhitening
  Amplify a signal in the freq band where the signal is weak

## Slide 39: Actuator noise
Actuator noise appears in the GW signal as an external disturbance
  Mitigation 1) Make the noise itself smaller 2) Make the actuator response smaller
  We need to keep sufficient actuator strength for lock acquisition => Transition to a low-noise mode after achieving lock

## Slide 40: Summary

## Slide 41: Summary
Summary
  There are such large number of noises
  They are quite omni-disciplinary
  Even only one noise can ruin our GW detection
  GW detection will be achieved by
    Careful design / knowledge / experience
    Logical, but inspirational trouble shooting
  Noise “hunting”