LIGO Laboratory / LIGO Scientific Collaboration

LIGO-T1700237-v1	LIGO	5/30/2017

Test Procedure for Bullseye Photo Sensor

Marc Pirello

Distribution of this document:

LIGO Scientific Collaboration

This is an internal working note

of the LIGO Laboratory.

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 159 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

http://www.ligo.caltech.edu/

# Introduction

The following Test Procedure describes the test of proper operation of the Bullseye Photo Sensor.

S/N ____________________	 	Tester ______________	Date __________________

# Test Equipment

- Voltmeter

- Oscilloscope

- Modulated light source near 60kHz.

  - Fluorescent overhead bench lamp (55kHz)

  - Etc…

- Board Schematics, LIGO D1700060-v2.

# Tests

- Verify the proper current draw.  Using a bench DC supply apply +- 18 Volts to the external power connector.

+18 Volt current ______________ 	100mA +/- 25mA

-18 Volt current ______________ 	80mA +/- 25mA

- Connect the DB9 signal output (J2) to an oscilloscope and measure channels S1-S4.  Adjust the oscilloscope to measure 5V per division and 1 second per division.  Excite each photo diode with a low power laser (1mw) and verify its response.

Center Segment, S1 (Pins 3, 8) ______________ -14V +/- 2V

Right Segment, S2 (Pins 2, 7) ______________ -17V +/- 2V

Bottom Segment, S3 (Pins 1, 6) ______________ -17V +/- 2V

Left Segment, S4 (Pins 4, 9) ______________ -17V +/- 2V

- Verify that the photo sensor is able to measure 60kHz modulation.  Apply frequency modulated light to the photo sensor opening.  Adjust oscilloscope to trigger at 60kHz and measure frequency on each channel S1-S4.

Light Source _________________________________

Frequency Measured S1_______________________________

Frequency Measured S2_______________________________

Frequency Measured S3_______________________________

Frequency Measured S4_______________________________