## how to use this workbook
How to use this workbook:
1) Get an E-number for this test report from https://dcc.ligo.org
a) Label the document "Optical Contamination Test Results: <name>", where <name> is a descriptive name of the article under test
b) For the E-<number>-X0 version DCC entry, give a COMPLETE description of the article(s) under test in the Abstract (not the "Changes and Notes"), including manufacturer/source, material grade or assembly/device model number, and quantity/amount. For example, it is not sufficient to say "photodiode", but rather state the manufacturer's name and model number, such as "Vishay TSTS7100 IR-LED". Likewise "PEEK" is not a sufficient material descriptor, but "Victrex PEEK grade 450GL" is a complete description.
c) For the E-<number>-X0 version DCC entry, give a URL link to the location where intermediate test results will be posted in the "Notes and Changes" field.
2) Rename this template workbook to "E<number>-v1 <name>.xls"
3) Fill in only the light yellow colored cells in the sheet labelled "scaled Loss rate".
You must decide the appropriate units to characterize the material, or assembly or device under test. For example, for a device or assembly the appropriate units is likely "each", whereas for a bulk material (say PEEK) the appropriate unit is likely surface area (cm^2). For a lubricant, such as Krytox, the appropriate units may be mass, or volume.
4) Fill in only the light yellow colored cells in the sheets labelled "d1", "d2", "d3", etc.
Each sheet is a data set taken at a particular day and time. In the first data set sheet (d1), the start date of the test and the cavity number is entered. Data for the TEM 00-01 beat frequency as a function of power and the ring down data are entered for each data set. The data is plotted and fitted automatically. The sheet labelled "data" summarizes the absorption and total loss measurements as a function of time in a tabular listing. Sheets "absorption" and "total loss" summarize the absortion and total loss in plots and a least squares ft, as the data is entered. The "scaled loss rate" sheet summarizes the results scaled to the LIGO Observatory for the intended use and is updated automatically as data is entered.
5) The test is terminated when a determination is made that the material/assembly/device under test has either passed or failed. At this point, use the "Convert to Adobe PDF" tool to create a file with just the spredsheets labelled "scaled loss rate", "absorption" and "total loss" which is named "E<number>-v1 <name>.pdf".
6) Select "Create E<number>-v1" in the DCC entry for this document and upload the pdf file "E<number>-v1 <name>.pdf" as the "Main file in document" and this Excel file "E<number>-v1 <name>.xls" as an "Other file" (not Main). Then in either the abstract or the "Notes and Changes" field indicate the test results (passed or failed).
7) File any other documents, which help to define the material/assembly/device which was tested, as "Other files" under the DCC entry for E<number>-v1, such as photos or descriptions of the test (if it was a unique test, such as taken at elevated temperatures).
8) File proprietary data (data not owned by LIGO Lab, even if posted to the web) about the material/assembly/device such as MSDS sheets, specifications, etc. This data should be filed under a separate E-number and then linked from the E-number which corresponds to the optical contamination cavity test results. Even if the MSDS, specifications, etc. are already on the web, they should be filed in the DCC, in case the material/assembly/device is discontinued (in which case the manufacturer often removes this data from their web sites.)

## scaled loss rate
LIGO E-Document Number: |  |  | E1800020
Sample Test:
| Material under test: |  | EPO-TEK 301-3M adhesive
| units |  |  | 135 | cm^2 surface area
| absorption | 0.38833 | ± | 0.49378 | ppm/yr | 1 sigma
| scatter | 27.8498 | ± | 19.69 | ppm/yr | 1 sigma
| max. normalized absorption |  |  | 0.010192 | ppm/yr/unit | 2 sigma
| max. normalized scatter |  |  | 0.498 | ppm/yr/unit | 2 sigma
| test turbopump speed (liter/s) |  |  | 24.3916 | torr/liter/sec
Scaled to LIGO:
| LIGO Vacuum Volume |  |  | Vertex | LHO Diagonal | End | Comments
| Quantity (units) |  |  |  |  | 0.004 | cm^2 (see abstract of E1800020 for area calculation)
| LIGO ion pumping speed (liter/s) |  |  | 6800 | 6800 | 1700 | see E0900398 or PSI V049-1-078 for pump rates
| pumping speed ratio (test/LIGO) |  |  | 0.003587 | 0.003587 | 0.014348 | does not include cryo-pump and effective pumping from the Beam Tube
| max. absorption (ppm/yr) |  |  | 0 | 0 | 5.8492e-07 | * Limit is < 0.02 ppm/yr for a single source
| max. scatter (ppm/yr) |  |  | 0 | 0 | 2.8581e-05 | * Limit is < 0.2 ppm/yr for a single source
* The overall limit on contamination loss on optics for AdL is < 0.5 ppm/yr absorption and < 4 ppm/yr scatter from all sources, per Table 4 of the COC Design Requirements Document (T000127-v1). It is assumed that ~20 sources could contribute.
|  |  |  | 2.75 | in, dia, flex hose
|  |  |  | 36 | in, length, flex hose
|  |  |  | 6.985 | cm, dia
|  |  |  | 91.44 | cm, length
|  |  |  | 463.669 | C1, aperture, liter/sec
|  |  |  | 45.0971 | C2, flex hose, liter/sec
|  |  |  | 41.0997 | Ct, liter/sec
|  |  |  | 60 | S, liter/sec [Varian Dry TPS-compact with TV81M]
|  |  |  | 24.3916 | Seff, liter/sec

## absorption
LIGO E-Document Number: |  | E1800020
Test Material/Assy./Device: |  | EPO-TEK 301-3M adhesive
Absorption fitting
| Slope | 4.433e-05 | 3.1451 | Y-intercept
| Standard Error | 5.6367e-05 | 0.040743 | Standard error
| r2 | 0.037218 | 0.084298 | sey
| F | 0.6185 | 16 | df
| ssreg | 0.0043952 | 0.1137 | ssresid
| Absorption change rate (ppm/yr) |  | ± sigma (ppm/yr)
| 0.38833 |  | 0.49378

## total loss
LIGO E-Document Number: |  | E1800020
Test Material/Assy./Device: |  | EPO-TEK 301-3M adhesive
Total loss fitting
| Slope | 0.0031792 | 182.4606 | Y-intercept
| Standard Error | 0.0022477 | 1.6247 | Standard error
| r2 | 0.11114 | 3.3615 | sey
| F | 2.0006 | 16 | df
| ssreg | 22.6057 | 180.796 | ssresid
| Total loss change rate (ppm/yr) |  | ± sigma (ppm/yr)
| 27.8498 |  | 19.69

## data
| Time (hr) | Absorption (ppm) | Total Loss (ppm) | EPO-TEK 301-3M adhesive
1 | 72.4167 | 3.2008 | 171.1851
2 | 96.4167 | 3.0498 | 185.1525
3 | 144.3542 | 3.2058 | 185.1546
4 | 216.3542 | 3.1258 | 185.2501
5 | 264.375 | 3.2504 | 185.1174
6 | 408.4167 | 3.2176 | 185.4566
7 | 480.4931 | 3.2112 | 185.1302
8 | 600.4167 | 3.1813 | 185.0694
9 | 648.375 | 3.0784 | 185.252
10 | 720.3125 | 3.0615 | 189.0594
11 | 768.375 | 3.1134 | 184.8745
12 | 816.375 | 3.1908 | 185.0403
13 | 888.3333 | 3.0665 | 184.9702
14 | 936.3125 | 3.1476 | 185.0348
15 | 984.3125 | 3.32 | 184.2369
16 | 1056.3125 | 3.1244 | 184.8446
17 | 1104.3333 | 3.3234 | 184.6214
18 | 1152.3125 | 3.2473 | 184.9522
19 | #N/A | #N/A | #N/A
20 | #N/A | #N/A | #N/A
21 | #N/A | #N/A | #N/A
22 | #N/A | #N/A | #N/A
23 | #N/A | #N/A | #N/A
24 | #N/A | #N/A | #N/A
25 | #N/A | #N/A | #N/A
26 | #N/A | #N/A | #N/A
27 | #N/A | #N/A | #N/A
28 | #N/A | #N/A | #N/A
29 | #N/A | #N/A | #N/A
30 | #N/A | #N/A | #N/A
31 | #N/A | #N/A | #N/A
32 | #N/A | #N/A | #N/A
33 | #N/A | #N/A | #N/A
34 | #N/A | #N/A | #N/A
35 | #N/A | #N/A | #N/A
36 | #N/A | #N/A | #N/A
37 | #N/A | #N/A | #N/A
38 | #N/A | #N/A | #N/A
39 | #N/A | #N/A | #N/A
40 | #N/A | #N/A | #N/A
41 | #N/A | #N/A | #N/A
42 | #N/A | #N/A | #N/A
43 | #N/A | #N/A | #N/A
44 | #N/A | #N/A | #N/A
45 | #N/A | #N/A | #N/A
46 | #N/A | #N/A | #N/A
47 | #N/A | #N/A | #N/A
48 | #N/A | #N/A | #N/A
49 | #N/A | #N/A | #N/A
50 | #N/A | #N/A | #N/A

## d1
Cavity No | Sample | Comment
1 | EPO-TEK-301-3M | 12 Micro scope slides 25mm X 45mm with EPO-TEK on onre side, See related docs for ferther info.
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2018-01-27 | 2018-01-30 | 10:00 | 3 | 72.4167 | 2.95
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
25.3
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.988 | 0.097 | 74.0145 | 1.9472e-05 | 1.02e-08
270 | 0.96 | 0.095 | 74.0162
274 | 0.887 | 0.087 | 74.025
278 | 0.797 | 0.072 | 74.0337
282 | 0.677 | 0.0681 | 74.046
286 | 0.551 | 0.0542 | 74.0583
290 | 0.413 | 0.0412 | 74.0705
294 | 0.277 | 0.0263 | 74.0845
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.2008 |  | 171.1851

## d2
Cavity No | Sample | Comment
1 | EPO-TEK-301-3M
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2018-01-27 | 2018-01-31 | 10:00 | 4 | 96.4167 | 2.95
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
25.6 | 6.3e-08 | 25.6 | 44.018
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.474 | 0.097 | 74.0162 | 1.8003e-05 | 1.04e-08
270 | 1.459 | 0.094 | 74.0198
274 | 1.385 | 0.0874 | 74.0268
278 | 1.253 | 0.0776 | 74.0355
282 | 1.057 | 0.0668 | 74.046
286 | 0.862 | 0.0552 | 74.0583
290 | 0.626 | 0.0385 | 74.0722
294 | 0.435 | 0.026 | 74.0845
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.0498 |  | 185.1525

## d3
Cavity No | Sample | Comment
1 | EPO-TEK-301-3M
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2018-01-27 | 2018-02-02 | 8:30 | 6 | 144.3542 | 2.95
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
25.8 | 5.9e-08 | 25.8 | 44.018
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.962 | 0.0736 | 74.0373 | 1.8003e-05 | 2.59e-08
270 | 1.904 | 0.0678 | 74.0408
274 | 1.748 | 0.0659 | 74.0443
278 | 1.557 | 0.0608 | 74.0512
282 | 1.323 | 0.0515 | 74.0583
286 | 1.061 | 0.0432 | 74.0687
290 | 0.793 | 0.0313 | 74.0775
294 | 0.555 | 0.0231 | 74.088
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.2058 |  | 185.1546

## d4
Cavity No | Sample | Comment
1 | EPO-TEK-301-3M
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2018-01-27 | 2018-02-05 | 8:30 | 9 | 216.3542 | 2.95
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
25.8 | 5.5e-08 | 25.8 | 44.018
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.862 | 0.0953 | 74.018 | 1.7994e-05 | 1.06e-08
270 | 1.19 | 0.0922 | 74.0198
274 | 1.776 | 0.0858 | 74.0268
278 | 1.563 | 0.0752 | 74.0355
282 | 1.319 | 0.0642 | 74.0477
286 | 1.065 | 0.0524 | 74.06
290 | 0.804 | 0.0391 | 74.0722
294 | 0.552 | 0.0263 | 74.0845
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.1258 |  | 185.2501

## d5
Cavity No | Sample | Comment
1 | EPO-TEK-301-3M
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2018-01-27 | 2018-02-07 | 9:00 | 11 | 264.375 | 2.95
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
26 | 5.4e-08 | 26 | 44.031
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.787 | 0.094 | 74.0162 | 1.8007e-05 | 1.07e-08
270 | 1.725 | 0.092 | 74.0198
274 | 1.602 | 0.085 | 74.025
278 | 1.431 | 0.0769 | 74.0355
282 | 1.221 | 0.0662 | 74.046
286 | 0.998 | 0.0541 | 74.0583
290 | 0.74 | 0.0401 | 74.0722
294 | 0.509 | 0.0269 | 74.0845
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.2504 |  | 185.1174
Formulas:
Cavity No = [1]d1!A4

## d6
Cavity No | Sample | Comment
1 | EPO-TEK-301-3M
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2018-01-27 | 2018-02-13 | 10:00 | 17 | 408.4167 | 2.95
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
26 | 4.6e-08 | 24.8 | 44.017
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.723 | 0.096 | 74.0127 | 1.7974e-05 | 1.7e-08
270 | 1.705 | 0.093 | 74.0162
274 | 1.602 | 0.087 | 74.0215
278 | 1.489 | 0.0777 | 74.032
282 | 1.269 | 0.0662 | 74.0443
286 | 1.025 | 0.0538 | 74.0565
290 | 0.776 | 0.0403 | 74.0687
294 | 0.55 | 0.0282 | 74.081
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.2176 |  | 185.4566
Formulas:
Cavity No = [1]d1!A4

## d7
Cavity No | Sample | Comment
1 | EPO-TEK-301-3M
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2018-01-27 | 2018-02-16 | 11:50 | 20 | 480.4931 | 2.95
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
21 | 4.1e-08 | 24.5 | 44.009
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.86 | 0.098 | 74.011 | 1.8005e-05 | 1.48e-08
270 | 1.795 | 0.094 | 74.0162
274 | 1.668 | 0.088 | 74.0215
278 | 1.506 | 0.0795 | 74.0302
282 | 1.27 | 0.0672 | 74.0425
286 | 1.025 | 0.0542 | 74.0547
290 | 0.767 | 0.0408 | 74.0687
294 | 0.53 | 0.0277 | 74.0828
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.2112 |  | 185.1302
Formulas:
Cavity No = [1]d1!A4

## d8
Cavity No | Sample | Comment
1 | EPO-TEK-301-3M
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2018-01-27 | 2018-02-21 | 10:00 | 25 | 600.4167 | 2.95
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
26 | 4e-08 | 26 | 44.009
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.837 | 0.094 | 74.0162 | 1.8011e-05 | 1.09e-08
270 | 1.776 | 0.091 | 74.0198
274 | 1.63 | 0.084 | 74.0285
278 | 1.483 | 0.0766 | 74.0355
282 | 1.275 | 0.0651 | 74.046
286 | 1.028 | 0.0522 | 74.0583
290 | 0.788 | 0.0401 | 74.0722
294 | 0.545 | 0.0274 | 74.0828
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.1813 |  | 185.0694
Formulas:
Cavity No = [1]d1!A4

## d9
Cavity No | Sample | Comment
1 | EPO-TEK-301-3M
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2018-01-27 | 2018-02-23 | 9:00 | 27 | 648.375 | 2.95
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
26 | 4e-08 | 26 | 44089
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.041 | 0.0938 | 74.018 | 1.7994e-05 | 1.06e-08
270 | 1.02 | 0.0911 | 74.0215
274 | 0.95 | 0.0853 | 74.0285
278 | 0.837 | 0.0755 | 74.0355
282 | 0.722 | 0.0659 | 74.046
286 | 0.583 | 0.0534 | 74.0583
290 | 0.443 | 0.0406 | 74.0705
294 | 0.296 | 0.027 | 74.0828
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.0784 |  | 185.252
Formulas:
Cavity No = [1]d1!A4

## d10
Cavity No | Sample | Comment
1 | EPO-TEK-301-3M
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2018-01-27 | 2018-02-26 | 7:30 | 30 | 720.3125 | 2.95
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
26 | 3.8e-08 | 26 | 44.009
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.168 | 0.0916 | 74.0215 | 1.7631e-05 | 1.478e-08
270 | 1.094 | 0.0889 | 74.025
274 | 1 | 0.083 | 74.0302
278 | 0.88 | 0.0741 | 74.039
282 | 0.75 | 0.0633 | 74.0495
286 | 0.619 | 0.0517 | 74.06
290 | 0.487 | 0.0377 | 74.074
294 | 0.332 | 0.0262 | 74.0845
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.0615 |  | 189.0594

## d11
Cavity No | Sample | Comment
1 | EPO-TEK-301-3M
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2018-01-27 | 2018-02-28 | 9:00 | 32 | 768.375 | 2.95
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
25.3 | 3.7e-08 | 25.3 | 43.896
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.054 | 0.0914 | 74.018 | 1.803e-05 | 1.61e-08
270 | 1.015 | 0.0888 | 74.0233
274 | 0.936 | 0.0827 | 74.0285
278 | 0.835 | 0.0742 | 74.0373
282 | 0.71 | 0.0636 | 74.0477
286 | 0.581 | 0.0517 | 74.0583
290 | 0.444 | 0.0392 | 74.0705
294 | 0.3 | 0.0266 | 74.0828
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.1134 |  | 184.8745

## d12
Cavity No | Sample | Comment
1 | EPO-TEK-301-3M
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2018-01-27 | 2018-03-02 | 9:00 | 34 | 816.375 | 2.95
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
25.5 | 3.6e-08 | 25.4 | 43.999
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.728 | 0.0918 | 74.0198 | 1.8014e-05 | 1.09e-08
270 | 1.647 | 0.089 | 74.0233
274 | 1.478 | 0.083 | 74.0302
278 | 1.303 | 0.0741 | 74.0373
282 | 1.111 | 0.0639 | 74.0477
286 | 0.884 | 0.0513 | 74.06
290 | 0.664 | 0.039 | 74.0722
294 | 0.448 | 0.0267 | 74.0863
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.1908 |  | 185.0403

## d13
Cavity No | Sample | Comment
1 | EPO-TEK-301-3M
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2018-01-27 | 2018-03-05 | 8:00 | 37 | 888.3333 | 2.95
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
25.6 | 3.5e-08 | 25.6 | 43.992
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.027 | 0.0926 | 74.0198 | 1.8021e-05 | 1.13e-08
270 | 1.043 | 0.09 | 74.0215
274 | 0.994 | 0.0844 | 74.0302
278 | 0.898 | 0.0754 | 74.0373
282 | 0.763 | 0.0645 | 74.0477
286 | 0.612 | 0.0523 | 74.06
290 | 0.455 | 0.0394 | 74.0722
294 | 0.305 | 0.0265 | 74.0828
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.0665 |  | 184.9702

## d14
Cavity No | Sample | Comment
1 | EPO-TEK 301-3M adhesive
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2018-01-27 | 2018-03-07 | 7:30 | 39 | 936.3125 | 2.95
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
25.5 | 3.5e-08 | 25.5 | 43.977
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.977 | 0.0866 | 74.025 | 1.8015e-05 | 1.16e-08
270 | 1.906 | 0.084 | 74.0268
274 | 1.762 | 0.0789 | 74.0337
278 | 1.562 | 0.0707 | 74.0408
282 | 1.272 | 0.059 | 74.0512
286 | 1.033 | 0.0484 | 74.0618
290 | 0.77 | 0.0372 | 74.074
294 | 0.509 | 0.0247 | 74.0863
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.1476 |  | 185.0348

## d15
Cavity No | Sample | Comment
1 | EPO-TEK 301-3M adhesive
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2018-01-27 | 2018-03-09 | 7:30 | 41 | 984.3125 | 2.95
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
25.4 | 3.5e-08 | 25.4 | 43.985
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.941 | 0.0841 | 74.0215 | 1.8093e-05 | 1.35e-08
270 | 1.877 | 0.0815 | 74.0233
274 | 1.75 | 0.0766 | 74.032
278 | 1.556 | 0.0683 | 74.039
282 | 1.33 | 0.0582 | 74.0495
286 | 1.1 | 0.0481 | 74.0583
290 | 0.806 | 0.0349 | 74.0722
294 | 0.563 | 0.0249 | 74.0845
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.32 |  | 184.2369

## d16
Cavity No | Sample | Comment
1 | EPO-TEK 301-3M adhesive
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2018-01-27 | 2018-03-12 | 7:30 | 44 | 1056.3125 | 2.95
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
25.5 | 3.4e-08 | 25.5 | 43.789
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.98 | 0.0924 | 74.018 | 1.8033e-05 | 1.05e-08
270 | 0.953 | 0.09 | 74.0198
274 | 0.89 | 0.0837 | 74.0268
278 | 0.797 | 0.0747 | 74.0355
282 | 0.689 | 0.0645 | 74.046
286 | 0.559 | 0.0524 | 74.0583
290 | 0.426 | 0.0399 | 74.0687
294 | 0.283 | 0.026 | 74.0828
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.1244 |  | 184.8446

## d17
Cavity No | Sample | Comment
1 | EPO-TEK 301-3M adhesive
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2018-01-27 | 2018-03-14 | 8:00 | 46 | 1104.3333 | 2.95
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
25.2 | 3.5e-08 | 25.2 | 43808
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.923 | 0.089 | 74.018 | 1.8055e-05 | 1.17e-08
270 | 1.861 | 0.0874 | 74.0198
274 | 1.727 | 0.081 | 74.0285
278 | 1.545 | 0.0732 | 74.0337
282 | 1.302 | 0.0613 | 74.0477
286 | 1.066 | 0.0515 | 74.0583
290 | 0.794 | 0.0376 | 74.0722
294 | 0.542 | 0.0267 | 74.0828
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.3234 |  | 184.6214

## d18
Cavity No | Sample | Comment
1 | EPO-TEK 301-3M adhesive
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2018-01-27 | 2018-03-16 | 7:30 | 48 | 1152.3125 | 2.95
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
25.4 | 3.2e-08 | 25.4 | 43.795
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.166 | 0.0935 | 74.0162 | 1.8023e-05 | 1.07e-08
270 | 1.137 | 0.0908 | 74.018
274 | 1.069 | 0.0851 | 74.0233
278 | 0.956 | 0.0755 | 74.0337
282 | 0.822 | 0.0652 | 74.0443
286 | 0.665 | 0.053 | 74.0565
290 | 0.5 | 0.0397 | 74.0705
294 | 0.342 | 0.0272 | 74.0828
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.2473 |  | 184.9522

## d19
Cavity No | Sample | Comment
1 | EPO-TEK 301-3M adhesive
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2018-01-27 |  |  | -43127 | -1035048 | 2.95
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264
270
274
278
282
286
290
294
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | #VALUE! |  | #DIV/0!

## d20
Cavity No | Sample | Comment
1 | EPO-TEK 301-3M adhesive
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2018-01-27 |  |  | -43127 | -1035048 | 2.95
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264
270
274
278
282
286
290
294
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | #VALUE! |  | #DIV/0!

## d21
Cavity No | Sample | Comment
1 | EPO-TEK 301-3M adhesive
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2018-01-27 |  |  | -43127 | -1035048 | 2.95
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264
270
274
278
282
286
290
294
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | #VALUE! |  | #DIV/0!