# Tested By: Luis E . Sánchez____			Date: _2-26-19_

# LSC Style Detector Measured Parameters

All transimpedance measurements are referred to plane of the physical output connector and include the effect of the voltage divider created by the 50 Ω termination.  All notch rejection ratios are relative to the magnitude of the transimpedance at the respective RF detection center frequency of the given RF output port.

| Parameter | Value |
| --- | --- |
| Detector serial number | S1900052 |
| Detector schematic D# and revision | D1101124-v8 Modified |
| Diode element manufacturer’s serial number | C30642G |
| Quiescent DC current (amps at +18 VDC) | 60mA |
| Quiescent DC current (amps at -18 VDC) | 90mA |
| PD bias regulator output voltage (VDC) | +5.04v |
| RF opamp positive voltage regulator (VDC) | +5.84v |
| RF opamp negative voltage regulator (VDC) | -6.04v |
| Audio opamp positive voltage regulator (VDC) | +14.77v |
| Audio opamp negative voltage regulator (VDC) | -15.22v |
| DC path transimpedance and zero light offset (Ω/mVDC at BNC out) | 96Ω | 0.096mVDC |
| DC path transimpedance and zero light offset (Ω/mVDC at differential out) | 196Ω | 0.196mVDC |
| DC path zero frequency (Hz) | 0.1Hz |
| DC path pole frequency (Hz) | 2.4Hz |
| Inferred DC path shot noise limited input photo sensitivity (mA) at 100Hz measured at differential output | 3.7mA |
| RF detection center frequency (MHz), f low | 6.25MHz |
| RF detection center frequency (MHz), f hi |  |
| Notch frequencies (MHz) used in design |  |
| F low feedback notch frequency |  |
| F hi feedback notch frequency |  |
| Rejection (dB) at notch1 (f low) |  |
| Rejection (dB) at notch2 (f low) |  |
| Rejection (dB) at notch3 (f low) |  |
| Rejection (dB) at notch4 (f low) |  |
| Rejection (dB) f low to f hi |  |
| Rejection (dB) at notch1 (f hi) |  |
| Rejection (dB) at notch2 (f hi) |  |
| Rejection (dB) at notch3 (f hi) |  |
| Rejection (dB) at notch4 (f hi) |  |
| Rejection (dB) f hi to f low |  |
| Transimpedance (Ω) at f low (note PD Current) | Based on Q 6.298KΩ | mA |
| RF dark/light noise used for f low Trans-Z | -128.8dBm/Hz | dBm/Hz |
| Transimpedance (Ω) at f hi (note PD Current) | Ω | mA |
| RF dark/light noise used for f hi Trans-Z | dBm/Hz | dBm/Hz |
| RF preamp used during testing (noise/gain) | dBm/Hz | dB |
| f low, shot-noise limited input sensitivity (mA) | Calculated 20.8µA/Hz |
| f hi, shot-noise limited input sensitivity (mA) |  |
| Test input transconductance at f1(mA/V) |  |
| Test switch isolation at f1 (dB) |  |
| Test input transconductance at f2(mA/V) |  |
| Test switch isolation at f2 (dB) |  |

For calculations please see document T1900075.