## Slide 1: Calibrating thE LIGO Interferometers: Third Time’s a Charm
J. Kissel, for the LIGO Calibration Team and the LIGO Scientific Collabortation

## Slide 2: This morning: The DARM Loop
Interferometer is non-linear unless extremely well controlled.
Detector readout is fed to control differential arm lengths.
To produce an estimate of uncontrolled input, h, we must have an exquisite model of the loop.
Interferometer
Sensor, C
Multi-stage
Actuator, A

## Slide 3: This morning: Absolute Reference
Let’s assume that …
Sudarshan told you everything about using photon radiation pressure as our absolute reference,
the PCAL team has delivered us a perfect reference (no systematic error, but still with statistical uncertainty)
we have digital signal, pdpc (the digitized voltage from receiver module’s integrating sphere and PD) that’s been converted to displacement of the test mass, xpc .
O2 Uncertainty = 0.76 %
O3 Uncertainty = 0.5%

## Slide 4: OVERALL Uncertainty
G = A D C
R ≡
1 + G
C
_________
L
1
__
h =
R
derr
The model, and therefore the estimated detector input, has uncertainty and systematic error:
∂ h 2 ≈ ∂ R 2 ≈
∂ A
A
____
1 + G
_________
1
(
)
1 + G
_________
G
(
)
2
∂ C
C
____
)
2
(
+
2
(
)
2
+ Systematic Errors
h L =
C(model)
____
1
derr
dctrl
+
A(model)

## Slide 5: The problem Gets HARDer…
Our measurements of C and A are
ratios of complex transfer functions
>> the frequency-dependent, magnitude and phase are important for A and C
With all loops closed and the detector running at its best sensitivity, we request a series of in-loop excitations to obtain direct measurements of the sensor and actuator
C =
dIN2
____
dexc
pcexc
____
derr
pdpc
____
xpc
x
Ai =
ai,exc
____
derr
pcexc
____
derr
pdpc
____
xpc
x
Complex Transfer
Functions w/ Full IFO
Absolute Scale Factor
1 + G
_________
1
)
(
1 + G
_________
C
)
(
1 + G
_________
Ai C
)
(
1 + G
_________
C
)
(
x
x

## Slide 6: The Simple version…
Evan mentioned the “free” parameters that dominate the uncertainty in A and C.
Here’s what you normally tell people in a talk:
gain 1 / f 4
gain 1 / f 6
gain 1 / f 2
x
x
x
gain
1 + i f / fc
______________
1
)
(
fc =“cavity pole
frequency”
x
Why focus on these?
(they’re “easily” accessible to the “average” audience)
They can only be measured with the interferometer running.

## Slide 7: …But there’s MorE
IFO measurements contain more than just simple functions.
There are WAY more details in
  C’s analog-to-digital conversation path
  A’s digital-to-analog path
that must be measured and modelled, in order to obtain a “clean” result that can be fit for free parameters
For C (two paths of):
  Detuning of optical plant
  PD trans-impedance amplifiers
  analog whitening filters,
  analog anti-aliasing filters,
  computational delays,
  digital down-sampling filters
For A (three stages of…):
digital up-sampling filters,
computational delays
analog anti-imaging filters
analog low-pass filters
trans-conductance/voltage drivers
drive signal to force actuators
non-trivial force to displacement dynamics of multi-stage pendula
… all must be included with systematic error quantifiably negligible in the model

## Slide 8: Assuming we’ve got that under control…
C =
dIN2
____
dexc
pcexc
____
derr
pdpc
____
xpc
x
A =
ai,exc
____
derr
pcexc
____
derr
pdpc
____
xpc
x
Complex Transfer
Functions w/ Full IFO
Absolute Scale Factor
x
x
Back to the full interferometer measurements…

## Slide 9: Example Measurement
Reconstructed
PCAL RX PD xpc
Reconstructed,
Live, DL
Cartoon
Target DL
“Swept Sine” Excitation
pcexc
____
derr
Taken during lowest noise state, in loop
Two measurements per function
Actuator strength vs. Detector Noise limited!
This is what it looks like when we measure

## Slide 10: Limitations
Absolute reference limits where detector is most sensitive
Actuator strength limits where noise is worse
Black Holes and Pulsars
Neutron Stars and Supernovae
pcexc
____
derr

## Slide 11: Uncertainty Component Breakdown
Gravitational Wave Band
10%
Actuator Model, A
Sensor Model, C

## Slide 12: Current Status, A
PUM Matters HERE
~1 / f 4
Systematic error we must fix!
Violins
Huh?
Known delays and HF electronics response
= 10 deg phase loss at 200 Hz…
PUM Matters HERE

## Slide 13: Statistical Uncertainty
there are 17 free parameters per IFO:
5 sensing function parameters
2 actuation parameters for each stage
5 parameters that we track as a function of time
1 parameter for PCAL uncertainty
So, we numerically evaluate the uncertainty on R
R =
1 + ADC
C
_________
L
1
__
Each has a posterior distribution, and a well-defined contribution to the response
example results for 5 sensing function parameters

## Slide 14: How to Handle Unknown Systematic Error?
Gaussian Process Regression (GPR):
It’s black magic, like any other transfer function fitting program
but importantly it gives you a posterior distribution of curves, do be sampled for numerical evaluation of total uncertainty in R
Instead of concatenating all the residual systematic errors from every single part of each function, we divide out the model from our measurement
Remaining frequency dependence is unknown systematic error, for which we can fit
example GPR fit to Sensing Residual
4 more distributions of free parameter functions…

## Slide 15: The Final Answer (Example)
Uncertainty is informed by individual, numerically evaluated draws from the collection of posterior distributions, divided by “nominal” model (model of maximum a posteriori values, and assuming no systematic error )
The 68% confidence interval bounds at each frequency are the response function’s frequency dependent uncertainty and systematic error.
Individual
Draws of
(R + ∂ R)i / R
68% C.I.
Median
∂ h 2 ≈ ∂ R 2 ≈
∂ A
A
____
1 + G
_________
1
(
)
1 + G
_________
G
(
)
2
∂ C
C
____
)
2
(
+
2
(
)
2
+ Systematic Errors

## Slide 16: We were able to get to “fundamental” limit last time…
68% C.I. for H1
68% C.I. for L1
68% C.I. for PCAL

## Slide 17: Where To?
We’re using photon radiation pressure to great success.
Improving PCAL uncertainty will directly improve interferometer response uncertainty, and thus astrophysical parameter uncertainty
We’re (on our 3rd round of) reducing as much of our other systematic error such that we can reach the PCAL’s “fundamental” limit (again)
Coordinate standards in the global GW Network
Let us know if there’s a systematic error in the estimate of the power…

## Slide 18: Thank You

## Slide 19: Bonus Material

## Slide 20: Current Status, C
C Matters HERE
Cavity Pole
Optical Plant Detuning
Parametric Fit Results

## Slide 21: Why Is it Different Between Runs?
The interferometers are constantly evolving between (and during!) runs to improve the noise, and due to reality of experiment!
Change for Noise between O2 and O3 | Consequence For Calibration
New Test Masses | new force to displacement dynamical model
Higher power for O3 | more complex interferometer response
Loss on optics accrued from vent | more complex interferometer response
Better sensor electronics for sensitivity features | New measurement of electronics
Better actuator electronics to reduce impact of DAC noise | New measurement of electronics
One of the O2 actuators are broken | New actuator scheme model
The calibrator’s job is never finished!

## Slide 22: Why Bother Being FullY Bayesian?
In detection era, it was safe to just quote a maximum in magnitude and phase “XX % / YY deg,” because detection is relatively insensitive to calibration uncertainty
In the observational era, astrophysical parameters depend ~proportionally to calibration uncertainty.
Astrophysical parameter estimation is fully Bayesian, with its own set of 14 parameters
New, on-going project: study the subtle interactions between astrophysical parameters and physical interferometer parameters.
  Can better marginalize (integrate) over Gaussian distribution of real uncertainty, instead of naïve box-car distribution
  Can retrace steps to ask “I want to improve estimation of astro-param M, so please improve estimation of interferometer-param N”

## Slide 23: LIGO Calibration References
Viets, A. D., et al. "Reconstructing the calibrated strain signal in the Advanced LIGO detectors." Classical and Quantum Gravity 35.9 (2018): 095015. P1700236
Cahillane, Craig, et al. "Calibration uncertainty for Advanced LIGO’s first and second observing runs." Physical Review D 96.10 (2017): 102001. P1600139
Abbott, B. P., and LIGO Scientific Collaboration. "Calibration of the Advanced LIGO detectors for the discovery of the binary black-hole merger GW150914." Phys. Rev. D 95, 062003 (2017). P1500248
Karki, S., et al. "The Advanced LIGO photon calibrators." Review of Scientific Instruments 87.11 (2016): 114503. P1500249