## Slide 1: Laser Power Meter Calibrations at NIST
Matthew.Spidell@NIST.GOV

## Slide 2
LIGO Gold Standard traceability to the SI
Details of the NIST Standard Calorimeter
Verification of the NIST standard
LIGO Gold Standard test conditions
Incident optical power inequivalence contributions

## Slide 3: LIGO Gold Standard traceability to the SI

## Slide 4: Representation of the optical Watt through electrical substitution
Defining Constants
Base/Derived Unit
V
e
h
Δν
Ω
s
Primary Standard
C-Calorimeter

## Slide 5: Details of the NIST Standard Calorimeter

## Slide 6: Calorimeter Design
Insulating shell
DC Amplifier
Isothermal vacuum shell
Carbon-Black absorbing cavity
Electrical-Substitution heater windings
Multi-Junction thermopile/Weak thermal link
1° Wedge Window
Heater windings
Thermal dampening shell

## Slide 7: Systematic uncertainties
Source | Standard Uncertainty (type) | 
Standard meter ratio | 0.29% | ()
Inequivalence | 0.087% | ()
Window Transmission | 0.064% | ()
Electronics | 0.058% | ()
Inject time | 0.029% | ()
Absorptivity | 0.0058% | ()
Heater Leads | 0.0058% | ()
Electronics | 0.0033% | ()(N=30)
Window Transmission | 0.0033% | ()(N=30)
Source | Standard Uncertainty (type) | 
Standard meter ratio | 0.29% | 
Inequivalence | 0.087% | 
Window Transmission | 0.064% | 
Electronics | 0.058% | 
Inject time | 0.029% | 
Absorptivity | 0.0058% | 
Heater Leads | 0.0058% | 
Electronics | 0.0033% | 
Window Transmission | 0.0033% | 

## Slide 8: Systematic Uncertainty Details
Standard meter ratio (0.29 %)
  Uncertainty attributable to pointing instability integrated with last external aperture and /or standard entrance aperture
Inequivalence (0.087 %)
  Difference in temperature response between optically delivered power and electrically delivered power (during calibration of the standard)
Window Transmission (0.064%)
  Composite polarization and spectral-absorptivity uncertainty
Electronics (0.058 %)
  Composite of the pre-amplifier and A-D converter uncertainty contributions

## Slide 9: Verification of the NIST standard

## Slide 10: Optical inter-comparisons
Laser Optimized Cryogenic Radiometer
  1 mW
  1 μm
  U (k=2) 0.86%

## Slide 11
Chart: C vs LOCR Transfer Standard Calibration Ratio
| 2007 | 2008 | 2009 | 2010 | 2015 | 2016 | 2017 | 2018
 | 1.0026 | 0.99433 | 0.99986 | 1.0069 | 0.99691 | 0.99701 | 1.0039 | 1.0005

## Slide 12: Optical inter-comparisons
K-Series Calorimeter
  2 W
  1 μm
  (C.F.= 0.9981)
  U (k=2) 1.5%

## Slide 13: Kg inter-comparison (via Radiation Pressure Power Meter & K-series calorimeter)
Kg via RPPM & K-Series Calorimeter
  (C.F.= 0.9941)
  U (k=2) 1.7%

## Slide 14: LIGO Gold Standard test conditions

## Slide 15: LIGO ‘Gold Standard’ Calibration Conditions
1047nm Diode Pumped Solid State Source
1/e2 diameter of approximately 4 mm
24 Hour+ thermal equilibrium
20 ± 1 °C
20 ± 5 % RH

## Slide 16: Measurement Procedures
Order of Operations:
  Calorimeter vs monitor
  Gold Standard vs monitor
  Calorimeter vs monitor
Each step consists of 4+ measurements to meet minimum degrees of freedom for a 96% confidence interval
DUT
Laser
Shutter
Calorimeter
Calorimeter
Data Acquisition and Control
Beamsplitter

## Slide 17: Detailed Calibration Geometry
GS
Shutter
Attenuator
Entrance Iris
M1
M2
Standard 2
Standard 1
Beamsplitter
Enclosure
Enclosure

## Slide 18: Inequivalence contributors to Standard and Device Under Test incident power

## Slide 19: Incident optical power inequivalence contributions
Setup geometry
  Intermediate optics (scatter, diffraction, alignment, etc…)
  Long paths between last optical element and standards preferred
    renders apertures effective against scatter
  Minimum 30 cm separation between integrating sphere and any normal surface (more on this latter)
Unintended etalons
  Avoid planar beamsplitters (to monitor)
  Avoid attenuators after beamsplitter
  example:
    Assume planar surfaces
    R=0.5%

## Slide 20: Incident optical power inequivalence contributions (cont.)
Laser induced heating of optics
  Absorbing ND filters and invar film attenuators can see laser induced heating and resulting transmission variation
Continually-variable attenuators after the beamsplitter (to monitor) yield splitting ratio variation due to pointing instability
Atmospheric absorption
  Typically water lines in Short Wave IR
  Negligible at 1047nm, 20 cm gas cell
Source instability
  Active power stabilization or Concurrent measurement of a monitor and Standard/Device under test preferred

## Slide 21: An Unintended Etalon
Assume planar surfaces
R=0.5%
Chart: Relative transmittance of an Anti-Reflection coated 3mm thick glass filter element
| 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53
 | 0.98543 | 0.9879 | 0.99054 | 0.99315 | 0.99554 | 0.99755 | 0.99902 | 0.99984 | 0.99996 | 0.99935 | 0.99808 | 0.99623 | 0.99393 | 0.99137 | 0.98873 | 0.98619 | 0.98393 | 0.98211 | 0.98085 | 0.98025 | 0.98033 | 0.9811 | 0.9825 | 0.98444 | 0.98678 | 0.98937 | 0.99201 | 0.99452 | 0.99672 | 0.99844 | 0.99956 | 1 | 0.99971 | 0.99872

## Slide 22: Beam cameras
Unreliable for defining ‘maximum extent’ beam diameter since significant energy may exist in large diameter “margin” of the beam
Gain conditions may suppress the signal in the “margin”
  May miss diffraction spikes and interference fringes
Significant variation between manufacturers
Easily damaged at the focus of even a low powered beam

## Slide 23: A special caution regarding integrating sphere calibrations

## Slide 24: Test Setup
Aperture stop typically located before standard and device under test to ensure scattered light does not yield inequivalence
Aperture diameter must exceed main beam diameter to bar against introducing edge diffraction

## Slide 25: Test Setup
Reflectance of PTFE >98% at 1047 nm
Integrating spheres can have a diffuse exitance of approximately 50%
Worst case conditions yielded a 2% increase in detected signal
Aperture stops must be located >30 cm from the detector entrance port (empirically determined)
Detector at top of sphere
Sphere Shell
Diffuse reflectance wall material (PTFE)
50% Exitance

## Slide 26: What grants a National Metrology Agency authority to disseminate standards?

## Slide 27: CIPM Mutual Recognition Arrangement (MRA)
Provides open, transparent, and comprehensive scheme to assess:
  Comparability of national metrology standards and services
  Technical basis for wider agreements for international trade, commerce, and regulatory affairs
Notes: Driven by emergence of more globalized economy
October 1999: The directors of the national metrology institutes (NMIs) of 38 Member States of the Metre Convention signed a Mutual Recognition Arrangement (MRA) for national measurement standards and for calibration and measurement certificates issued by national metrology institutes.
Today, 101 institutes have signed, covering 101 institutes plus 153 designated institutes
MRA drawn up by International Committee of Weights and Measures (CIPM) under authority of Metre Convention
By signing, accept the scheme for assessing comparability and recognize results/outcomes of the scheme

## Slide 28: Regional Metrology Organizations RMOs
Run key comparisons and supplementary comparisons to support the CIPM MRA
Validate NMI Quality Systems
Review NMI Calibration and Measurement Capability Statements (CMCs)

## Slide 29: Components for international recognition/approval of measurement capabilities
Signing of the CIPM Mutual Recognition Agreement
Peer-Reviewed Quality Management System
Inter-comparisons

## Slide 30: The Americas: Sistema Interamericano de Metrologia
5 Year review cycle
Past and planned key comparisons conducted with low power (100 μW) high precision cryogenic radiometers

## Slide 31: What about ISO 17043
NIST quality system follows ISO 17043 guidelines
Not reviewed by a third party corporation
  SIM supersedes such entities