## how to use this workbook
How to use this workbook:
1) Get an E-number for this test report from https://dcc.ligo.org
a) Label the document "Optical Contamination Test Results: <name>", where <name> is a descriptive name of the article under test
b) For the E-<number>-X0 version DCC entry, give a COMPLETE description of the article(s) under test in the Abstract (not the "Changes and Notes"), including manufacturer/source, material grade or assembly/device model number, and quantity/amount. For example, it is not sufficient to say "photodiode", but rather state the manufacturer's name and model number, such as "Vishay TSTS7100 IR-LED". Likewise "PEEK" is not a sufficient material descriptor, but "Victrex PEEK grade 450GL" is a complete description.
c) For the E-<number>-X0 version DCC entry, give a URL link to the location where intermediate test results will be posted in the "Notes and Changes" field.
2) Rename this template workbook to "E<number>-v1 <name>.xls"
3) Fill in only the light yellow colored cells in the sheet labelled "scaled Loss rate".
You must decide the appropriate units to characterize the material, or assembly or device under test. For example, for a device or assembly the appropriate units is likely "each", whereas for a bulk material (say PEEK) the appropriate unit is likely surface area (cm^2). For a lubricant, such as Krytox, the appropriate units may be mass, or volume.
4) Fill in only the light yellow colored cells in the sheets labelled "d1", "d2", "d3", etc.
Each sheet is a data set taken at a particular day and time. In the first data set sheet (d1), the start date of the test and the cavity number is entered. Data for the TEM 00-01 beat frequency as a function of power and the ring down data are entered for each data set. The data is plotted and fitted automatically. The sheet labelled "data" summarizes the absorption and total loss measurements as a function of time in a tabular listing. Sheets "absorption" and "total loss" summarize the absortion and total loss in plots and a least squares ft, as the data is entered. The "scaled loss rate" sheet summarizes the results scaled to the LIGO Observatory for the intended use and is updated automatically as data is entered.
5) The test is terminated when a determination is made that the material/assembly/device under test has either passed or failed. At this point, use the "Convert to Adobe PDF" tool to create a file with just the spredsheets labelled "scaled loss rate", "absorption" and "total loss" which is named "E<number>-v1 <name>.pdf".
6) Select "Create E<number>-v1" in the DCC entry for this document and upload the pdf file "E<number>-v1 <name>.pdf" as the "Main file in document" and this Excel file "E<number>-v1 <name>.xls" as an "Other file" (not Main). Then in either the abstract or the "Notes and Changes" field indicate the test results (passed or failed).
7) File any other documents, which help to define the material/assembly/device which was tested, as "Other files" under the DCC entry for E<number>-v1, such as photos or descriptions of the test (if it was a unique test, such as taken at elevated temperatures).
8) File proprietary data (data not owned by LIGO Lab, even if posted to the web) about the material/assembly/device such as MSDS sheets, specifications, etc. This data should be filed under a separate E-number and then linked from the E-number which corresponds to the optical contamination cavity test results. Even if the MSDS, specifications, etc. are already on the web, they should be filed in the DCC, in case the material/assembly/device is discontinued (in which case the manufacturer often removes this data from their web sites.)

## scaled loss rate
LIGO E-Document Number: |  |  | E1900080
Sample Test:
| Material under test: |  | SR3 ROC Actuator, Ceramic Heater Assy (NOT THE ENTIRE ACTUATOR ASSY.)
| units |  |  | 1 | unit
| absorption | 2.3131 | ± | 0.22919 | ppm/yr | 1 sigma
| scatter | 9.6568 | ± | 4.3372 | ppm/yr | 1 sigma | from 2nd set of 87C data after transient/instability
| max. normalized absorption |  |  | 2.7715 | ppm/yr/unit | 2 sigma
| max. normalized scatter |  |  | 18.3312 | ppm/yr/unit | 2 sigma
| test turbopump speed (liter/s) |  |  | 24.3916 | torr/liter/sec
Scaled to LIGO:
| LIGO Vacuum Volume |  |  | Vertex | LHO Diagonal | End | Comments
| Quantity (units) |  |  | 1 | 0 | 0
| LIGO ion pumping speed (liter/s) |  |  | 6800 | 6800 | 1700 | see E0900398 or PSI V049-1-078 for pump rates
| pumping speed ratio (test/LIGO) |  |  | 0.003587 | 0.003587 | 0.014348 | does not include cryo-pump and effective pumping from the Beam Tube
| max. absorption (ppm/yr) |  |  | 0.0099413 | 0 | 0 | * Limit is < 0.02 ppm/yr for a single source
| max. scatter (ppm/yr) |  |  | 0.065754 | 0 | 0 | * Limit is < 0.2 ppm/yr for a single source
* The overall limit on contamination loss on optics for AdL is < 0.5 ppm/yr absorption and < 4 ppm/yr scatter from all sources, per Table 4 of the COC Design Requirements Document (T000127-v1). It is assumed that ~20 sources could contribute.
|  |  |  | 2.75 | in, dia, flex hose
|  |  |  | 36 | in, length, flex hose
|  |  |  | 6.985 | cm, dia
|  |  |  | 91.44 | cm, length
|  |  |  | 463.669 | C1, aperture, liter/sec
|  |  |  | 45.0971 | C2, flex hose, liter/sec
|  |  |  | 41.0997 | Ct, liter/sec
|  |  |  | 60 | S, liter/sec [Varian Dry TPS-compact with TV81M]
|  |  |  | 24.3916 | Seff, liter/sec

## absorption
LIGO E-Document Number: |  | E1900080
Test Material/Assy./Device: |  | SR3 ROC Actuator, Ceramic Heater Assy (NOT THE ENTIRE ACTUATOR ASSY.)
Absorption fitting
| Slope | 0.00026405 | 3.5535 | Y-intercept
| Standard Error | 2.6164e-05 | 0.075066 | Standard error
| r2 | 0.67969 | 0.26299 | sey
| F | 101.8548 | 48 | df
| ssreg | 7.0448 | 3.3199 | ssresid
| Absorption change rate (ppm/yr) |  | ± sigma (ppm/yr)
| 2.3131 |  | 0.22919

## total loss
LIGO E-Document Number: |  | E1900080
Test Material/Assy./Device: |  | SR3 ROC Actuator, Ceramic Heater Assy (NOT THE ENTIRE ACTUATOR ASSY.)
First 4 data points are at room temperature
| Fit of data points 5 through 30 (until apparently unstable results) |  |  |  | Fit of data pts 5 - 15, heater @87C (except first two pts at 67C and 80C) |  |  |  | Fit of data points 16-23, heater at 120C
| Total loss fitting |  |  |  | Total loss fitting |  |  |  | Total loss fitting
| Slope | -0.00010054 | 156.1119 | Y-intercept | Slope | 0.00055839 | 155.2629 | Y-intercept | Slope | -0.00030783 | 156.421 | Y-intercept
| Standard Error | 7.1515e-05 | 0.12596 | Standard error | Standard Error | 0.00015492 | 0.14692 | Standard error | Standard Error | 0.00019685 | 0.37494 | Standard error
| r2 | 0.065938 | 0.34878 | sey | r2 | 0.59074 | 0.15122 | sey | r2 | 0.28955 | 0.17548 | sey
| F | 1.9766 | 28 | df | F | 12.9911 | 9 | df | F | 2.4453 | 6 | df
| ssreg | 0.24045 | 3.4061 | ssresid | ssreg | 0.29708 | 0.20581 | ssresid | ssreg | 0.075295 | 0.18475 | ssresid
| Total loss change rate (ppm/yr) |  | ± sigma (ppm/yr) |  | Total loss change rate (ppm/yr) |  | ± sigma (ppm/yr) |  | Total loss change rate (ppm/yr) |  | ± sigma (ppm/yr)
| -0.88076 |  | 0.62647 |  | 4.8915 |  | 1.3571 |  | -2.6966 |  | 1.7244
| time | loss |  |  | time | loss |  |  | time | loss
| 456.4167 | 156.066 | trend/fit line |  | 456.4167 | 155.5178 | trend/fit line |  | 1464.3889 | 155.9702 | trend/fit line
| 3024.5833 | 155.8078 |  |  | 1368.3646 | 156.027 |  |  | 2496.5417 | 155.6525
Fit of data points 37-53, heater at 87C (2nd time)
Total loss fitting
[9 empty] | Slope | 0.0011024 | 130.6323 | Y-intercept
[9 empty] | Standard Error | 0.00049511 | 2.1884 | Standard error
[9 empty] | r2 | 0.2484 | 1.0543 | sey
[9 empty] | F | 4.9574 | 15 | df
[9 empty] | ssreg | 5.51 | 16.672 | ssresid
[9 empty] | Total loss change rate (ppm/yr) |  | ± sigma (ppm/yr)
[9 empty] | 9.6568 |  | 4.3372
[9 empty] | time | loss
[9 empty] | 3552.5417 | 134.5486 | trend/fit line
[9 empty] | 5208.5833 | 136.3741

## data
| Time (hr) | Absorption (ppm) | Total Loss (ppm) | Temp (deg C) | SR3 ROC Actuator, Ceramic Heater Assy (NOT THE ENTIRE ACTUATOR ASSY.)
1 | 288.3542 | 3.6645 | 156.7926 | 22.5
2 | 312.3958 | 3.6447 | 156.6483 | 21.9
3 | 336.375 | 3.713 | 156.7062 | 21.9
4 | 360.5417 | 3.5979 | 156.7411 | 21.9
5 | 456.4167 | 3.6388 | 155.7378 | 67.9
6 | 504.7222 | 3.7108 | 155.5043 | 80.4
7 | 624.5208 | 3.6603 | 155.6622 | 86.7
8 | 696.375 | 3.638 | 155.6097 | 86.9
9 | 840.6042 | 3.6755 | 155.6359 | 86.8
10 | 936.4375 | 3.4456 | 155.4644 | 86.8
11 | 984.3125 | 3.4397 | 155.7346 | 86.7
12 | 1032.4375 | 3.7704 | 155.9537 | 86.1
13 | 1176.4583 | 3.9428 | 155.915 | 86
14 | 1296.4375 | 3.8931 | 156.088 | 86.1
15 | 1368.3646 | 3.892 | 156.1244 | 86.4 Start - 89.3 End
16 | 1464.3889 | 3.9438 | 156.168 | 120.9
17 | 1512.3021 | 3.4397 | 156.0047 | 120.1
18 | 1632.4271 | 4.5693 | 155.8394 | 119.3
19 | 1776.625 | 4.6826 | 155.8203 | 119.3
20 | 1968.3542 | 4.5502 | 155.6494 | 119.2
21 | 2040.3854 | 4.3003 | 155.5467 | 119.5
22 | 2304.3958 | 4.3629 | 155.8891 | 119.4
23 | 2328.4167 | 4.3224 | 155.8246 | 119.4
24 | 2496.5417 | 4.354 | 155.9613 | 85.4
25 | 2544.3125 | 4.3804 | 155.9584 | 86.2
26 | 2664.3333 | 4.3209 | 156.0643 | 86.3
27 | 2784.5208 | 4.2455 | 155.871 | 86.5
28 | 2856.5521 | 4.3049 | 155.8797 | 87.1
29 | 2976.3958 | 4.4233 | 156.0015 | 87.2
30 | 3024.5833 | 4.4017 | 155.9772 | 87.2
31 | 3120.4375 | 4.3852 | 147.8332 | 87.3 | unstable?
32 | 3168.6042 | 4.4269 | 142.3753 | 87.4 | unstable?
33 | 3216.5417 | 4.3916 | 141.1689 | 87.2 | unstable?
34 | 3312.5417 | 4.3363 | 140.7652 | 87.2 | unstable?
35 | 3456.5833 | 4.1822 | 137.6531 | 87.2 | unstable?
36 | 3504.3194 | 4.421 | 138.4118 | 87.2 | unstable?
37 | 3552.5417 | 4.0331 | 134.4035 | 87.2
38 | 3648.5417 | 4.2715 | 134.2589 | 87.1
39 | 3696.5833 | 4.287 | 134.4415 | 87.1
40 | 3792.5833 | 4.0861 | 134.3941 | 87.1
41 | 3840.5833 | 4.2467 | 133.8696 | 87.3
42 | 4224.5833 | 4.4253 | 138.1548 | 88.1
43 | 4320.625 | 4.3905 | 134.8407 | 88
44 | 4368.625 | 4.8644 | 135.7846 | 88.2
45 | 4464.5417 | 5.2686 | 135.0234 | 88.3
46 | 4512.5417 | 5.233 | 135.2009 | 88.2
47 | 4560.5417 | 4.7705 | 136.8602 | 88.3
48 | 4680.5417 | 4.7378 | 135.5494 | 88.2
49 | 4752.5556 | 5.1499 | 136.8432 | 88.2
50 | 4824.4792 | 4.7395 | 137.1082 | 88.2
51 | 5016.5833 | 4.9886 | 135.0851 | 88.3
52 | 5160.5833 | 4.7085 | 135.3369 | 88.2
53 | 5208.5833 | 4.6241 | 135.8601 | 88.2
54 | #N/A | #N/A | #N/A | #N/A
55 | #N/A | #N/A | #N/A | #N/A
56 | #N/A | #N/A | #N/A | #N/A
57 | #N/A | #N/A | #N/A | #N/A
58 | #N/A | #N/A | #N/A | #N/A
59 | #N/A | #N/A | #N/A | #N/A
60 | #N/A | #N/A | #N/A | #N/A
61 | #N/A | #N/A | #N/A | #N/A
62 | #N/A | #N/A | #N/A | #N/A

## Least Sqr Fitting
std (ppm)
|  | 0.2 | error | error | error
histogram | x | y | LS slope | LS slope | LS slope/std
bins | day | ppm | (ppm/day) | (ppm/yr) | (ppm/yr)/ppm
-1 | 12 | 3.6645
-0.9 | 13 | 3.6447 | -0.019791 | -0.47497 | -2.3749
-0.8 | 14 | 3.713 | 0.024277 | 0.58266 | 2.9133
-0.7 | 15 | 3.5979 | -0.013135 | -0.31523 | -1.5762
-0.6 | 19 | 3.6388 | -0.0046993 | -0.11278 | -0.56391
-0.5 | 21 | 3.7108 | 0.0028039 | 0.067294 | 0.33647
-0.4 | 26 | 3.6603 | 0.0010724 | 0.025737 | 0.12869
-0.3 | 29 | 3.638 | -0.00027635 | -0.0066324 | -0.033162
-0.2 | 35 | 3.6755 | 0.00033153 | 0.0079567 | 0.039784
-0.1 | 39 | 3.4456 | -0.0041351 | -0.099244 | -0.49622
-1.3878e-16 | 41 | 3.4397 | -0.0059458 | -0.1427 | -0.71349
0.1 | 43 | 3.7704 | -0.0028452 | -0.068285 | -0.34142
0.2 | 49 | 3.9428 | 0.001259 | 0.030215 | 0.15108
0.3 | 54 | 3.8931 | 0.0031555 | 0.075732 | 0.37866
0.4 | 57 | 3.892 | 0.0041625 | 0.0999 | 0.4995
0.5 | 61 | 3.9438 | 0.0050374 | 0.1209 | 0.60449
0.6 | 63 | 3.4397 | 0.0027076 | 0.064982 | 0.32491
0.7 | 68 | 4.5693 | 0.0067869 | 0.16289 | 0.81443
0.8 | 74 | 4.6826 | 0.010092 | 0.2422 | 1.211
0.9 | 82 | 4.5502 | 0.011547 | 0.27713 | 1.3856
1 | 85 | 4.3003 | 0.011409 | 0.27381 | 1.3691
| 96 | 4.3629 | 0.011071 | 0.2657 | 1.3285
| 97 | 4.3224 | 0.010701 | 0.25683 | 1.2842
| 104 | 4.354 | 0.010306 | 0.24734 | 1.2367
| 106 | 4.3804 | 0.010045 | 0.24107 | 1.2054
| 111 | 4.3209 | 0.009605 | 0.23052 | 1.1526
| 116 | 4.2455 | 0.009018 | 0.21643 | 1.0822
| 119 | 4.3049 | 0.0086322 | 0.20717 | 1.0359
| 124 | 4.4233 | 0.008452 | 0.20285 | 1.0142 | Bin | Frequency
| 126 | 4.4017 | 0.0082544 | 0.1981 | 0.99052 | -1 | 0
| 130 | 4.3852 | 0.0080236 | 0.19257 | 0.96283 | -0.9 | 0
| 132 | 4.4269 | 0.0078722 | 0.18893 | 0.94466 | -0.8 | 0
| 134 | 4.3916 | 0.007688 | 0.18451 | 0.92256 | -0.7 | 0
| 138 | 4.3363 | 0.0074343 | 0.17842 | 0.89211 | -0.6 | 0
| 144 | 4.1822 | 0.0069937 | 0.16785 | 0.83924 | -0.5 | 0
| 146 | 4.421 | 0.0068586 | 0.16461 | 0.82303 | -0.4 | 0
| 148 | 4.0331 | 0.0063631 | 0.15272 | 0.76358 | -0.3 | 1
| 152 | 4.2715 | 0.0061269 | 0.14705 | 0.73523 | -0.2 | 9
| 154 | 4.287 | 0.0059287 | 0.14229 | 0.71145 | -0.1 | 8
| 158 | 4.0861 | 0.0055693 | 0.13366 | 0.66831 | -1.3878e-16 | 10
| 160 | 4.2467 | 0.005376 | 0.12902 | 0.64512 | 0.1 | 10
| 176 | 4.4253 | 0.0052608 | 0.12626 | 0.63129 | 0.2 | 12
| 180 | 4.3905 | 0.0051178 | 0.12283 | 0.61414 | 0.3 | 4
| 182 | 4.8644 | 0.0053438 | 0.12825 | 0.64126 | 0.4 | 4
| 186 | 5.2686 | 0.0058118 | 0.13948 | 0.69742 | 0.5 | 1
| 188 | 5.233 | 0.0061803 | 0.14833 | 0.74164 | 0.6 | 0
| 190 | 4.7705 | 0.0062006 | 0.14881 | 0.74407 | 0.7 | 0
| 195 | 4.7378 | 0.0061799 | 0.14832 | 0.74159 | 0.8 | 1
| 198 | 5.1499 | 0.00639 | 0.15336 | 0.76681 | 0.9 | 0
| 201 | 4.7395 | 0.0063375 | 0.1521 | 0.76049 | 1 | 0
| 209 | 4.9886 | 0.0063986 | 0.15357 | 0.76784 | More | 0
| 215 | 4.7085 | 0.0062816 | 0.15076 | 0.75379
| 217 | 4.6241 | 0.0061301 | 0.14712 | 0.73562
|  | #N/A | #N/A | #N/A | #N/A
|  | #N/A | #N/A | #N/A | #N/A
|  | #N/A | #N/A | #N/A | #N/A
|  | #N/A | #N/A | #N/A | #N/A
|  | #N/A | #N/A | #N/A | #N/A
|  | #N/A | #N/A | #N/A | #N/A
|  | #N/A | #N/A | #N/A | #N/A
#N/A

## d1
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-07-23 | 8:30 | 12 | 288.3542 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C)
20.9 | 3.7e-07 | 21 | 43.576 | 22.5
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.342 | 0.0883 | 73.9555 | 2.1259e-05 | 2.32e-07
270 | 1.262 | 0.0857 | 73.958
274 | 1.114 | 0.0805 | 73.9633
278 | 0.931 | 0.0716 | 73.9726
282 | 0.765 | 0.0613 | 73.983
286 | 0.588 | 0.0491 | 73.9946
290 | 0.422 | 0.0353 | 74.0081
294 | 0.281 | 0.0198 | 74.021
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.6645 |  | 156.7926

## d2
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-07-24 | 9:30 | 13 | 312.3958 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C)
20.8 | 3.4e-07 | 20.9 | 43.537 | 21.9
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.887 | 0.0883 | 73.9551 | 2.1279e-05 | 2.32e-07
270 | 1.775 | 0.0856 | 73.9578
274 | 1.555 | 0.0799 | 73.9633
278 | 1.309 | 0.072 | 73.9713
282 | 1.062 | 0.0613 | 73.982
286 | 0.791 | 0.0496 | 73.9933
290 | 0.569 | 0.0364 | 74.0058
294 | 0.374 | 0.0205 | 74.0199
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.6447 |  | 156.6483

## d3
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-07-25 | 9:00 | 14 | 336.375 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C)
21 | 3.1e-07 | 21 | 43.537 | 21.9
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.808 | 0.0892 | 73.9555 | 2.1271e-05 | 2.32e-07
270 | 1.636 | 0.0867 | 73.9575
274 | 1.428 | 0.0812 | 73.9629
278 | 1.207 | 0.0724 | 73.9714
282 | 0.945 | 0.0615 | 73.9825
286 | 0.738 | 0.0499 | 73.9938
290 | 0.531 | 0.0372 | 74.0059
294 | 0.341 | 0.0248 | 74.0185
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.713 |  | 156.7062

## d4
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-07-26 | 13:00 | 15 | 360.5417 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C)
21.1 | 2.9e-07 | 21.1 | 43.547 | 21.9
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.677 | 0.0885 | 73.9559 | 2.1266e-05 | 2.33e-07
270 | 1.515 | 0.0859 | 73.9585
274 | 1.306 | 0.08 | 73.9642
278 | 1.104 | 0.0717 | 73.9725
282 | 0.899 | 0.0613 | 73.9829
286 | 0.681 | 0.0488 | 73.9949
290 | 0.497 | 0.0365 | 74.0068
294 | 0.324 | 0.0192 | 74.0209
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.5979 |  | 156.7411

## d5x
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy | Power applied to heater - 16.0Vdc, 237mA, 68.5ohm, 3.74W. Lock lost, cavity realigned.nLock relost
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-07-29 | 13:30 | 18 | 432.5625 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C)
21.4 | 5.6e-07 | 21.4 | 43.562 | 64.8
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.439 | 0.0881 | 73.9582 | 2.1284e-05 | 2.32e-07
270 | 1.436 | 0.0857 | 73.961
274 | 1.349 | 0.0806 | 73.9665
278 | 1.22 | 0.0718 | 73.9752
282 | 1.029 | 0.0612 | 73.9849
286 | 0.883 | 0.0497 | 73.9956
290
294
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | #VALUE! |  | 156.6122

## d5
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy | Lock lost, cavity realigned. Heater voltage raised after testing
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-07-30 | 10:00 | 19 | 456.4167 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C)
21.4 | 3.6e-07 | 21.4 | 43.564 | 67.9
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.611 | 0.0891 | 73.9579 | 2.1403e-05 | 2.36e-07
270 | 1.581 | 0.0866 | 73.9599
274 | 1.503 | 0.0809 | 73.9665
278 | 1.34 | 0.0725 | 73.974
282 | 1.141 | 0.0622 | 73.984
286 | 0.906 | 0.0503 | 73.9958
290 | 0.669 | 0.0376 | 74.0072
294 | 0.447 | 0.0252 | 74.0194
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.6388 |  | 155.7378

## d6
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-08-01 | 17:20 | 21 | 504.7222 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C)
21.2 | 3.2e-07 | 21.2 | 43.564 | 80.4
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.948 | 0.0898 | 73.9545 | 2.1436e-05 | 2.39e-07
270 | 0.901 | 0.0874 | 73.9568
274 | 0.87 | 0.0816 | 73.9628
278 | 0.825 | 0.073 | 73.9713
282 | 0.742 | 0.0624 | 73.9814
286 | 0.629 | 0.0501 | 73.9939
290 | 0.484 | 0.0375 | 74.0054
294 | 0.339 | 0.0255 | 74.0178
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.7108 |  | 155.5043

## d7
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-08-06 | 12:30 | 26 | 624.5208 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C)
21.3 | 2.5e-07 | 21.2 | 43.548 | 86.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.026 | 0.0916 | 73.9534 | 2.1414e-05 | 2.38e-07
270 | 0.962 | 0.0897 | 73.9552
274 | 0.872 | 0.0837 | 73.9613
278 | 0.768 | 0.0751 | 73.9702
282 | 0.65 | 0.0637 | 73.9808
286 | 0.527 | 0.0516 | 73.9927
290 | 0.398 | 0.0386 | 74.0046
294 | 0.286 | 0.0255 | 74.0175
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.6603 |  | 155.6622

## d8
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-08-09 | 9:00 | 29 | 696.375 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C)
21.5 | 2.3e-07 | 21.5 | 43.548 | 86.9
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.495 | 0.092 | 73.9547 | 2.1421e-05 | 2.37e-07
270 | 1.609 | 0.0892 | 73.9576
274 | 1.613 | 0.0829 | 73.9633
278 | 1.493 | 0.074 | 73.9726
282 | 1.301 | 0.0634 | 73.9827
286 | 1.068 | 0.0513 | 73.9942
290 | 0.799 | 0.0381 | 74.0064
294 | 0.544 | 0.0258 | 74.0185
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.638 |  | 155.6097

## d9
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy | Output Power Monitor fixed (batteries installed were leaking)
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-08-15 | 14:30 | 35 | 840.6042 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C)
22.4 | 2e-07 | 22.4 | 43.779 | 86.8
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.075 | 0.0905 | 73.9578 | 2.1418e-05 | 2.37e-07
270 | 1.198 | 0.0883 | 73.9602
274 | 1.166 | 0.0823 | 73.9657
278 | 1.055 | 0.0739 | 73.9744
282 | 0.949 | 0.0637 | 73.9845
286 | 0.786 | 0.0523 | 73.995
290 | 0.578 | 0.0395 | 74.0072
294 | 0.404 | 0.0288 | 74.018
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.6755 |  | 155.6359

## d10
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-08-19 | 10:30 | 39 | 936.4375 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C)
22.1 | 1.87e-07 | 22.2 | 43.779 | 86.8
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.054 | 0.0913 | 73.9589 | 2.1441e-05 | 2.38e-07
270 | 1.099 | 0.0884 | 73.9614
274 | 1.074 | 0.0824 | 73.9669
278 | 1.135 | 0.0737 | 73.9754
282 | 1.003 | 0.062 | 73.9853
286 | 0.843 | 0.0508 | 73.9956
290 | 0.659 | 0.0373 | 74.0078
294 | 0.459 | 0.0248 | 74.0197
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.4456 |  | 155.4644

## d11
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-08-21 | 7:30 | 41 | 984.3125 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C)
22.1 | 1.87e-07 | 22.1 | 43.781 | 86.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.03 | 0.0902 | 73.9592 | 2.1404e-05 | 2.37e-07
270 | 1.059 | 0.0876 | 73.9618
274 | 1.091 | 0.0818 | 73.9675
278 | 1.065 | 0.0726 | 73.9745
282 | 0.934 | 0.0617 | 73.9851
286 | 0.778 | 0.0504 | 73.9956
290 | 0.602 | 0.0372 | 74.0079
294 | 0.419 | 0.0249 | 74.0185
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.4397 |  | 155.7346

## d12
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-08-23 | 10:30 | 43 | 1032.4375 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
21.3 | 1.6e-07 | 21.3 | 43.738 | 86.1 | 0.275 | 20.2
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.094 | 0.0894 | 73.9526 | 2.1374e-05 | 2.36e-07
270 | 1.005 | 0.0872 | 73.9549
274 | 0.895 | 0.0816 | 73.9604
278 | 0.763 | 0.0724 | 73.9689
282 | 0.648 | 0.0619 | 73.9799
286 | 0.519 | 0.0504 | 73.9914
290 | 0.391 | 0.0376 | 74.0043
294 | 0.266 | 0.0252 | 74.0164
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.7704 |  | 155.9537

## d13
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-08-29 | 11:00 | 49 | 1176.4583 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
20.9 | 1.5e-07 | 20.9 | 43.722 | 86 | 0.275 | 20.2
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.889 | 0.085 | 73.9539 | 2.1379e-05 | 2.37e-07
270 | 1.708 | 0.0827 | 73.9567
274 | 1.597 | 0.0769 | 73.9619
278 | 1.366 | 0.069 | 73.9707
282 | 1.13 | 0.0596 | 73.9807
286 | 0.905 | 0.0479 | 73.9925
290 | 0.665 | 0.0355 | 74.005
294 | 0.433 | 0.0241 | 74.0178
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.9428 |  | 155.915

## d14
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-09-03 | 10:30 | 54 | 1296.4375 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
20.8 | 1.3e-07 | 20 | 43.722 | 86.1 | 0.275 | 20.2
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.847 | 0.0845 | 73.9545 | 2.1355e-05 | 2.36e-07
270 | 1.824 | 0.0824 | 73.9565
274 | 1.72 | 0.0773 | 73.9626
278 | 1.527 | 0.0691 | 73.9703
282 | 1.282 | 0.0587 | 73.9812
286 | 0.993 | 0.0475 | 73.9925
290 | 0.718 | 0.0355 | 74.0055
294 | 0.475 | 0.0244 | 74.0162
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.8931 |  | 156.088

## d15
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy | Heater temperature increase to 120C began. Vin to 21.5, I=292mA
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-09-06 | 8:45 | 57 | 1368.3646 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
21 | 1.3e-07 | 21 | 43.722 | 86.4 Start - 89.3 End | 0.292 | 21.5
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.889 | 0.0846 | 73.9548 | 2.135e-05 | 2.36e-07
270 | 1.816 | 0.0822 | 73.9573
274 | 1.645 | 0.0772 | 73.9626
278 | 1.383 | 0.0691 | 73.971
282 | 1.128 | 0.0586 | 73.9813
286 | 0.894 | 0.0479 | 73.9926
290 | 0.645 | 0.0353 | 74.0058
294 | 0.433 | 0.0245 | 74.0165
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.892 |  | 156.1244

## d16
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy | Heater at 120C
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-09-10 | 9:20 | 61 | 1464.3889 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
20.9 | 1.7e-07 | 20.9 | 43.722 | 120.9 | 0.35 | 27
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.483 | 0.0841 | 73.9536 | 2.1345e-05 | 2.35e-07
270 | 1.592 | 0.0818 | 73.9565
274 | 1.559 | 0.0765 | 73.962
278 | 1.459 | 0.0683 | 73.9706
282 | 1.289 | 0.0589 | 73.9806
286 | 1.048 | 0.0474 | 73.9924
290 | 0.793 | 0.0358 | 74.0044
294 | 0.529 | 0.0242 | 74.0162
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.9438 |  | 156.168

## d17
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-09-12 | 7:15 | 63 | 1512.3021 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
23.3 | 1.7e-07 | 23.3 | 43.712 | 120.1 | 0.36 | 27.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.879 | 0.0794 | 73.9687 | 2.1367e-05 | 2.37e-07
270 | 1.583 | 0.0785 | 73.9703
274 | 1.156 | 0.0733 | 73.9751
278 | 0.808 | 0.066 | 73.9818
282 | 0.687 | 0.0565 | 73.9906
286 | 0.731 | 0.0451 | 74.0006
290 | 0.668 | 0.0332 | 74.0112
294 | 0.516 | 0.0228 | 74.0207
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 3.4397 |  | 156.0047

## d18
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-09-17 | 10:15 | 68 | 1632.4271 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
23.3 | 1.2e-07 | 19.5 | 43.667 | 119.3 | 0.363 | 27.9
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.515 | 0.0745 | 73.9524 | 2.139e-05 | 2.37e-07
270 | 1.291 | 0.0739 | 73.9529
274 | 1.116 | 0.0694 | 73.958
278 | 0.94 | 0.0624 | 73.9667
282 | 0.749 | 0.0531 | 73.9775
286 | 0.576 | 0.0433 | 73.9898
290 | 0.421 | 0.0325 | 74.0025
294 | 0.28 | 0.0224 | 74.0154
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.5693 |  | 155.8394

## d19
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-09-23 | 15:00 | 74 | 1776.625 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
19.4 | 1.1e-07 | 19.4 | 43.667 | 119.3 | 0.363 | 27.9
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.925 | 0.0736 | 74.0145 | 2.1392e-05 | 2.36e-07
270 | 1.884 | 0.0728 | 74.0151
274 | 1.732 | 0.0686 | 74.0208
278 | 1.509 | 0.0615 | 74.0297
282 | 1.256 | 0.0527 | 74.0403
286 | 0.983 | 0.043 | 74.0524
290 | 0.72 | 0.0325 | 74.0652
294 | 0.479 | 0.0221 | 74.0782
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.6826 |  | 155.8203

## d20
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-10-01 | 8:30 | 82 | 1968.3542 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
19.3 | 9.1e-08 | 19.3 | 43.644 | 119.2 | 0.364 | 27.9
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.57 | 0.0741 | 73.952 | 2.1416e-05 | 2.37e-07
270 | 1.679 | 0.0729 | 73.9539
274 | 1.655 | 0.0679 | 73.9594
278 | 1.517 | 0.0607 | 73.9682
282 | 1.312 | 0.0518 | 73.9787
286 | 1.078 | 0.0419 | 73.991
290 | 0.812 | 0.0315 | 74.0029
294 | 0.543 | 0.0207 | 74.0166
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.5502 |  | 155.6494

## d21
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy | lost lost on 10/2. reestablished 10/3
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-10-04 | 9:15 | 85 | 2040.3854 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
19.1 | 8.6e-08 | 19.3 | 44.656 | 119.5 | 0.363 | 27.9
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.25 | 0.0763 | 73.9557 | 2.143e-05 | 2.36e-07
270 | 1.3236 | 0.0744 | 73.9577
274 | 1.339 | 0.0694 | 73.9635
278 | 1.271 | 0.0618 | 73.9721
282 | 1.14 | 0.0529 | 73.9821
286 | 0.951 | 0.0427 | 73.9937
290 | 0.728 | 0.032 | 74.0061
294 | 0.508 | 0.0218 | 74.0175
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.3003 |  | 155.5467

## d22
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy | lock lost 10/14. Unable to maintaine lock until the cavity was realigned.
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-10-15 | 9:30 | 96 | 2304.3958 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
19.2 | 6.9e-08 | 19.2 | 44.463 | 119.4 | 0.364 | 27.9
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.058 | 0.0701 | 73.9627 | 2.1383e-05 | 2.34e-07
270 | 1.175 | 0.0681 | 73.9642
274 | 1.202 | 0.0637 | 73.9698
278 | 1.182 | 0.0569 | 73.978
282 | 1.078 | 0.0484 | 73.9875
286 | 0.924 | 0.0392 | 73.998
290 | 0.719 | 0.0291 | 74.0098
294 | 0.501 | 0.0196 | 74.0205
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.3629 |  | 155.8891

## d23
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy | lock lost on the 294 measurment. Lock re-established at 44.460
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-10-16 | 10:00 | 97 | 2328.4167 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
19.9 | 6.8e-08 | 19.9 | 44.463 | 119.4 | 0.364 | 27.9
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.023 | 0.0705 | 73.9625 | 2.1392e-05 | 2.34e-07
270 | 1.101 | 0.0685 | 73.9645
274 | 1.128 | 0.0636 | 73.9697
278 | 1.059 | 0.0572 | 73.9785
282 | 0.947 | 0.0488 | 73.9864
286 | 0.782 | 0.0393 | 73.9985
290 | 0.606 | 0.0295 | 74.0097
294 | 0.429 | 0.0194 | 74.0203
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.3224 |  | 155.8246

## d24
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy | Heater temp set to 87 10/21/19 12:00
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-10-23 | 13:00 | 104 | 2496.5417 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
20.1 | 4.9e-08 | 20.1 | 44.442 | 85.4 | 0.278 | 20.5
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.735 | 0.0714 | 73.9604 | 2.1373e-05 | 2.33e-07
270 | 1.794 | 0.0699 | 73.9627
274 | 1.727 | 0.0653 | 73.9677
278 | 1.569 | 0.0581 | 73.9756
282 | 1.332 | 0.0493 | 73.9855
286 | 1.082 | 0.0398 | 73.9968
290 | 0.819 | 0.0302 | 74.0081
294 | 0.559 | 0.0204 | 74.0194
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.354 |  | 155.9613

## d25
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy | Heater temp set to 87 10/21/19 12:00
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-10-25 | 7:30 | 106 | 2544.3125 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
20.1 | 4.8e-08 | 20 | 44.446 | 86.2 | 0.278 | 20.5
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.332 | 0.0717 | 73.9605 | 2.1373e-05 | 2.32e-07
270 | 1.422 | 0.0701 | 73.9616
274 | 1.456 | 0.0653 | 73.9674
278 | 1.342 | 0.0582 | 73.9754
282 | 1.17 | 0.0498 | 73.9855
286 | 0.977 | 0.0401 | 73.9964
290 | 0.753 | 0.0302 | 74.0083
294 | 0.519 | 0.0203 | 74.0195
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.3804 |  | 155.9584

## d26
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-10-30 | 8:00 | 111 | 2664.3333 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
4.6e-08 | 2.8e-07 | 20 | 44.441 | 86.3 | 0.278 | 20.6
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.172 | 0.0714 | 73.9603 | 2.1359e-05 | 2.32e-07
270 | 1.301 | 0.0692 | 73.9635
274 | 1.299 | 0.0648 | 73.968
278 | 1.222 | 0.0578 | 73.9758
282 | 1.061 | 0.0492 | 73.9859
286 | 0.879 | 0.04 | 73.9961
290 | 0.679 | 0.0301 | 74.0078
294 | 0.462 | 0.0203 | 74.019
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.3209 |  | 156.0643

## d27
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-11-04 | 12:30 | 116 | 2784.5208 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
4.6e-08 | 4.1e-08 | 19.9 | 44.441 | 86.5 | 0.279 | 20.6
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.02 | 0.0659 | 73.9688 | 2.1385e-05 | 2.33e-07
270 | 0.894 | 0.0635 | 73.9696
274 | 0.802 | 0.0601 | 73.9742
278 | 0.716 | 0.0536 | 73.9808
282 | 0.655 | 0.0465 | 73.9895
286 | 0.571 | 0.0371 | 73.9996
290 | 0.443 | 0.0277 | 74.0105
294 | 0.306 | 0.0184 | 74.0213
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.2455 |  | 155.871

## d28
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-11-07 | 13:15 | 119 | 2856.5521 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
20.1 | 4.1e-08 | 20.1 | 44.434 | 87.1 | 0.281 | 20.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.575 | 0.069 | 73.9641 | 2.1384e-05 | 2.33e-07
270 | 1.703 | 0.0662 | 73.9662
274 | 1.645 | 0.0621 | 73.9714
278 | 1.493 | 0.0557 | 73.9785
282 | 1.28 | 0.0477 | 73.9879
286 | 1.036 | 0.0382 | 73.9984
290 | 0.783 | 0.029 | 74.0091
294 | 0.536 | 0.0201 | 74.0193
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.3049 |  | 155.8797

## d29
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-11-12 | 9:30 | 124 | 2976.3958 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
20 | 4e-08 | 20 | 44.248 | 87.2 | 0.281 | 20.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.285 | 0.0683 | 73.9628 | 2.1367e-05 | 2.33e-07
270 | 1.38 | 0.067 | 73.9646
274 | 1.38 | 0.062 | 73.9696
278 | 1.284 | 0.0557 | 73.9781
282 | 1.135 | 0.0474 | 73.9875
286 | 0.954 | 0.0389 | 73.9979
290 | 0.746 | 0.0294 | 74.0084
294 | 0.509 | 0.0197 | 74.0193
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.4233 |  | 156.0015

## d30
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-11-14 | 14:00 | 126 | 3024.5833 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
19.9 | 4e-08 | 19.9 | 44.225 | 87.2 | 0.281 | 20.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.928 | 0.0695 | 73.9628 | 2.1371e-05 | 2.33e-07
270 | 0.965 | 0.0679 | 73.964
274 | 0.962 | 0.0627 | 73.9695
278 | 0.911 | 0.0563 | 73.9772
282 | 0.825 | 0.0486 | 73.9862
286 | 0.687 | 0.0391 | 73.9972
290 | 0.519 | 0.029 | 74.0096
294 | 0.359 | 0.019 | 74.0209
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.4017 |  | 155.9772

## d31
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-11-18 | 10:30 | 130 | 3120.4375 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
19.8 | 3.6e-08 | 19.8 | 44.225 | 87.3 | 0.281 | 20.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.65 | 0.0679 | 73.9638 | 2.2548e-05 | 3.21e-07
270 | 1.677 | 0.0662 | 73.9654
274 | 1.589 | 0.0618 | 73.9703
278 | 1.416 | 0.0555 | 73.9776
282 | 1.211 | 0.0477 | 73.9869
286 | 0.97 | 0.0385 | 73.9977
290 | 0.722 | 0.0292 | 74.0086
294 | 0.475 | 0.0194 | 74.0195
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.3852 |  | 147.8332

## d32
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-11-20 | 14:30 | 132 | 3168.6042 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
20 | 3.7e-08 | 20 | 44.225 | 87.4 | 0.281 | 20.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.233 | 0.0683 | 73.9635 | 2.3412e-05 | 3.19e-07
270 | 1.289 | 0.0661 | 73.9653
274 | 1.293 | 0.0615 | 73.9715
278 | 1.206 | 0.0549 | 73.978
282 | 1.084 | 0.0474 | 73.9874
286 | 0.931 | 0.0392 | 73.9973
290 | 0.704 | 0.0287 | 74.0095
294 | 0.494 | 0.02 | 74.0196
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.4269 |  | 142.3753

## d33
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-11-22 | 13:00 | 134 | 3216.5417 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
19.9 | 3.6e-08 | 19.9 | 44.225 | 87.2 | 0.281 | 20.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.885 | 0.0682 | 73.9635 | 2.3612e-05 | 3.29e-07
270 | 0.849 | 0.0669 | 73.9653
274 | 0.794 | 0.0621 | 73.9706
278 | 0.715 | 0.0554 | 73.978
282 | 0.62 | 0.047 | 73.9877
286 | 0.502 | 0.0378 | 73.9988
290 | 0.386 | 0.0285 | 74.0095
294 | 0.266 | 0.019 | 74.0208
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.3916 |  | 141.1689

## d34
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-11-26 | 13:00 | 138 | 3312.5417 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
19.7 | 3.3e-08 | 19.7 | 44.225 | 87.2 | 0.281 | 20.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.093 | 0.067 | 73.9654 | 2.368e-05 | 3.31e-07
270 | 1.002 | 0.0659 | 73.9672
274 | 0.874 | 0.0609 | 73.9721
278 | 0.76 | 0.0555 | 73.9783
282 | 0.634 | 0.0473 | 73.9874
286 | 0.505 | 0.0388 | 73.9975
290 | 0.371 | 0.0288 | 74.0094
294 | 0.254 | 0.02 | 74.0194
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.3363 |  | 140.7652

## d35
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy | Lost lock when pulse was applied, had to walk in mirrors to increase power and maintain lock
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-12-02 | 14:00 | 144 | 3456.5833 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
19.7 | 3.1e-08 | 19.7 | 44.207 | 87.2 | 0.281 | 20.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.967 | 0.069 | 73.9632 | 2.4215e-05 | 3.59e-07
270 | 0.945 | 0.0686 | 73.9641
274 | 0.897 | 0.0681 | 73.965
278 | 0.815 | 0.0631 | 73.9702
282 | 0.714 | 0.055 | 73.9796
286 | 0.589 | 0.0447 | 73.9909
290 | 0.445 | 0.0332 | 74.004
294 | 0.311 | 0.0205 | 74.0166
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.1822 |  | 137.6531

## xxd36xx
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy | Lock lost at ringdown. Measurement abandoned due to Power Out Monitor measurement. Output meter readjusted, BFPD1 adjusted, and cavity input mirrors re-optimized
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-12-04 | 7:00 | 146 | 3504.2917 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
19.6 | 3.5e-08 | 19.6 | 44.02 | 87.2 | 0.281 | 20.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.451 | 0.067 | 73.9647 | 2.4202e-05 | 3.58e-07
270 | 1.301 | 0.066 | 73.9654 | 2.4113e-05 | 3.53e-07
274 | 1.099 | 0.0665 | 73.964
278
282
286
290
294
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | #VALUE! |  | 137.7316

## d36
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-12-04 | 7:40 | 146 | 3504.3194 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
19.9 | 3.5e-08 | 20 | 44.02 | 87.2 | 0.281 | 20.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.901 | 0.0727 | 73.9598 | 2.4083e-05 | 3.53e-07
270 | 0.877 | 0.0717 | 73.9604
274 | 0.826 | 0.068 | 73.9642
278 | 0.741 | 0.0606 | 73.9733
282 | 0.638 | 0.0518 | 73.9829
286 | 0.524 | 0.0422 | 73.9994
290 | 0.389 | 0.031 | 74.0073
294 | 0.266 | 0.0209 | 74.0188
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.421 |  | 138.4118

## d37
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy | Lost lock on the 94 measurement, regained lock at 44.007
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-12-06 | 13:00 | 148 | 3552.5417 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
19.5 | 3e-08 | 19.5 | 44.02 | 87.2 | 0.281 | 20.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.702 | 0.0721 | 73.9617 | 2.4801e-05 | 4.93e-07
270 | 1.503 | 0.0705 | 73.9634
274 | 1.317 | 0.0663 | 73.9682
278 | 1.1 | 0.0586 | 73.976
282 | 0.898 | 0.0504 | 73.986
286 | 0.695 | 0.0404 | 73.9964
290 | 0.498 | 0.0295 | 74.0097
294 | 0.318 | 0.0151 | 74.0208
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.0331 |  | 134.4035

## d38
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-12-10 | 13:00 | 152 | 3648.5417 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
19.6 | 2.9e-08 | 19.6 | 44.007 | 87.1 | 0.281 | 20.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.783 | 0.0708 | 73.9625 | 2.4828e-05 | 4.95e-07
270 | 1.669 | 0.0688 | 73.9644
274 | 1.502 | 0.0648 | 73.9694
278 | 1.295 | 0.0583 | 73.9768
282 | 1.078 | 0.0498 | 73.9862
286 | 0.849 | 0.0398 | 73.9979
290 | 0.635 | 0.0304 | 74.0081
294 | 0.414 | 0.0202 | 74.0194
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.2715 |  | 134.2589

## d39
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-12-12 | 14:00 | 154 | 3696.5833 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
19.6 | 2.8e-08 | 19.6 | 44.02 | 87.1 | 0.281 | 20.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.822 | 0.0716 | 73.9619 | 2.4794e-05 | 4.91e-07
270 | 1.688 | 0.0699 | 73.9634
274 | 1.514 | 0.0658 | 73.9682
278 | 1.293 | 0.0591 | 73.9763
282 | 1.083 | 0.051 | 73.9848
286 | 0.839 | 0.0406 | 73.9969
290 | 0.611 | 0.0303 | 74.0087
294 | 0.409 | 0.0207 | 74.0193
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.287 |  | 134.4415

## d40
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-12-16 | 14:00 | 158 | 3792.5833 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
19.5 | 2.6e-08 | 19.5 | 44.02 | 87.1 | 0.281 | 20.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.894 | 0.072 | 73.9612 | 2.4803e-05 | 4.92e-07
270 | 1.765 | 0.0703 | 73.9691
274 | 1.605 | 0.0659 | 73.968
278 | 1.351 | 0.059 | 73.9756
282 | 1.108 | 0.0504 | 73.9858
286 | 0.876 | 0.0414 | 73.9954
290 | 0.632 | 0.0301 | 74.0084
294 | 0.426 | 0.0203 | 74.0193
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.0861 |  | 134.3941

## d41
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2019-12-18 | 14:00 | 160 | 3840.5833 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
19.6 | 2.7e-08 | 19.6 | 44.02 | 87.3 | 0.281 | 20.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.908 | 0.0713 | 73.962 | 2.49e-05 | 5e-07
270 | 1.815 | 0.0707 | 73.962
274 | 1.651 | 0.0696 | 73.9636
278 | 1.452 | 0.0629 | 73.9706
282 | 1.219 | 0.0534 | 73.9813
286 | 0.98 | 0.0436 | 73.9927
290 | 0.73 | 0.0326 | 74.0051
294 | 0.508 | 0.0227 | 74.016
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.2467 |  | 133.8696

## d42
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2020-01-03 | 14:00 | 176 | 4224.5833 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
21.1 | 2.1e-08 | 21.1 | 44.096 | 88.1 | 0.281 | 20.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.981 | 0.0759 | 73.9544 | 2.4128e-05 | 4.47e-07
270 | 1.932 | 0.0749 | 73.9553
274 | 1.799 | 0.0709 | 73.96
278 | 1.597 | 0.0632 | 73.9689
282 | 1.345 | 0.0542 | 73.9792
286 | 1.089 | 0.0444 | 73.9912
290 | 0.805 | 0.0328 | 74.0049
294 | 0.549 | 0.0224 | 74.0165
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.4253 |  | 138.1548

## d43
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2020-01-07 | 15:00 | 180 | 4320.625 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
21 | 2.4e-08 | 21 | 44.096 | 88 | 0.281 | 20.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.992 | 0.0681 | 73.9639 | 2.4721e-05 | 4.96e-07
270 | 1.936 | 0.0675 | 73.9642
274 | 1.783 | 0.066 | 73.9658
278 | 1.571 | 0.0639 | 73.9684
282 | 1.328 | 0.0558 | 73.978
286 | 1.048 | 0.0446 | 73.9907
290 | 0.756 | 0.0332 | 74.004
294 | 0.503 | 0.0223 | 74.0168
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.3905 |  | 134.8407

## d44
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2020-01-09 | 15:00 | 182 | 4368.625 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
21.3 | 2.5e-08 | 21.3 | 44.096 | 88.2 | 0.281 | 20.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.505 | 0.078 | 73.9432 | 2.4549e-05 | 3.11e-07
270 | 1.378 | 0.0765 | 73.9437
274 | 1.202 | 0.0746 | 73.9464
278 | 1.039 | 0.0727 | 73.9491
282 | 0.879 | 0.0679 | 73.9552
286 | 0.694 | 0.0544 | 73.9724
290 | 0.509 | 0.0411 | 73.9901
294 | 0.347 | 0.0278 | 74.0068
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.8644 |  | 135.7846

## d45
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2020-01-13 | 13:00 | 186 | 4464.5417 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
21.2 | 2.6e-08 | 21.2 | 44.096 | 88.3 | 0.281 | 20.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.655 | 0.0767 | 74.0015 | 2.4687e-05 | 3.22e-07
270 | 1.567 | 0.0759 | 74.003
274 | 1.43 | 0.0738 | 74.0058
278 | 1.246 | 0.0723 | 74.0082
282 | 0.991 | 0.0669 | 74.0151
286 | 0.789 | 0.0547 | 74.033
290 | 0.587 | 0.041 | 74.0512
294 | 0.386 | 0.0282 | 74.0693
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 5.2686 |  | 135.0234

## d46
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2020-01-15 | 13:00 | 188 | 4512.5417 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
21.3 | 2.5e-08 | 21.3 | 44.096 | 88.2 | 0.281 | 20.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.466 | 0.0772 | 74.0015 | 2.4655e-05 | 3.2e-07
270 | 1.367 | 0.0765 | 74.003
274 | 1.241 | 0.0742 | 74.0062
278 | 1.063 | 0.0728 | 74.0081
282 | 0.865 | 0.0673 | 74.0156
286 | 0.673 | 0.0541 | 74.0341
290 | 0.51 | 0.0409 | 74.0521
294 | 0.341 | 0.0274 | 74.0707
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 5.233 |  | 135.2009

## d47
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2020-01-17 | 13:00 | 190 | 4560.5417 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
21.3 | 2.3e-08 | 21.3 | 44.096 | 88.3 | 0.281 | 20.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.542 | 0.0853 | 73.9343 | 2.4356e-05 | 2.99e-07
270 | 1.422 | 0.0856 | 73.9343
274 | 1.273 | 0.084 | 73.9357
278 | 1.078 | 0.0784 | 73.9435
282 | 0.867 | 0.0672 | 73.9575
286 | 0.675 | 0.0546 | 73.9731
290 | 0.504 | 0.0403 | 73.9909
294 | 0.354 | 0.0277 | 74.0072
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.7705 |  | 136.8602

## d48
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2020-01-22 | 13:00 | 195 | 4680.5417 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
21.4 | 2.3e-08 | 21.4 | 44.096 | 88.2 | 0.281 | 20.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.193 | 0.0785 | 73.944 | 2.4591e-05 | 3.13e-07
270 | 1.113 | 0.0777 | 73.9456
274 | 0.999 | 0.0754 | 73.9476
278 | 0.868 | 0.074 | 73.9497
282 | 0.719 | 0.0672 | 73.9584
286 | 0.578 | 0.055 | 73.9733
290 | 0.432 | 0.0412 | 73.9904
294 | 0.285 | 0.0277 | 74.008
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.7378 |  | 135.5494

## d49
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2020-01-25 | 13:20 | 198 | 4752.5556 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
21.2 | 2.5e-08 | 21.2 | 44.096 | 88.2 | 0.281 | 20.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.092 | 0.0795 | 74.0016 | 2.4359e-05 | 3e-07
270 | 1.006 | 0.0777 | 74.0025
274 | 0.908 | 0.0753 | 74.0056
278 | 0.794 | 0.0737 | 74.0088
282 | 0.667 | 0.0686 | 74.0159
286 | 0.538 | 0.0556 | 74.0331
290 | 0.402 | 0.0415 | 74.0524
294 | 0.274 | 0.028 | 74.0708
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 5.1499 |  | 136.8432

## d50
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2020-01-28 | 11:30 | 201 | 4824.4792 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
21.2 | 2.4e-08 | 21.2 | 44.096 | 88.2 | 0.281 | 20.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.293 | 0.0805 | 73.9437 | 2.4312e-05 | 2.96e-07
270 | 1.107 | 0.0797 | 73.944
274 | 0.964 | 0.0778 | 73.9462
278 | 0.81 | 0.0765 | 73.9483
282 | 0.668 | 0.0695 | 73.9569
286 | 0.539 | 0.0569 | 73.9721
290 | 0.403 | 0.0429 | 73.9904
294 | 0.276 | 0.029 | 74.0078
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.7395 |  | 137.1082

## d51
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2020-02-05 | 14:00 | 209 | 5016.5833 | 2.49
7 | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
21.7 | 1.8e-08 | 21.7 | 43.922 | 88.3 | 0.281 | 20.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.965 | 0.0797 | 74.0043 | 2.4676e-05 | 3.2e-07
270 | 0.969 | 0.0794 | 74.0054
274 | 0.961 | 0.0763 | 74.0085
278 | 0.93 | 0.0738 | 74.0123
282 | 0.828 | 0.0678 | 74.0201
286 | 0.708 | 0.0551 | 74.037
290 | 0.55 | 0.0415 | 74.0547
294 | 0.373 | 0.0283 | 74.0723
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.9886 |  | 135.0851

## d52
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2020-02-11 | 14:00 | 215 | 5160.5833 | 2.49
7 | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
21.6 | 2.1e-08 | 21.6 | 43.908 | 88.2 | 0.281 | 20.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.26 | 0.0781 | 73.9449 | 2.463e-05 | 3.15e-07
270 | 1.175 | 0.078 | 73.9453
274 | 1.022 | 0.0751 | 73.9483
278 | 0.878 | 0.0726 | 73.9513
282 | 0.712 | 0.0678 | 73.9576
286 | 0.587 | 0.0566 | 73.9714
290 | 0.431 | 0.0418 | 73.9897
294 | 0.288 | 0.0283 | 74.007
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.7085 |  | 135.3369

## d53
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 | 2020-02-13 | 14:00 | 217 | 5208.5833 | 2.49
7 | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C) | Drive Current (A) | Drive Voltage (V)
21.6 | 2.1e-08 | 21.6 | 43.908 | 88.2 | 0.281 | 20.7
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.113 | 0.0791 | 73.9451 | 2.4535e-05 | 2.09e-07
270 | 1.048 | 0.0781 | 73.946
274 | 0.926 | 0.0758 | 73.9483
278 | 0.808 | 0.0735 | 73.9515
282 | 0.668 | 0.0675 | 73.9583
286 | 0.534 | 0.0549 | 73.9741
290 | 0.399 | 0.041 | 73.9912
294 | 0.276 | 0.0285 | 74.0066
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 4.6241 |  | 135.8601

## d0
Cavity No | Sample | Comment
1 | SR3 ROC Actuator, Ceramic Heater Assy
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2019-07-11 |  |  | -43657 | -1047768 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C) | Heater Temp (°C)
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264
270
274
278
282
286
290
294
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | #VALUE! |  | #DIV/0!