## Slide 1
CDS Hardware Update
High Reliability Ultra-Fast Shutter (HRUF)
D2200426-v1
Update 11/22/23
Dean Schaetzl for the CDS group
LIGO-G23xxxxx-v1

## Slide 2: Fast Shutter Upgrade High Reliability Ultra Fast (HRUF)
Responsible: D. Schaetzl
Updated: September 25, 2023
ECR: E1900177 DCC: E1900180
Why do we have a Fast Shutter?
  Optical energy stored in the LIGO arms during normal interferometer (IFO) operation is dumped into the antisymmetric (AS) port upon loss of lock. The stored energy can be of order 50J, which appears at the AS port as an optical transient lasting a few tens of milliseconds. In order to protect sensitive photodetectors in the AS port beampath from damage, a shutter is required to quickly block the laser beam upon a lock-loss event
Current Fast Shutter design (D1003318)
  Wires are attached to the moving parts
    Work Hardening
  Arresting of the shutter is done mechanically with hard-stops
    Stress
    Debris
    Physical Shock (Stopping)
New electro-magnetic ultra-fast shutter design of a moving
payload consisting of magnets and a mirror with no wires attached
    A non-moving coil requiring electrical attachment.
    Self Arresting

## Slide 3: Fast Shutter Upgrade High Reliability Ultra Fast (HRUF)
How do we make it better?
  We first want to build a mathematical model of the
system to understand the parameters we can tune
to optimize the behavior. (modeling by Serkay Ölmez)
Real-Time Fast Shutter Simulator

## Slide 4: Fast Shutter Upgrade High Reliability Ultra Fast (HRUF)
A prototype was constructed and studied to create a moving shutter without the need for wires attached to the payload
High-Reliability Ultra-Fast Mechanical Shutter (HRUF) Assy D2200426

## Slide 5: Fast Shutter Upgrade High Reliability Ultra Fast (HRUF)
Mirror
Shoulder
Magnet
Coil
How Does it Work?
  Magnet and flag rest on the flag shoulder just inside the coil
  When activated the magnet pulls up into the coil raising the flag
  The magnetic field interacts with the copper plate and arrests the flag
  A holding current keeps the flag raised
  Falls back under gravity
Flag

## Slide 6: Fast Shutter Upgrade High Reliability Ultra Fast (HRUF)
What have the issues been?
  The throw to the beam-dump is relatively long so any rotation of the FS mirror on deployment could swing and “paint” the chamber
  Tolerances must be kept very tight while still allowing free movement of the flag
Beam Dump
Shutter

## Slide 7: Fast Shutter Upgrade High Reliability Ultra Fast (HRUF)
What have the issues been?
  How do you assemble it?
  What order?
    Magnet-to-Flag
    Mirror-to-Flag
  The magnet doesn’t fit through the copper plate (square peg in a round hole)
  The mirror cannot be attached when the flag is in the copper

## Slide 8: Fast Shutter Upgrade High Reliability Ultra Fast (HRUF)
What have the issues been?
  The mirror cannot be attached when the flag is in the copper
Mirror-to-Flag Fixture D2300217
  Holds the flags still and positions the mirrors correctly for glue-up

## Slide 9: Fast Shutter Upgrade High Reliability Ultra Fast (HRUF)
First Attempt
  We first attempted to glue a magnet-to-flag
    The EP-30 Epoxy used was past expiration and did not properly cure
    Precision pins were used to hold the flag relative to the copper and fixture but were still too magnetic and twisted during assembly

## Slide 10: Fast Shutter Upgrade High Reliability Ultra Fast (HRUF)
Magnet-to-Flag Fixture D2300218
  The “Pac-Man” was developed to hold the flag and copper relative to the fixture
  Must be short enough to be removed when the magnet is glued, but long enough to bridge the transition of the pieces
Pac-Man

## Slide 11: Fast Shutter Upgrade High Reliability Ultra Fast (HRUF)
Gluing the Mirror-to-Flag
  Gluing successful
Mirror Edges not beveled
  Tolerance revenge!
  Manual shaping of test optics
The depth of the mirror on the flag with the epoxy is
tight and does not always fall smoothly
v2 = 0.005” taken out the flag thickness at the mirror

## Slide 12: Fast Shutter Upgrade High Reliability Ultra Fast (HRUF)
Initial test firing
  Does not fall back down
  Does not seem to break
Slow-Mo(ish)
Real-Time

## Slide 13: Fast Shutter Upgrade High Reliability Ultra Fast (HRUF)
Timeline
  All of the physical parts are in-hand or in production as of 9/13
  Magnets and mirrors need to be epoxied together
  Testing once the assembly is complete
  Expenditures for 4 complete assemblies
    Main Bodies S615433 $4809.11
    Flag-to-Mirror Jig S612944 $1132.27
    Magnet-to-Flag Jig S607572 $523.69
    Copper Plates S587290 $1821.33
    Bobbins S586150 $1296.54
    Flags (Ti & PEEK) S603440 $2443.14
    Terminal Cover S616093 $1885.25
    Ti Flags, Pac-Man S621897 $3569.90
10/25/23
  First epoxy test failed. Expired epoxy?
  Parts sent to C&B for a proper Class B clean
  10/23 Parts back from C&B
  10/24 Test optic glued to flag
    One PEEK and one Titanium
  10/27 Glue successful
    Flag thickness possible issue
      Flag redesigned
  11/7 New flags, M-to-F jigs arrived
    Class B clean
  11/20 Glue –up failed
  11/22 Magnet-to-Flag glue-up curing