## Slide 1
LIGO DAC Design Results and Prospects
Daniel Sigg and Marc Pirello
LIGO Hanford Observatory
CDS Meeting October 16, 2024

## Slide 2: Reminder Goals
Technology:
  New ADC/DAC chips typically support faster rates not necessarily more bits
    iLIGO: ~20kHz, aLIGO: ~200kHz, now: ~1MHz.
  Lower noise achieved by faster acquisition
  Avoid higher data rates for IO chassis by implementing FPGA decimation filters
Clocking Improvements:
  No rogues clocks: synchronized switchers, FPGA clock is 226Hz
  Use timing distribution signal as input clock: Converter knows GPS time
  No special startup synchronization procedure required, always synchronized
  Allows for time stamped data transfers
    Watchdog: Outputs can be stopped when no valid data is received
Others:
  Replace failing 18-bit DAC (GS 20-bit DAC very expensive, i.e., ~$7500 for 8 channels)
  Running out of space in some IO chassis

## Slide 3: Status
DAC Prototype done
  Uses SD-Technology, on chip FIR filters replaced by FPGA IIR filters
  Runs at 1MHz data rate, but max useful rate is 128kHz
    Runs at 64kHz in RTCDS: Up-sampling to 1MHz and 30KHz low pass filters in FPGA
  FPGA firmware debugged, DMA and filters working
  RTCDS drivers fully functional
In-Situ DAC Testing
  1 DAC board installed at EX driving ETMX ESD, L1, L2
  Reliability: Installed on 7/25/24, running with connected outputs since 9/17/24
  No visible effect on sensitivity or operations
Test stand results
  Superior to 20bit DAC in all tests performed
  See next pages for details
Report T2400066, compare to G1500761 (20-bit) and G1401399 (18-bit)

## Slide 4: Proposal
Build 20 new 32-chn DACs to replace old 8-chn 18-bit DACs at LHO
  Cost: ~$2500/DAC when ordering 20 or more (includes new adapter/interface boards for AI)
Still use 33 18-bit DACs in H1
  Don’t have enough 20-bit DACs, would have to buy ~20 more to complete the transition
  Would cost ~$150k
Will have extra channels to cover any A+ short fall
  A+ is using 16-bit DACs in some places to save costs, wouldn’t be needed
Will have extra channels to cover JAC/POPX upgrades
  sus2a and sus2b IO chassis can be combined into sus12

## Slide 5: Results from Test Stand
Top: New DAC in empty IO chassis
Bottom: Multiple old 16-bit ADC cards added
Beat notes of oscillators on ADC boards

## Slide 6: 20-bit DAC
Fair amount of excess noise with high bias voltage
Channel 6 probably broken

## Slide 7: New DAC
Chart: Frequency (Hz)

## Slide 8: Four Simple Poles/Zeroes at 1 Hz/10 Hz
Significantly better than 20-bit DAC G1500761

## Slide 9: Four Simple Poles/Zeroes at 5 Hz/50 Hz
Low frequency similar to 20-bit DAC
Not totally clear if we are seeing ADC noise

## Slide 10: cheby1(“LowPass”,8,1,3)
Mostly similar to 20-bit DAC
10-7

## Slide 11: Summary
New DAC has superior performance in most tests
No rogue oscillators
Costs 15x less per channel than 20-bit DAC
Has performed well in ETMX test
Recommend going forward with 1st purchase of 20 boards for H1

## Slide 12: Front-End Chassis
Timing Board
Backplane
Adapter Board
Converter
AA/AI
Data
Sampling Clock
Watchdog
PCIe