LIGO Laboratory / LIGO Scientific Collaboration

LIGO-T2500372-v1	Advanced LIGO	4 November 2025

Fast Digital Servo: Timing Diagrams

Daniel Sigg

Distribution of this document:

LIGO Scientific Collaboration

This is an internal working note

of the LIGO Laboratory.

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 1970 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

http://www.ligo.caltech.edu/

# Overview

This document shows proposed timing diagrams to communicate with the converters in the fast digital servo board. The main clock is assumed to be 228 Hz (or roughly 268 MHz), which is derived from the 226 Hz timing clock through the help of a PLL.

# ADAQ23876 and ADAQ23878 ADCs

These ADCs have an LVDS interface that can be clocked at rates up to 400 MHz. Since we are using two lanes to transfer the data a 134MHz clock frequency is sufficient.

See Figure 1 for the ADAQ23876, and figure 2 for the ADAQ23878.

A full conversion plus readout takes 32 cycles and supports a 223 Hz (8.4 MHz) conversion clock.

# AD3551R and AD3552R

These DACs have quad SPI interface that can be clocked at rates up to 66 MHz. This is slightly lower than a 226 Hz (67.1 MHz) clock, which means that we cannot easily support the AD3552R in a non-streaming mode unless we run the interface 1.6% faster than recommended. For the AD3551R we can use 6 clock cycles at 268 MHz for one SPI clock cycle which results in a 44.7 MHz SPI clock.

See Figure 3 for the AD3551R, and figure 4 for the AD3552R.

A full conversion plus readout takes 32 cycles and supports a 223 Hz (8.4 MHz) conversion clock when the AD3552R is operated at 67.1 MHz. In principle, it is possible to run these devices in streaming mode that does not require a full address/data cycle but instead keeps sending the data only. This would allow the AD3552R to meet the timing requirements but could result in prolonged invalid data outputs if a SPI cycle is ever missed.

# DAC11001B

These DACs have SPI interface that can be clocked at rates up to 50 MHz assuming DVDD is 5V and IOVDD is 3.3V. Since 32 bits need to be transmitted per conversion cycle, we can support a 220 Hz (1 MHz) sampling clock for the precision DAC.

See Figure 5 for the DAC11001B timing diagram.

We again use 6 clock cycles at 268 MHz for one SPI clock cycle resulting in a 44.7 MHz SPI clock. Since this interface is single lane, without DDR and requires 32-bits per cycle, we can support a 220 Hz (1 MHz) sampling clock. With a reduced IOVDD at 2.5V or 1.8V we would have to half the sampling clock.

Figure 1: Timing diagram from the ADAQ23876 data sheet (top), timing requirements (middle) and proposed FPGA timing.

Figure 2: Timing diagram from the ADAQ23878 data sheet (top), timing requirements (middle) and proposed FPGA timing.

Figure 3: Timing diagram from the AD3551R data sheet (top), timing requirements (middle) and proposed FPGA timing.

Figure 4: Timing diagram from the AD3552R data sheet (top), timing requirements (middle) and proposed FPGA timing. Notice the slight timing violation.

Figure 5: Timing diagram from the DAC11001B data sheet (top), timing requirements (second) and proposed FPGA timing (full acquisition cycle, zoomed in at start of acquisition cycle and end of acquisition cycle).